US06969351 | B2 | 2005-11-29 | Device and method for removing liquid from endogenic tissue and determining the concentrations of substance in said liquid | 2001-01-10 | The present invention relates to a device for removing liquid from endogenic tissue and determining concentrations of substances in this liquid, in which a hollow chamber (3), which is placeable in direct contact on the skin (1) covering the tissue with at least one contact opening (3.1) disposed on a contact surface (3.2) and the hollow chamber (3) is connected to a low pressure-producing element (P) via a channel (4.1). | 1. A method for removing liquid from endogenic tissue and determining concentrations of substances in this liquid, comprising: placing a device on a skin surface, the device having at least one contact opening configured on at least one hollow chamber in a contact surface,setting a low pressure at the hollow chamber and, forming a liquid reservoir underneath the skin surface,collecting the liquid in the liquid reservoir,subsequently flowing the liquid from the liquid reservoir out of the tissue and guiding the liquid with suction forces occasioned by low pressure to a sensor or a sensor system, which is disposed on the channel or is integrated in the wall of the hollow chamber, andmonitoring a lymph flow by determining the number of drops per unit of time, which are sucked into a capillary body. | 1. A method for removing liquid from endogenic tissue and determining concentrations of substances in this liquid, comprising: placing a device on a skin surface, the device having at least one contact opening configured on at least one hollow chamber in a contact surface,setting a low pressure at the hollow chamber and, forming a liquid reservoir underneath the skin surface,collecting the liquid in the liquid reservoir,subsequently flowing the liquid from the liquid reservoir out of the tissue and guiding the liquid with suction forces occasioned by low pressure to a sensor or a sensor system, which is disposed on the channel or is integrated in the wall of the hollow chamber, andmonitoring a lymph flow by determining the number of drops per unit of time, which are sucked into a capillary body. 2. The method according to claim 1, comprising setting a pressure in the hollow chamber, which is 0.05 to 1 bar below the ambient pressure, for the formation of the liquid reservoir and the transportation of the liquid. 3. The method according to claim 1, comprising transporting the liquid with a volume flow of at least 0.01 μl/min to the sensor or to the sensor system. 4. The method according to claim 1, comprising continually conveying the liquid. 5. The method according to claim 1, further comprising penetrating at least an uppermost skin layer lying above the liquid reservoir after formation of the liquid reservoir with the help of a penetration element disposed in the hollow chamber or guidable through the contact opening of the hollow chamber in the direction of the skin. 6. A device for removing liquid from endogenic tissue and determining concentrations of substances in this liquid, comprising: at least one hollow chamber having at least one contact opening disposed on a contact surface, the hollow chamber being placeable in direct contact on the skin covering the tissue, and the hollow chamber being connected to a low pressure-producing element via a channel, anda sensor or sensor system disposed on the channel, and/or a sensor or sensor system being integratable into a wall of the hollow chamber;wherein the device comprises two parts, wherein one part is present with an element producing low pressure and a second part is present with at least one sensor; and wherein both parts are interchangeable. 7. The device according to claim 6, wherein the sensor or the sensor system is disposed on the device such that it is in contact with the liquid. 8. The device according to claim 6, wherein the contact opening has an inner width of from 0.01 to 10 mm. 9. The device according to claim 6, further comprising a penetration element disposed on the hollow chamber in the direction of the skin. 10. The device according to claim 9, wherein the penetration element comprises at least a blade, a lancet, a needle, or a pin. 11. The device according to claim 9, wherein the contact opening includes an external edge, the external edge of the contact opening ... | Meinhard Knoll | A61B005/00 A61B005/02
US06904303 | B2 | 2005-06-07 | Apparatus for mapping and coagulating soft tissue in or around body orifices | 2002-12-11 | A probe that may be used to create circumferential lesions in body tissue and, in some implementations, may also be used to perform mapping functions. The probe includes a collapsible/expandable structure that supports electrodes or other operative elements against the body tissue. | 1. A mapping and coagulation apparatus, comprising: a tissue coagulation probe including a tissue coagulation probe support body defining at least one interior lumen, and an inflatable tissue coagulation structure supported on the tissue coagulation probe support body; and a mapping probe passable through the tissue coagulation probe support body lumen including a mapping probe support body, and an expandable/collapsible mapping basket supported on the mapping probe support body. | 1. A mapping and coagulation apparatus, comprising: a tissue coagulation probe including a tissue coagulation probe support body defining at least one interior lumen, and an inflatable tissue coagulation structure supported on the tissue coagulation probe support body; and a mapping probe passable through the tissue coagulation probe support body lumen including a mapping probe support body, and an expandable/collapsible mapping basket supported on the mapping probe support body. 2. A mapping and coagulation apparatus as claimed in claim 1, wherein the tissue coagulation probe support body comprises a catheter. 3. A mapping and coagulation apparatus as claimed in claim 1, wherein the inflatable tissue coagulation structure defines an interior and the tissue coagulation probe support body includes a fluid lumen operably connected to the interior of the inflatable tissue coagulation structure. 4. A mapping and coagulation apparatus as claimed in claim 1, wherein the inflatable tissue coagulation structure comprises a non-porous, relatively thermally conductive structure. 5. A mapping and coagulation apparatus as claimed in claim 4, further comprising: a fluid heating element located within the interior of the inflatable tissue coagulation structure. 6. A mapping and coagulation apparatus as claimed in claim 1, wherein the inflatable tissue coagulation structure comprises a porous structure. 7. A mapping and coagulation apparatus as claimed in claim 6, wherein the porous structure includes a porous region and a non-porous region. 8. A mapping and coagulation apparatus as claimed in claim 6, further comprising: an electrode located within the interior of the inflatable tissue coagulation structure. 9. A mapping and coagulation apparatus as claimed in claim 1, wherein the expandable/collapsible mapping basket comprises at least two splines respectively supporting at least two electrodes. 10. A mapping and coagulation apparatus, comprising: a tissue coagulation probe including a tissue coagulation probe support body defining at least one interior lumen and a distal end, and an expandable/collapsible tissue coagulation structure supported on the tissue coagulation probe support body at least 3 cm from the distal end of the tissue coagulation probe support body; and a mapping probe passable through the tissue coagulation probe support body lumen including a mapping probe support body, and a mapping structure supported on the mapping probe support body. 11. A mapping and coagulation apparatus as claimed in claim 10, wherein the tissue coagulation probe support body comprises a catheter. 12. A mapping and coagulation apparatus as claimed in claim 10, wherein the expandable/collapsible tissue coagulation structure comprises an inflatable structure. 13. A mapping and coagulation apparatus as claimed in claim 12, wherein the inflatable structure defines an interior and the tissue coagulation probe support body includes a fluid lumen operab... | Boston Scientific SciMed, Inc. | A61B005/042 A61B018/14
US06847841 | B1 | 2005-01-25 | Detector of living tissue strength and electrical resistance and activity | 2000-11-24 | A diagnostic cylindrical probe introduced through the body surfaces to the target tissue to detect the resistance of the tissues to piercing, its electrical activity and its resistance to passage of electrical current to identify the nature of the tissue and predict the nature of its pathology during needle biopsy has a sharp piercing tip attached to its end through a coiled wire to slide over the surface of an inbuilt changeable electrical resistance incorporated in an electrical circuit to detect the mechanical resistance of the tissues to piercing, a sensor at its tip incorporated in a second electrical circuit to detect the electrical impedance of the target tissue and a secondary electrically isolated sensor incorporated with the first sensor in a third electrical circuit to detect the electrical resistance of the tissues to passage of electrical current. | 1. A diagnostic cylindrical probe introduced through the body to detect the mechanical resistance of the tissue to piercing, comprising a cylindrical probe body a compressible pointed piercing tip having a base mounted at a distal end of the cylindrical probe body through a coiled wire sliding over the surface of an inbuilt changeable electrical resistance and over a metal blade, a first wire connecting the metal blade to an electrical source, the first wire passing through the body of the probe; a second wire connecting the electrical resistance to the electrical source, the second wire passing through the body of the probe; wherein the electrical source is located outside the probe; and a monitor comprising an Ammeter or Voltmeter to detect the electrical current intensity or voltage with the ability to add a registering unit using sensitive paper to monitor and record the electrical resistance, wherein the nature of the target tissue is detected by changing the mechanical resistance faced by the tip of the probe during its passage in the target tissues into a change in the electrical resistance. | 1. A diagnostic cylindrical probe introduced through the body to detect the mechanical resistance of the tissue to piercing, comprising a cylindrical probe body a compressible pointed piercing tip having a base mounted at a distal end of the cylindrical probe body through a coiled wire sliding over the surface of an inbuilt changeable electrical resistance and over a metal blade, a first wire connecting the metal blade to an electrical source, the first wire passing through the body of the probe; a second wire connecting the electrical resistance to the electrical source, the second wire passing through the body of the probe; wherein the electrical source is located outside the probe; and a monitor comprising an Ammeter or Voltmeter to detect the electrical current intensity or voltage with the ability to add a registering unit using sensitive paper to monitor and record the electrical resistance, wherein the nature of the target tissue is detected by changing the mechanical resistance faced by the tip of the probe during its passage in the target tissues into a change in the electrical resistance. 2. A diagnostic cylindrical probe introduced through the body to detect the electrical resistance of the target tissue comprising a pointed piercing tip including two electrically isolated electrodes connected to an electrical circuit to detect the electrical resistance of the tissue to passage of an electrical current; a first wire running inside the body of the probe with one of its terminals at the tip of the probe and the other terminal connected to an electrical source; the electrical source is located outside the probe; an Ammeter or Voltmeter to detect the electrical current intensity or voltage with possibility of adding a registering unit on sensitive paper and a second wire running inside the body of the probe with one end connected to the electrical source &amp; and the other end is located at the tip of the probe near the end of the said first wire, wherein the nature of the target tissue is detected by monitoring the electrical resistance exerted by the tissue surrounding the tip. 3. A diagnostic cylindrical probe introduced through the bode to detect the electrical activity of target tissue, comprising an electrical circuit to detect the electrical activity; a pointed piercing tip electrically isolated from the probe by a transverse insulator to detect the electrical activity of the target tissue; a first wire running inside the body of the probe with one of its terminals at the tip of the probe and the other terminal connected to an electrical activity monitor; and a second wire connecting the electrical activity monitor to the body of the probe, which operates as a neutral isoelectric point; wherein the nature of the target tissue is detected by monitoring the electrical activity of the tissue surrounding the tip. | Mohamed Khaled Mohamed El Hatw | A61B005/05
US06911916 | B1 | 2005-06-28 | Method and apparatus for accessing medical data over a network | 2000-07-13 | A method and apparatus for retrieving, accessing, and storing medical data relating to a patient during a medical procedure. The invention provides a single interface to many disparate forms of medical data, which is accessible over a local area network, wide area network, direct connection, or combinations thereof. In one embodiment, an operating room control system for use during a medical procedure on a patient includes an input device, a display device, and a controller that is coupled to the input device and the display device. The controller receives one or more user inputs, transmits a command to a server located outside of the operating room to retrieve medical data, receives the medical data from the server, and displays the medical data on the display device. | 1. An operating room control system for use during a medical procedure on a patient, comprising: a plurality of operating room devices; a voice input device capable of receiving a spoken device command that has an associated device qualifier, and a data retrieval command having an associated patient record identifier; a display device; and a controller coupled to the input device and the display device, the controller capable of receiving the data retrieval command and the device command, wherein the controller transmits a command to a server to retrieve patient information in response to the data retrieval command so that desired patient data can be selected in response to the patient record identifier, wherein the controller receives the patient information from the server and displays the patient information on the display device, and wherein the controller transmits a control command to a selected operating room device in response to the device command, the medical operating room device selected from the plurality of operating room devices in response to the device qualifier so as to allow voice manipulation of an operating room environment. | 1. An operating room control system for use during a medical procedure on a patient, comprising: a plurality of operating room devices; a voice input device capable of receiving a spoken device command that has an associated device qualifier, and a data retrieval command having an associated patient record identifier; a display device; and a controller coupled to the input device and the display device, the controller capable of receiving the data retrieval command and the device command, wherein the controller transmits a command to a server to retrieve patient information in response to the data retrieval command so that desired patient data can be selected in response to the patient record identifier, wherein the controller receives the patient information from the server and displays the patient information on the display device, and wherein the controller transmits a control command to a selected operating room device in response to the device command, the medical operating room device selected from the plurality of operating room devices in response to the device qualifier so as to allow voice manipulation of an operating room environment. 2. The operating room control system of claim 1 wherein the server is located outside the operating room. 3. The operating room control system of claim 1 wherein the patient information includes one or more of the following: an MRI, CT, X-ray, still-frame video, moving video, ultrasound image, patient records, patient notes, and patient schedule. 4. The operating room control system of claim 1 wherein the server is coupled to the controller via one or more of the following: a local area network, wide area network, Internet, and Intranet. 5. The operating room control system of claim 1 wherein the controller includes a memory containing an application program, and a processor, said processor to execute the application program to cause the patient information to be displayed on the display device in a web browseable format. 6. The operating room control system of claim 1 wherein the input device comprises one or more of the following: a microphone, keyboard, mouse, foot pedal, and hand-held device. 7. The operating room control system of claim 1, wherein the controller compares the device command with a list of stored control commands associated with the selected operating room device, and transmits a control signal to a medical instrument for controlling the operating room device if there is a match, wherein the operating room devices comprise one or more of the following: a robotic arm, electrocautery device, operating room table, operating room lights, insufflator, and camera. 8. The operating room control system of claim 1 wherein the controller captures one or more units of patient data, transmits the one or more units of data and a request to a second server located outside the operating room to store the one or more units of patient data. 9. The operating room control system of claim 8 where... | The Cleveland Clinic Foundation Intuitive Surgical, Inc. | G05B023/02 A61B005/00 G10L021/00
US06845259 | B2 | 2005-01-18 | MRI compatible guide wire | 2003-01-02 | An improved intracorporeal device such as a guide wire or other guiding member for use within a patient's body that is at least in part visible under magnetic resonance imaging (MRI) but is not detrimentally affected by the imaging is disclosed. The intercorporeal device has a non-conductive proximal core section, an essentially non-magnetic metallic distal core section that is preferably more flexible than the proximal core section, and that has an MRI visible member or coil in the distal section. | 1. A guidewire, comprising: a core having a proximal core section with proximal and distal ends and a distal core section with proximal and distal ends; and a polymer joint securing the distal end of the proximal core section to the proximal end of the distal core section; wherein the polymer joint is not covered by a hypotube. | 1. A guidewire, comprising: a core having a proximal core section with proximal and distal ends and a distal core section with proximal and distal ends; and a polymer joint securing the distal end of the proximal core section to the proximal end of the distal core section; wherein the polymer joint is not covered by a hypotube. 2. The guidewire of claim 1, wherein the polymer joint encases the distal end of the proximal core section and the proximal end of the distal core section. 3. The guidewire of claim 1, wherein the distal end of the proximal core section and the proximal end of the distal core section both include a taper. 4. The guidewire of claim 1, wherein the distal end of the proximal core section and the proximal end of the distal core section have an undulated shape. 5. The guidewire of claim 1, wherein the polymer joint is formed at least in part of one or more thermoplastic polymeric materials selected from the group consisting of polyester, polyetheretherketone, ABS, epoxy, copolymers, and blends thereof. 6. A process for joining two guide wire core materials without using a hypotube, comprising: providing a proximal core section having a distal end and a distal core section having a proximal end; grinding at least one of the distal end of the proximal core section and the proximal end of the distal core section; shaping at least one of the distal end of the proximal core section and the proximal end of the distal core section into a geometric pattern; and encasing the distal end of the proximal core section and the proximal end of the distal core section with a polymer. 7. The process of claim 6, wherein the polymer is formed at least in part of one or more thermoplastic polymeric materials selected from the group consisting of polyester, polyetheretherketone, ABS, epoxy, copolymers, and blends thereof. 8. The process of claim 6, further comprising forming a taper on at least one of the distal end of the proximal core and the proximal end of the distal core. 9. The process of claim 6, wherein the geometric pattern is an undulated shape. 10. A process for joining two wire core materials, comprising: providing a proximal core section having a distal end and a distal core section having a proximal end; shaping at least one of the distal end of the proximal core section and the proximal end of the distal core section into an undulated shape; and encasing the distal end of the proximal core section and the proximal end of the distal core section with a polymer. 11. The process of claim 10, further comprising grinding at least one of the distal end of the proximal core section and the proximal end of the distal core section. 12. The process of claim 11, wherein grinding at least one of the distal end of the proximal core section and the proximal end of the distal core section to include a taper. 13. The process of claim 10, wherein the polymer is formed at least in part of one or more thermoplastic polymeric materials se... | Advanced Cardiovascular Systems, Inc. | A61B005/05
US06875174 | B2 | 2005-04-05 | General-purpose medical instrumentation | 2002-04-18 | A general-purpose, low-cost system provides comprehensive physiological data collection, with extensive data object oriented programmability and configurability for a variety of medical as well as other analog data collection applications. In a preferred embodiment, programmable input signal acquisition and processing circuits are used so that virtually any analog and/or medical signal can be digitized from a common point of contact to a plurality of sensors. A general-purpose data routing and encapsulation architecture supports input tagging and standardized routing through modern packet switch networks, including the Internet; from one of multiple points of origin or patients, to one or multiple points of data analysis for physician review. The preferred architecture further supports multiple-site data buffering for redundancy and reliability, and real-time data collection, routing, and viewing (or slower than real-time processes when communications infrastructure is slower than the data collection rate). Routing and viewing stations allow for the insertion of automated analysis routines to aid in data encoding, analysis, viewing, and diagnosis. | 1. A remote patient management system, comprising: at least one sensor outputting patient physiologic data; a patient monitor operative to receive the physiologic data and automatically download the data in encrypted form to a network over a wired telephone connection; and a server interfaced to the network providing a viewing environment enabling a clinician to access and decrypt the patient physiologic data for private analysis or diagnostic purposes. | 1. A remote patient management system, comprising: at least one sensor outputting patient physiologic data; a patient monitor operative to receive the physiologic data and automatically download the data in encrypted form to a network over a wired telephone connection; and a server interfaced to the network providing a viewing environment enabling a clinician to access and decrypt the patient physiologic data for private analysis or diagnostic purposes. 2. The remote patient management system according to claim 1, wherein the patient physiologic data output by the sensor is cardiac-related. 3. The remote patient management system according to claim 2, wherein the patient physiologic data output by the sensor includes electrocardiogram information. 4. The remote patient management system according to claim 1, wherein the patient monitor is portable or wearable. 5. The remote patient management system according to claim 1, wherein the patient physiologic data is self-descriptive to facilitate proper routing and retrieval through the server. 6. The remote patient management system according to claim 1, wherein the network is the Internet. 7. The remote patient management system according to claim 1, wherein at least a portion of the transfer of the physiologic data monitor from the sensor to the server occurs through a wireless communication link. 8. A remote patient management system, comprising: at least one sensor outputting cardiac-related patient data; a portable patient monitor operative to receive the patient data and automatically download the data in encrypted form to an Internet server over a wired telephone connection; and a viewing environment resident on the server enabling a clinician to access and decrypt the patient's cardiac data for private analysis or diagnostic purposes. 9. The remote patient management system according to claim 8, wherein the cardiac-related data includes electrocardiogram information. 10. The remote patient management system according to claim 8, wherein the cardiac-related data is self-descriptive to facilitate proper routing and retrieval through the server. 11. The remote patient management system according to claim 8, wherein at least a portion of the transfer of the physiologic data monitor from the sensor to the server occurs through a wireless communication link. | Cybernet Systems Corporation | A61B005/00
US07225010 | B1 | 2007-05-29 | System and method for confocal imaging within dermal tissue | 2000-09-11 | An improved system for confocal imaging within dermal tissue of a patient is provided which minimizes instability in confocal images by reducing the relative motion of the tissue with respect to the confocal imaging optics of the system. The system includes a mechanism for maintaining an area of skin tissue under stress by application of force at the edges of the area, and an imaging head coupled to this mechanism for imaging the stressed skin. The mechanism includes a mechanical structure, such as a platen, brace, or attachment, which both supports the imaging head of the system and applies stress to a limited surface area of the tissue to minimize skin motion during confocal imaging. | 1. A system for dermatological examination of the skin tissue of a patient comprising: a plate;means for maintaining an area of the skin tissue under stress by application of force against said plate, in which said maintaining means comprises a member extending from said plate; andan imaging head, coupled to said maintaining means, capable of producing one or more images representing optically formed sections of said stressed skin tissue from light returned through said plate from focused light under the surface of said skin tissue to provide therewith an integrated assembly in which said stress applied to said tissue stabilizes said skin tissue with respect to said imaging head, wherein said imaging head comprises at least an objective lens for focusing the light under the surface of said skin tissue and collecting the returned light through said plate, and at least one actuator capable of moving said objective lens to change location of said focusing in said tissue. | 1. A system for dermatological examination of the skin tissue of a patient comprising: a plate;means for maintaining an area of the skin tissue under stress by application of force against said plate, in which said maintaining means comprises a member extending from said plate; andan imaging head, coupled to said maintaining means, capable of producing one or more images representing optically formed sections of said stressed skin tissue from light returned through said plate from focused light under the surface of said skin tissue to provide therewith an integrated assembly in which said stress applied to said tissue stabilizes said skin tissue with respect to said imaging head, wherein said imaging head comprises at least an objective lens for focusing the light under the surface of said skin tissue and collecting the returned light through said plate, and at least one actuator capable of moving said objective lens to change location of said focusing in said tissue. 2. The system according to claim 1 wherein said member comprises a platen which is positionable with respect to said patient, said platen having an orifice, and said imaging head being coupled to said platen and positioned for imaging through said surface and said plate; and first means for moving said platen from an up position in which said platen is spaced from said patient to a down position onto the surface of said skin tissue of the patient, whereby in said down position the force of said platen provides said stressed skin tissue within said orifice. 3. The system according to claim 2 wherein said maintaining means further comprises: second means coupled to said platen for moving said platen with respect to said patient and temporarily fixing the position of said platen so that said skin tissue to be imaged is located within said orifice when said first means moves said platen to said down position. 4. The system according to claim 2 wherein said maintaining means further comprises: means coupled to said imaging head and said platen for translating said imaging head with respect to said orifice of said platen. 5. The system according to claim 2 wherein said orifice has said plate positioned therein. 6. The system according to claim 1 wherein said imaging head is a confocal imaging head having confocal imaging optics for providing images of section of said stressed skin tissue. 7. The system according to claim 1 wherein said member comprises a platen having means through which said tissue is imaged by said imaging head. 8. The system according to claim 7 wherein said platen is positionable with respect to said tissue. 9. The system according to claim 1 further comprising means for fixing the position of said member with respect to said patient. 10. The system according to claim 1 wherein said imaging head is movable with respect to said maintaining means to position said imaging head with respect to said stressed tissue. 11. The system according to claim 1 wherein sai... | Lucid, Inc. | A61B005/00
US07231245 | B2 | 2007-06-12 | System and method of assessment of neurological conditions using EEG | 2003-01-06 | The present invention is a system and method that produces features and indices that indicate the presence or absence of a disease or condition, or of the progression of a disease or condition. The system and method of the present invention also produce features and indices that predict responsiveness to medication from a premedication baseline. The system and method of the present invention further incorporates a testing methodology to improve the performance characteristics of the features or indices. To obtain such features and indices, time domain, power spectrum, bispectrum and higher order spectrum values are derived from biopotential signals taken from the subject being tested. | 1. A system of assessing neurological conditions comprising: one or more electrodes for acquiring biopotential signals,a processor for deriving, from said biopotential signals, Kth-order spectral values, where K is an integer greater than 1, and deriving from said Kth order spectral values features indicative of neurological conditions, said processor using a differential testing methodology to derive said features. | 1. A system of assessing neurological conditions comprising: one or more electrodes for acquiring biopotential signals,a processor for deriving, from said biopotential signals, Kth-order spectral values, where K is an integer greater than 1, and deriving from said Kth order spectral values features indicative of neurological conditions, said processor using a differential testing methodology to derive said features. 2. The system of claim 1 further comprising an infusion device connected to said processor for administering a controlled dose of a pharmacological agent as part of said differential testing methodology. 3. A method of assessing neurological conditions comprising the steps of acquiring biopotential signals,deriving, from said biopotential signals, Kth-order spectral values, where K is an integer greater than 1, andderiving at least one feature indicative of neurological conditions from said Kth order spectral using a differential test methodology. 4. The method of claim 3 further comprising the step of administering a controlled dose of a pharmacological agent as part of said differential test methodology. 5. A method of assessing neurological conditions comprising the steps of: acquiring biopotential signals from a human subject,deriving, from said biopotential signals, Kth-order spectral values, where K is an integer greater than 1, andderiving from said Kth order spectral values a prediction of the effectiveness of a specific therapeutic intervention at least one hour after said intervention. 6. A system of assessing neurological conditions comprising one or more electrodes for acquiring biopotential signals anda processor for deriving, from said biopotential signals, Kth-order spectral values, where K is an integer greater than 1 and for deriving from said Kth order spectral values a prediction of the effectiveness of a specific therapeutic intervention at least one hour after said intervention. 7. A system of assessing neurological conditions comprising: one or more electrodes for acquiring biopotential signals, anda processor for deriving, from said biopotential signals, at least one feature indicative of neurological conditions, one of said at least one feature being a prediction of the effectiveness of a therapeutic intervention in achieving a desired behavioral effect at least one hour after said intervention,wherein said at least one feature indicative of neurological conditions is a Kth -order spectral value, where K is an integer greater than 1. 8. A system of assessing neurological conditions comprising: one or more electrodes for acquiring biopotential signals, anda processor for deriving, from said biopotential signals, at least one feature indicative of neurological conditions, one of said at least one feature being a prediction of the effectiveness of a therapeutic intervention in achieving a desired behavioral effect at least one hour after said intervention,wherein said at least one feature indicative of neur... | Aspect Medical Systems, Inc. | A61B005/04
US07047057 | B2 | 2006-05-16 | Simultaneous multiwavelength TPSF-based optical imaging | 2002-01-22 | The TPSF-based imaging technique uses multiple wavelengths to image an object simultaneously. Acquisition time of an image can be shortened without sacrificing the effective amount or quality of raw imaging data acquired. A plurality of distinguishable wavelengths may be used simultaneously at different injection-detection positions to acquire simultaneously a plurality of TPSF-based imaging data points for the different injection-detection positions. The multiple wavelengths may provide complementary information about the object being imaged. | 1. A method of Temporal Point Spread Function (TPSF)-based optical imaging comprising the steps of: injecting light at a plurality of wavelengths into an object to be imaged at one or more injection positions;collecting the injected light after diffusing in the object at one or more collection positions;spectrally separating said collected light into individual wavelength signals;directing individual wavelength signals at distinct locations on a CCD camera;simultaneously detecting said individual wavelength signals using said CCD camera to obtain separate TPSF-based data for each of the wavelengths; andwherein said separating and said detecting preserves time-based information of each individual wavelength signal. | 1. A method of Temporal Point Spread Function (TPSF)-based optical imaging comprising the steps of: injecting light at a plurality of wavelengths into an object to be imaged at one or more injection positions;collecting the injected light after diffusing in the object at one or more collection positions;spectrally separating said collected light into individual wavelength signals;directing individual wavelength signals at distinct locations on a CCD camera;simultaneously detecting said individual wavelength signals using said CCD camera to obtain separate TPSF-based data for each of the wavelengths; andwherein said separating and said detecting preserves time-based information of each individual wavelength signal. 2. The method as claimed in claim 1, wherein said plurality of wavelengths provide different information about said object. 3. The method as claimed in claim 2, wherein said imaging is medical imaging, and said different information is complementary to provide physiological information. 4. The method as claimed in claim 1, wherein said step of injecting comprises: generating light from a laser light source; andswitching said light from said light source onto one of a plurality of optical fibers each corresponding to one of said injection positions. 5. The method as claimed in claim 4, wherein said light source comprises a plurality of lasers each operating at one of said plurality of wavelengths. 6. The method as claimed in claim 5, wherein an output of said lasers are combined prior to said switching, said injection position being the same at one time for all of said wavelengths. 7. The method as claimed in claim 6, wherein said plurality of wavelengths provide different information about said object. 8. The method as claimed in claim 1, wherein said step of collecting comprises positioning a bundle of optical fibers at each one of said one or more collection positions, each fiber of said bundle being for a separate one of said wavelengths. 9. The method as claimed in claim 8, wherein said step of collecting comprises placing a faceplate bandpass filter over said CCD camera and positioning together groups of optical fibers for a same wavelength over said filter. 10. The method us claimed in claim 1, wherein said imaging is medical imaging. 11. The method as claimed in claim 1, wherein said step of injecting comprises injecting said light at each of said wavelengths simultaneously at different injection positions, wherein said injected light traveling to at least some of said collection positions comprises light of more than one of said wavelengths. 12. The method as claimed in claim 11, wherein said plurality of wavelengths provide a same information about said object, said method allowing faster acquisition of said data over an imaging area. 13. A Temporal Point Spread Function (TPSP)-based optical imaging apparatus comprising: at least one source providing light at a plurality of wavelengths;at least one injection port co... | ART. Advanced Research Technologies Inc. | A61B005/00
US07107253 | B1 | 2006-09-12 | Computer architecture and process of patient generation, evolution and simulation for computer based testing system using bayesian networks as a scripting language | 2000-04-05 | A method and system for patient generation and evolution for a computer-based testing system and/or expert system. One or more belief networks, which describe parallel health state networks are accessed by a user or a computer. A knowledge base, at least in part, is scripted from the one or more belief networks by the computer. A model patient at least in part, is instantiated by the computer from the scripted knowledge base. Optionally, the model patient is evolved by the computer in accordance with the parallel health state networks and responsive to a received course of action. | 1. A method for evaluating or educating a user comprising the steps of: (a) generating a plurality of parallel health state networks; (b) generating at least one first Bayesian network which describes each of the plurality of parallel health state networks; (c) generating at least one second Bayesian network which describes rates of progression within and/or between said plurality of parallel health state networks, and describes task factors that affect the rates of progression; (d) generating at least one third Bayesian network which supports reveal structures to limit display of patient test data to patient test data specifically requested by the user; (e) generating at least one fourth Bayesian network which supports plan critiques of queries of and treatment prescribed by the user; (f) scripting a knowledge base from the at least one first Bayesiaii network and the at least one second Bayesian network; (g) instantiating a model patient, at least in part, from the scripted knowledge base; (h) receiving one of a course of action and a query for a specific medical finding concerning the model patient from the user responsive to the instantiated model patient; (i) displaying, if the query is received, the specific medical finding to the user based at least in part on the at least one third Bayesian network, and repeating step (h); (j) evolving the model patient in accordance with the plurality of parallel health state networks and responsive to the received course of action; (k) repeating the steps (h) through (j) until the user has completed treatment of the model patient; (l) generating an optimum combination of treatment and queries based, at least in part, on the at least one fourth Bayesian network and the instantiated model patient; and (m) evaluating the query and the treatment by the user in comparison to the generated optimum combination of treatment and queries. | 1. A method for evaluating or educating a user comprising the steps of: (a) generating a plurality of parallel health state networks; (b) generating at least one first Bayesian network which describes each of the plurality of parallel health state networks; (c) generating at least one second Bayesian network which describes rates of progression within and/or between said plurality of parallel health state networks, and describes task factors that affect the rates of progression; (d) generating at least one third Bayesian network which supports reveal structures to limit display of patient test data to patient test data specifically requested by the user; (e) generating at least one fourth Bayesian network which supports plan critiques of queries of and treatment prescribed by the user; (f) scripting a knowledge base from the at least one first Bayesiaii network and the at least one second Bayesian network; (g) instantiating a model patient, at least in part, from the scripted knowledge base; (h) receiving one of a course of action and a query for a specific medical finding concerning the model patient from the user responsive to the instantiated model patient; (i) displaying, if the query is received, the specific medical finding to the user based at least in part on the at least one third Bayesian network, and repeating step (h); (j) evolving the model patient in accordance with the plurality of parallel health state networks and responsive to the received course of action; (k) repeating the steps (h) through (j) until the user has completed treatment of the model patient; (l) generating an optimum combination of treatment and queries based, at least in part, on the at least one fourth Bayesian network and the instantiated model patient; and (m) evaluating the query and the treatment by the user in comparison to the generated optimum combination of treatment and queries. 2. The method according to claim 1, wherein said plurality of parallel health state networks describe at least one of a plurality of primary networks defining disease evolutions, a plurality of secondary networks defining risk factors affecting progression through a primary network of the plurality of primary networks, and a plurality of tertiary networks defining at least one of causal probabilistic medical complications attributed to at least one stage in the primary network and medical complications attributed to management of the at least one stage. 3. A computer readable medium including instructions being executed by a computer, the instructions instructing the computer to execute an educational or testing system for physicians, the instructions including: (a) accessing at least one first belief network, which describes a plurality of parallel health state networks; (b) scripting a knowledge base, at least in part, from the at least one first belief network; and (c) instantiating a model patient, at least in part, from the scripted knowledge base. 4. The computer read... | American Board of Family Practice, Inc. | G06F017/00 A61B005/00
US07103398 | B2 | 2006-09-05 | Electrical sensing and/or signal application device | 2000-11-01 | The electrical signal sensing and/or signal application device for sensing electrical signals on a surface and/or for applying electrical signals to a surface, particularly of human or animal skin or other organs or tissues, comprises a substrate (12) having first and second major surfaces (14), and a plurality of electrodes (16) arranged on the first major surface (14) of the substrate (12) and projecting therefrom, each of the electrodes (16) comprising a pointed contact end (36) facing away from the substrate (12) for contacting the surface, and a base end (26) facing towards the substrate (12). On the first or second or both major surfaces (14) of the substrate (12) there is arranged a plurality of first contact pads (22) for electrical connection to contact elements of an evaluation and/or driving means for evaluating the sensed electric signals and/or applying signals to the electrodes (16), the first contact pads (22) being electrically connected to the electrodes (16). The first major surface (14) of the substrate (12) comprises, for each electrode (16), a second contact pad (18) the base ends (26) of the electrodes electrically and mechanically connected to the second contact pads (18). Each electrode (18) is generated by bonding an electrically conductive bonding wire to a second contact pad (18) for generating a thickened socket portion (24) of the electrode comprising its base end (26), and by tearing the bonded bonding wire off the socket portion (24) so as to generate a pointed portion (34) of the electrode (16) comprising its pointed contact end (36) and projecting from the socket portion (24) thereof. | 1. An electrical signal sensing and/or signal application device for sensing electrical signals on a surface and/or for applying electrical signals to a surface, of human or animal skin or other biological organs or tissues, comprising a substrate having first and second major surfaces,a plurality of electrodes arranged on the first major surface of said substrate and projecting therefrom, each of said electrodes comprising a pointed contact end facing away from said substrate for contacting the surface, and a base end facing towards said substrate, anda plurality of first contact pads arranged on said first or second or both major surfaces of said substrate for electrical connection to contact elements of an evaluation and/or driving means for evaluating the sensed electric signals and/or applying signals to said electrodes, said first contact pads being electrically connected to said electrodes,wherein said first major surface of said substrate comprises, for each electrode, a second contact pad, said base ends of said electrodes being electrically and mechanically connected to said second contact pads, andwherein each electrode is generated by bonding an electrically conductive bonding wire to a second contact pad for generating a thickened socket portion of said electrode comprising its base end, and by tearing the bonded bonding wire off the socket portion so as to generate a pointed portion of said electrode comprising its pointed contact end and projecting from the socket portion thereof. | 1. An electrical signal sensing and/or signal application device for sensing electrical signals on a surface and/or for applying electrical signals to a surface, of human or animal skin or other biological organs or tissues, comprising a substrate having first and second major surfaces,a plurality of electrodes arranged on the first major surface of said substrate and projecting therefrom, each of said electrodes comprising a pointed contact end facing away from said substrate for contacting the surface, and a base end facing towards said substrate, anda plurality of first contact pads arranged on said first or second or both major surfaces of said substrate for electrical connection to contact elements of an evaluation and/or driving means for evaluating the sensed electric signals and/or applying signals to said electrodes, said first contact pads being electrically connected to said electrodes,wherein said first major surface of said substrate comprises, for each electrode, a second contact pad, said base ends of said electrodes being electrically and mechanically connected to said second contact pads, andwherein each electrode is generated by bonding an electrically conductive bonding wire to a second contact pad for generating a thickened socket portion of said electrode comprising its base end, and by tearing the bonded bonding wire off the socket portion so as to generate a pointed portion of said electrode comprising its pointed contact end and projecting from the socket portion thereof. 2. The electrical signal sensing and/or signal application device according to claim 1, wherein said substrate is flexible. 3. The electrical signal sensing and/or signal application device according to claim 1, wherein said substrate is rigid. 4. The electrical signal sensing and/or signal application device according to claim 1, wherein said second contact pads are electrically connected to the first contact pads of said substrate via electrically conductive traces arranged on said substrate. 5. The electrical signal sensing and/or signal application device according to claim 1, wherein said first contact pads are arranged on said second major surface of said substrate and flush with the second contact pads and are electrically connected to said second contact pads by conductor portions extending though said substrate. 6. The electrical signal sensing and/or signal application device according to claim 1, wherein said socket portions of said electrodes are connected to said second contact pads by thermal bonding or ultrasonic bonding or both. 7. The device according to claim 1 in a biomedical apparatus. 8. The device according to claim 7 wherein the biomedical apparatus is a dermal diagnostic device. 9. The device according to claim 7 wherein the biomedical apparatus does one or both of sensing and stimulating matter selected from the group consisting of biological tissue, organs, and cells. 10. A method for manufacturing an electrical signal s... | 3M Innovative Properties Company | A61B005/04 A61N001/04
US07029655 | B2 | 2006-04-18 | Magnetic resonance imaging agents for in vivo labeling and detection of amyloid deposits | 2001-10-04 | The invention relates to a novel class of magnetic resonance imaging agents that can cross the blood brain barrier and provide accurate magnetic resonance imaging of the brain, especially magnetic resonance images of amyloid deposits associated with Alzheimer's disease. | 1. An MRI agent comprising: a) a chelator; b) a lanthanide metal ion; and, c) an amyloid binding moiety (ABM) comprising (trans, trans)-1-bromo-2,5-bis-(3-hydroxycarbonyl-4-hydroxy)-styrylbenzene (BSB) or a derivative thereof. | 1. An MRI agent comprising: a) a chelator; b) a lanthanide metal ion; and, c) an amyloid binding moiety (ABM) comprising (trans, trans)-1-bromo-2,5-bis-(3-hydroxycarbonyl-4-hydroxy)-styrylbenzene (BSB) or a derivative thereof. 2. An MRI agent according to claim 1 wherein said lanthanide metal ion is gadolinium III (Gd+3 or Gd(III)). 3. An MRI agent according to claim 1 wherein said lanthanide metal ion is dysprosium (Dy+3 or Dy(III)). 4. An MRI agent according to claim 1 wherein said lanthanide metal ion is europium (Eu+3 or Eu(III)). 5. An MRI agent according to claim 1, 2, 3, or 4 wherein said chelator is 1,4,7,10-tetraazacyclododecane-N,N′N″,N′″-tetracetic acid (DOTA) or a derivative thereof. 6. An MRI agent according to claim 1, 2, 3 or 4 wherein said chelator is diethylenetriaminepentaacetic acid (DTPA) or a derivative thereof. 7. An MRI agent according to claim 1, 2, 3, or 4 wherein said chelator is 1,4,7,10-tetraazacyclododecane-N,N′N″,N′″-tetraethylphosphorus (DOTEP) or a derivative thereof. 8. An MRI agent composition according to claim 1 wherein said chelator is directly attached to said ABM. 9. An MRI agent composition according to claim 1 wherein said chelator is indirectly attached to said ABM via a linker. 10. An MRI agent composition according to claim 9 wherein said linker is attached to an internal position of said ABM. 11. An MRI agent composition according to claim 9 wherein said linker is attached to a terminus of said ABM. 12. An MRI agent composition according to claim 9, 10, or 11 wherein said linker is selected from the group consisting of an aryl or alkyl group. 13. An MRI agent composition according to claim 1 having the formula: 14. A pharmaceutical composition comprising an MRI agent composition as in any one of claims 1, 4, 8-11 and 13, and a pharmaceutically acceptable carrier. 15. A method of magnetic resonance imaging of a cell, tissue or patient comprising administering an MRI agent as in any one of claims 1-4, 8-11 and 13 to a cell, tissue or patient and rendering a magnetic resonance image of said cell, tissue or patient. | California Institute of Technology | A61B005/055
US07010340 | B2 | 2006-03-07 | Methods, systems, and associated implantable devices for dynamic monitoring of physiological and biological properties of tumors | 2002-02-18 | Methods of monitoring and evaluating the status of a tumor undergoing treatment includes monitoring in vivo at least one physiological parameter associated with a tumor in a subject undergoing treatment, transmitting data from an in situ located sensor to a receiver external of the subject, analyzing the transmitted data, repeating the monitoring and transmitting steps at sequential points in time and evaluating a treatment strategy. The method provides dynamic tracking of the monitored parameters over time. The method can also include identifying in a substantially real time manner when conditions are favorable for treatment and when conditions are unfavorable for treatment and can verify or quantify how much of a known drug dose or radiation dose was actually received at the tumor. The method can include remote transmission from a non-clinical site to allow oversight of the tumor's condition even during non-active treatment periods (in between active treatments). The disclosure also includes monitoring systems with in situ in vivo biocompatible sensors and telemetry based operations and related computer program products. | 1. An implantable biocompatible sensor, comprising: a sensor body having at least one sensor element configured for in vivo placement proximate a tumor site and to provide data responsive to at least one sensed internal parameter, and a transmitter coil and associated electronic components configured for wireless transmittal of the sensor data to a spatially remote receiver, said sensor body comprising a biocompatible material, wherein the sensor is inductively powered, wherein said sensor body is sized and configured to be chronically implanted in a subject, and wherein the at least one sensor element comprises a MOSFET-based radiation sensor element that is configured to detect ionizing radiation delivered to the subject. | 1. An implantable biocompatible sensor, comprising: a sensor body having at least one sensor element configured for in vivo placement proximate a tumor site and to provide data responsive to at least one sensed internal parameter, and a transmitter coil and associated electronic components configured for wireless transmittal of the sensor data to a spatially remote receiver, said sensor body comprising a biocompatible material, wherein the sensor is inductively powered, wherein said sensor body is sized and configured to be chronically implanted in a subject, and wherein the at least one sensor element comprises a MOSFET-based radiation sensor element that is configured to detect ionizing radiation delivered to the subject. 2. An implantable biocompatible sensor unit according to claim 1, wherein the sensor is configured to be implanted in the subject for greater than four weeks, and wherein the MOSFET-based radiation sensor element is configured to detect radiation present proximate the tumor site for each of a plurality of temporally spaced apart therapeutic oncology radiation treatment sessions. 3. An implantable biocompatible sensor unit according to claim 1, wherein the at least one sensed internal parameter is a plurality of different internal parameters including radiation dose and temperature proximate the tumor site. 4. An implantable biocompatible sensor device comprising: a first sensor body configured for chronic in viva implantation proximate a tumor site in a subject and to output data corresponding to at least one predetermined parameter and a transmitter coil and associated electronic components configured for wireless transmittal of the sensor output data to a spatially remote receiver, wherein the first sensor body comprises a radiation sensor element with a MOSFET-based radiation detection circuit, the radiation detection circuit being configured to provide MOSFET-based radiation data corresponding to detected radiation delivered to the subject in an oncology therapeutic treatment; anda second sensor body having at least one sensor element configured for in vivo implantation at a location that is spaced apart from the first sensor body, the second sensor body configured to output data corresponding to at least one sensed internal predetermined parameter, the second sensor body comprising a transmitter coil and associated electronic components configured for wireless transmittal of the sensor output data to the spatially remote receiver. 5. An implantable biocompatibe sensor device according to claim 4, wherein said first sensor body is configured with an anchor portion to positionally attach to a desired portion thereby allowing attachment to a surface of or at a depth into the tumor. 6. An implantable biocompatible sensor device according to claim 4, wherein the first and second sensor bodies include internal electronic circuitry that is operative over an oncology therapeutic dose range of external beam radiation, and whe... | North Carolina State University Sicel Technologies, Inc | A61B006/00 A61B005/07
US07092748 | B2 | 2006-08-15 | System and method for the tomography of the primary electric current of the brain and of the heart | 2002-07-26 | A three-dimensional map of the probability of brain or heart functional states is obtained based on electric or magnetic signals, or a combination of both measured in the surface of body. From signals, the statistical descriptive parameters are obtained and a map of its distribution is calculated. The map is the inverse solution of the problem based upon: a) the restriction of the solution to structures with high probability of generating electric activity using for this restriction an Anatomical Atlas and b) imposing that the solution belong to a pre-specified functional space. The probability is determined that this map belongs to a test group. The spatial and temporal correlations of the map are modeled as well as their dependence on experimental covariables. The resulting probabilities are coded in a pseudocolor scale and they are superimposed on a Anatomical Atlas for their interactive three-dimensional visualization. | 1. Method for the Tomography of the Primary Electric Current (TPEC) of the brain (TPECc) and of the heart (TPECk) comprising the following steps of: pre-processing of a plurality of signals recorded from a subject for the elimination of artifacts and of non-physiological frequency components;conversion of a plurality of coordinates of sensors to a reference system of a respective head or torso by means of the application of a visco-elastic deformation, that puts into correspondence the coordinates of a plurality of external anatomical markers of the subject with those of a structural image stored in a computing unit (CU);calculation of respective electric Kernel and magnetic Kernel linear operators that predict the o(t) that would be produced in the subject by the presence of the PEC j(t) in any part of the brain or heart;calculation of the TPEC by means of Bayesian Hierarchical Estimation Procedure that determines the primary electric current j(t) of the brain or heart using structural information obtained from an anatomical atlas, metabolic information obtained from functional images, and an assumption that the solution belongs to a space of Besov space of given smoothness or determined by a norm in a Megadictionary;calculation of descriptive parameters either on the basis of the observations o(t) directly, or of the j(t);calculation of a probability that all or some of the DP are similar to those of a given test group; andperforming the step wherein the obtained TPECc is coded by means of a pseudocolor scale and is overlaid on the anatomical atlas, a scale being fitted according to the probability previously calculated, whereby only statistically significant anatomical sites are highlighted. | 1. Method for the Tomography of the Primary Electric Current (TPEC) of the brain (TPECc) and of the heart (TPECk) comprising the following steps of: pre-processing of a plurality of signals recorded from a subject for the elimination of artifacts and of non-physiological frequency components;conversion of a plurality of coordinates of sensors to a reference system of a respective head or torso by means of the application of a visco-elastic deformation, that puts into correspondence the coordinates of a plurality of external anatomical markers of the subject with those of a structural image stored in a computing unit (CU);calculation of respective electric Kernel and magnetic Kernel linear operators that predict the o(t) that would be produced in the subject by the presence of the PEC j(t) in any part of the brain or heart;calculation of the TPEC by means of Bayesian Hierarchical Estimation Procedure that determines the primary electric current j(t) of the brain or heart using structural information obtained from an anatomical atlas, metabolic information obtained from functional images, and an assumption that the solution belongs to a space of Besov space of given smoothness or determined by a norm in a Megadictionary;calculation of descriptive parameters either on the basis of the observations o(t) directly, or of the j(t);calculation of a probability that all or some of the DP are similar to those of a given test group; andperforming the step wherein the obtained TPECc is coded by means of a pseudocolor scale and is overlaid on the anatomical atlas, a scale being fitted according to the probability previously calculated, whereby only statistically significant anatomical sites are highlighted. 2. The method of claim 1, wherein vectors of descriptive parameters (DP) are calculated that are statistics of models of stochastic processes for o(t) or j(t), wherein instances, of DP include at least one of: subsets of o(t) or j(t);subsets of o(t) or J(t) or μ(t), or γ(t);subsets of averages, variances and covariances or higher order o(t) or J(t) or μ(t), or γ(t) or of their representation in the basis of a Megadictionary;subsets of cross polyspectra of o(t) or J(t) or μ(t), or γ(t) of their representation in the basis of a Megadictionary;subsets of parametric multivariate Autoregressive functions fitted to o(t) or J(t) or μ(t), or γ(t) or their representation in the basis of a Megadictionary;subsets of non parametric multivariate Autoregressive functions fitted to o(t) or J(t) or μ(t), or γ(t) of their representation in the basis of a Megadictionary;correction measures, mutual information, Granger causality among components of o(t) or J(t) or μ(t), or γ(t) of their representation in the basis of a Megadictionary; andregression coefficients of any one of the previously described types of DP with stimuli emitted by the system or with the subject's responses. 3. The method of claim 1, wherein the probability is calculated that all or part of the TPEC bel... | Centro Nacional de Investigaciones Cientificas (CNIC) | A61B005/00
US07197357 | B2 | 2007-03-27 | Wireless ECG system | 2001-11-30 | A cardiac monitoring system and, more particularly, a wireless electrocardiograph (ECG) system. The present invention detects electrical signals from a patient's heart and transmits the signals digitally to a remote base station via telemetry. The base station converts the digital signals back to an analog electrical signals that can be read by an ECG monitor. | 1. A system for converting a conventional, non-wireless electrocardiograph monitoring system to a wireless electrocardiograph monitoring system comprising: a body electronics unit for acquiring electrical signals from a chest or precordial assembly and wirelessly transmitting the electrical signals to a base station, the base station a having a plurality of terminals for directly transmitting the electrical signals to any conventional electrocardiograph monitor, the base station having a user interface for communicating information to the user and further comprising an apparatus for pairing the base station with the body electronics unit wherein the apparatus is a removable token key being physically separable from said body electronics unit and said base station. | 1. A system for converting a conventional, non-wireless electrocardiograph monitoring system to a wireless electrocardiograph monitoring system comprising: a body electronics unit for acquiring electrical signals from a chest or precordial assembly and wirelessly transmitting the electrical signals to a base station, the base station a having a plurality of terminals for directly transmitting the electrical signals to any conventional electrocardiograph monitor, the base station having a user interface for communicating information to the user and further comprising an apparatus for pairing the base station with the body electronics unit wherein the apparatus is a removable token key being physically separable from said body electronics unit and said base station. 2. The system of claim 1 wherein the base station controls the data collected by the chest assembly. 3. The system of claim 1 wherein the chest assembly has five electrode connectors for connecting to electrodes. 4. The system of claim 3 wherein an electrode is positioned on the right side of a patient's chest about level of the first and second intercostal space, an electrode is positioned on the left side of the patient's chest about level of the first and second intercostals space, an electrode is positioned in the middle of the patient's chest about level of the fourth and fifth intercostals space, and two electrodes are positioned on the left side of the patient's torso. 5. The system of claim 1 wherein the on body electronics unit and the base station each have a battery port for removably retaining a battery. 6. The system of claim 5 wherein the battery is compatible with the battery port of the body electronics unit and the battery port of the base station. 7. A method of pairing the body electronics unit of claim 1 with the base station by: inserting the token key into a token key port of the base station to record an identification number of the base station; removing the token key from the token key port of the base station; inserting the token key into the token key port of the body electronics unit to record an identification number of the body electronics unit and to transmit the identification number of the base station to the body electronics unit; removing the token key port from the token key port of the body electronics unit; and inserting the token key into the token key port of the base station to transfer the identification number of the body electronics unit to the base station. 8. A method of pairing the body electronics unit of claim 1 with the base station by inserting the token key into the token key port of the body electronics unit to record an identification number or the body electronics unit; removing the token key from the token key port of the body electronics unit; inserting the token key into a token key port of the base station to record an identification number of the base station and to transmit the identification number of the body electr... | Life Sync Corporation | A61B005/04
US07175595 | B1 | 2007-02-13 | Method for adjusting of controlling the diet and/or a person's consumption | 2000-06-29 | The invention relates to a method for adjusting or controlling the diet and/or consumption of carbohydrates and/or fats and/or proteins of a person who is subjected to a certain physical strain taking into consideration his/her individual performance potential to be determined. | 1. Method for adjusting nutrition in a person subjected to physical stress, comprising the steps of: determining performance capacity of the person by determining individual anaerobic threshold of the person and measuring lactate accumulation rate ΔA at and above the individual anaerobic threshold;determining a stress state of the person in relation to the measured individual anaerobic threshold and measured lactate accumulation rate ΔA; andregulating at least one of fat, protein and carbohydrate consumption of the person as a function of the determined stress state, individual anaerobic threshold and lactate accumulation rate,wherein the step of determining the lactate accumulation rate ΔA comprises the steps of: measuring time-dependent lactate concentration change beyond the individual anaerobic threshold,plotting a measurement curve of measured lactate concentration in relation to time,determining a first gradient in the measurement curve at a time tIAT that corresponds to the individual anaerobic threshold,determining at least one second gradient in the measurement curve at a time tx with tx>tIAT; andsubtracting the second gradient from the first gradient to determine a difference, which represents the lactate accumulation rate ΔA. | 1. Method for adjusting nutrition in a person subjected to physical stress, comprising the steps of: determining performance capacity of the person by determining individual anaerobic threshold of the person and measuring lactate accumulation rate ΔA at and above the individual anaerobic threshold;determining a stress state of the person in relation to the measured individual anaerobic threshold and measured lactate accumulation rate ΔA; andregulating at least one of fat, protein and carbohydrate consumption of the person as a function of the determined stress state, individual anaerobic threshold and lactate accumulation rate,wherein the step of determining the lactate accumulation rate ΔA comprises the steps of: measuring time-dependent lactate concentration change beyond the individual anaerobic threshold,plotting a measurement curve of measured lactate concentration in relation to time,determining a first gradient in the measurement curve at a time tIAT that corresponds to the individual anaerobic threshold,determining at least one second gradient in the measurement curve at a time tx with tx>tIAT; andsubtracting the second gradient from the first gradient to determine a difference, which represents the lactate accumulation rate ΔA. 2. Method according to claim 1, wherein the stress occurs in a person over an extended period of time below the determined individual anaerobic threshold, and the fat and the carbohydrate percentage of the nutrition are adjusted comparatively higher than the protein percentage. 3. Method according to claim 1, wherein the stress occurs with a lactate accumulation rate ΔA approaching ΔAmax, and the protein percentage of the nutrition is adjusted up to several times as with stress at ΔA=0. 4. Method according to claim 1, wherein the performance capacity is determined under a stress selected from the group consisting of a running test, a swimming test, a stepping test and ergometry with graduated or continuous stress increase with and without breaks. | - | A61B005/00
US07127280 | B2 | 2006-10-24 | Apparatus and process for examining a surface | 2001-06-22 | An apparatus designed to examine a surface includes a polarization analyser element, or analyser, placed in the path of a light beam reflected by the surface, a device configured to take digital images and placed in the path of the beam reflected by the surface downstream of the analyser, and a processing unit capable of calculating the color and the intensity of a plurality of pixels of at least one image. | 1. An apparatus for examining a surface, comprising: a polarization analyser element placed in a path of a light beam reflected by the surface, the polarization analyser element constructed and arranged to alternately transmit a crossed polarization state and a parallel polarization state;a digital image acquisition device disposed in the path of the beam reflected by the surface downstream of the polarization analyser element; anda processing unit configured to calculate a brightness and a color of a plurality of points of the surface from pixels of at least two images of the surface;wherein the apparatus is constructed and arranged to examine the surface without contacting the surface. | 1. An apparatus for examining a surface, comprising: a polarization analyser element placed in a path of a light beam reflected by the surface, the polarization analyser element constructed and arranged to alternately transmit a crossed polarization state and a parallel polarization state;a digital image acquisition device disposed in the path of the beam reflected by the surface downstream of the polarization analyser element; anda processing unit configured to calculate a brightness and a color of a plurality of points of the surface from pixels of at least two images of the surface;wherein the apparatus is constructed and arranged to examine the surface without contacting the surface. 2. An apparatus according to claim 1, further comprising a source of polarized light configured to emit a beam incident on the surface to be examined. 3. An apparatus according to claim 2, wherein the light emanating from the source is substantially isotropic. 4. An apparatus according to claim 2, wherein the light emanating from the source is substantially white. 5. An apparatus according to claim 2, wherein a spectrum of the light emanating from the source is substantially the same as a solar spectrum. 6. An apparatus according to claim 1, wherein the polarization analyser element comprises a first transmitter configured to transmit crossed polarization and a second transmitter configured to transmit parallel polarization, the first and second transmitters being alternatively active. 7. An apparatus according to claim 6, wherein the polarization analyser element is rotatable. 8. An apparatus according to claim 6, the polarization analyser element further comprises an electrical switching component. 9. A process for a non-contact examination of a keratinous surface, comprising: (i) analysing crossed and parallel polarizations of a light beam reflected by the surface;(ii) taking digital images of the crossed and parallel polarizations of the reflected beam; and(iii) calculating a brightness and a color of a plurality of points of the surface from pixels of at least two images of the surface. 10. A process according to claim 9, wherein the surface is uneven. 11. A process according to claim 9, wherein the digital images are monochromatic digital images. 12. A process according to claim 9, wherein the digital images are polychromatic digital images. 13. The process of claim 9, wherein the process is performed by a computer. 14. A computer-readable medium bearing a program code embodied thereon for performing the process of claim 9. 15. An apparatus for examining a surface comprising: a source of polarized light constructed and arranged to emit a beam incident on the surface to be examined, a spectrum of the light being substantially the same as a solar spectrum;a polarization analyzer element placed in a path of a light beam reflected by the surface;a digital image acquisition device disposed in the path of the beam reflected by the surface downstream... | L'Oreal | A61B005/00
US07025949 | B2 | 2006-04-11 | In-vivo diagnostic method by means of near infrared radiation | 2001-05-07 | This invention relates to an in-vivo diagnostic method based on near infrared radiation (NIR radiation) that uses water-soluble dyes and their biomolecule adducts, each having specific photophysical and pharmaco-chemical properties, as a contrast medium for fluorescence and transillumination diagnostics in the NIR range, to new dyes and pharmaceuticals containing such dyes. | 1. A cyanine dye as represented by the general formula IIa: wherein r represents the numbers 0, 1 or 2, wherein, for r=2, the respective L6 and L7 groups that occur in duplicate may be the same or different, L1 to L7 are the same or different, each independently representing a CH or CR group,where R is a halogen atom, a hydroxy, carboxy acetoxy, amino, nitro, cyano or sulfonic acid group or an alkyl, alkenyl, hydroxyalkyl, carboxyalkyl, alkoxy, alkoxycarbonyl, sulfoalkyl, alkylamino, dialkylamino or halogenalkyl group containing up to 6 carbon atoms, aryl, alkylaryl, hydroxyaryl, carboxyaryl, sulfoaryl, arylamino, diarylamino, nitroaryl or halogenaryl group containing up to 9 carbon atoms, or where R represents a bond that bonds to another group R and forms a 4- to 6-member ring together with the interspersed groups L1 to L7, or two of L1 and L7 that are linked via a —CO— group,R3 to R12 are the same or different, each independently representing a hydrogen atom, a group B or W, or an alkyl or alkenyl group containing up to 6 carbon atoms or an aryl or aralkyl group optionally carrying an additional group W, or to each pair of adjacent groups R3 to R10 are annealed, with due regard for the interspersed C atoms, 5- to 6-member rings that may be saturated, unsaturated or aromatic, and that optionally carry an additional group R, and X and Y are the same or different, each independently representing an O, Se or Te or a —C(CH3)2—, —CH═CH— or —CR13R14— group, wherein at least one of R3 to R12 comprises a group W,where R13 and R14 independently represent a hydrogen atom, a group B or W, or an alkyl or alkylene group containing up to 6 carbon atoms or an aryl or aralkyl group containing up to 9 carbon atoms, the alkyl, alkenyl, aryl or aralkyl group optionally carrying an additional group W,wherein at least one of R3 to R14 is a group B,wherein B is a biological detecting unit having a molecular weight of up to 30,000 that bonds to specific cell populations or selectively to receptors, or accumulates in tissues or tumors, and W is a hydrophilic group that improves water-solubility selected from the group consisting of(i) a carboxylic acid group;(ii) a sulfonic acid group;(iii) a carboxyalkyl group;(iv) an alkoxycarbonyl or alkoxyoxoalkyl group containing up to 12 carbon atoms; and(v) a group of the general formula III —(CH2)n-O—Z or —(CH2—CH2—O)a-Z (III)whereina represents the numbers 0 to 6,Z represents a hydrogen atom or an alkyl group containing 3 to 6 C atoms having 2 to n−1 hydroxy groups, with n being the number of C atoms or an aryl or aralkyl group containing 6 to 10 C atoms and having 2 to 4 additional hydroxy groups, or an alkyl group containing 1 to 6 C atoms and having 1 to 3 additional carboxy groups, or an aryl group containing 6 to 9 C atoms and having 1 to 3 additional carboxyl groups or an aralkyl group or a nitroaryl or a nitroaralkyl group containing 6 to 15 C atoms;(vi) a group of the general formulas IIIa or IIIb; (vii) a group of t... | 1. A cyanine dye as represented by the general formula IIa: wherein r represents the numbers 0, 1 or 2, wherein, for r=2, the respective L6 and L7 groups that occur in duplicate may be the same or different, L1 to L7 are the same or different, each independently representing a CH or CR group,where R is a halogen atom, a hydroxy, carboxy acetoxy, amino, nitro, cyano or sulfonic acid group or an alkyl, alkenyl, hydroxyalkyl, carboxyalkyl, alkoxy, alkoxycarbonyl, sulfoalkyl, alkylamino, dialkylamino or halogenalkyl group containing up to 6 carbon atoms, aryl, alkylaryl, hydroxyaryl, carboxyaryl, sulfoaryl, arylamino, diarylamino, nitroaryl or halogenaryl group containing up to 9 carbon atoms, or where R represents a bond that bonds to another group R and forms a 4- to 6-member ring together with the interspersed groups L1 to L7, or two of L1 and L7 that are linked via a —CO— group,R3 to R12 are the same or different, each independently representing a hydrogen atom, a group B or W, or an alkyl or alkenyl group containing up to 6 carbon atoms or an aryl or aralkyl group optionally carrying an additional group W, or to each pair of adjacent groups R3 to R10 are annealed, with due regard for the interspersed C atoms, 5- to 6-member rings that may be saturated, unsaturated or aromatic, and that optionally carry an additional group R, and X and Y are the same or different, each independently representing an O, Se or Te or a —C(CH3)2—, —CH═CH— or —CR13R14— group, wherein at least one of R3 to R12 comprises a group W,where R13 and R14 independently represent a hydrogen atom, a group B or W, or an alkyl or alkylene group containing up to 6 carbon atoms or an aryl or aralkyl group containing up to 9 carbon atoms, the alkyl, alkenyl, aryl or aralkyl group optionally carrying an additional group W,wherein at least one of R3 to R14 is a group B,wherein B is a biological detecting unit having a molecular weight of up to 30,000 that bonds to specific cell populations or selectively to receptors, or accumulates in tissues or tumors, and W is a hydrophilic group that improves water-solubility selected from the group consisting of(i) a carboxylic acid group;(ii) a sulfonic acid group;(iii) a carboxyalkyl group;(iv) an alkoxycarbonyl or alkoxyoxoalkyl group containing up to 12 carbon atoms; and(v) a group of the general formula III —(CH2)n-O—Z or —(CH2—CH2—O)a-Z (III)whereina represents the numbers 0 to 6,Z represents a hydrogen atom or an alkyl group containing 3 to 6 C atoms having 2 to n−1 hydroxy groups, with n being the number of C atoms or an aryl or aralkyl group containing 6 to 10 C atoms and having 2 to 4 additional hydroxy groups, or an alkyl group containing 1 to 6 C atoms and having 1 to 3 additional carboxy groups, or an aryl group containing 6 to 9 C atoms and having 1 to 3 additional carboxyl groups or an aralkyl group or a nitroaryl or a nitroaralkyl group containing 6 to 15 C atoms;(vi) a group of the general formulas IIIa or IIIb; (vii) a group of t... | Institut fur Diagnostikforschung GmbH an der Freien Universitat Berlin | A61B010/00 A61B005/00 A61B008/00
US07074183 | B2 | 2006-07-11 | Method and system for improving vascular systems in humans using biofeedback and network data communication | 2002-06-05 | Methods for treating vascular disease in humans and generating data representing treatment plans are disclosed. A first set of clinical vascular health data from a healthcare provider and representing a vascular health condition of a patient is received at a data center server that is communicatively coupled to a public data network. One or more vascular disease analysis algorithms are applied to the first set of vascular health data, to result in creating and storing an initial treatment plan for the patient. A second set of vascular health data is received from a monitoring device that is associated with the patient and that is communicatively coupled to the data network; the second set of data include Doppler monitor data obtained from the peripheral vascular system of the patient. One or more vascular analysis algorithms are applied to result in creating one or more supplementary treatment plans for the patient. At least one of the treatment plans includes a biofeedback interaction. The treatment plans are provided to the patient over the data network. The foregoing steps are iteratively repeated one or more times as determined by the physician and patient, resulting in improved vascular health. | 1. A method of treating vascular disease in humans, the method comprising the computer-implemented steps of: (a) receiving, at a data center server that is communicatively coupled to a public data network, a first set of clinical vascular health data from a healthcare provider and representing a vascular health condition of a patient;(b) applying one or more vascular disease analysis algorithms to the first set of vascular health data, to result in creating and storing an initial treatment plan for the patient;(c) receiving a second set of vascular health data from a monitoring device that is associated with the patient and that is communicatively coupled to the data network, wherein the second set of vascular health data includes data from a Doppler vascular monitor;(d) applying one or more vascular analysis algorithms to result in creating one or more supplementary treatment plans for the patient, wherein at least one of the treatment plans includes a biofeedback interaction component;(e) providing the treatment plans to the patient over the data network;(f) iteratively repeating steps (c), (d) and (e) one or more times as determined by the patient, wherein one or more iterations include receiving patient data created after and resulting from the patient performing a biofeedback interaction and modifying the treatment plan based upon the resulting patient data, wherein the biofeedback interaction comprises the patient performing an action identified in the treatment plan while simultaneously (g) self-monitoring using the monitoring device and (h) observing output from the monitoring device, and the monitoring device generating the patient data during the performing of the action. | 1. A method of treating vascular disease in humans, the method comprising the computer-implemented steps of: (a) receiving, at a data center server that is communicatively coupled to a public data network, a first set of clinical vascular health data from a healthcare provider and representing a vascular health condition of a patient;(b) applying one or more vascular disease analysis algorithms to the first set of vascular health data, to result in creating and storing an initial treatment plan for the patient;(c) receiving a second set of vascular health data from a monitoring device that is associated with the patient and that is communicatively coupled to the data network, wherein the second set of vascular health data includes data from a Doppler vascular monitor;(d) applying one or more vascular analysis algorithms to result in creating one or more supplementary treatment plans for the patient, wherein at least one of the treatment plans includes a biofeedback interaction component;(e) providing the treatment plans to the patient over the data network;(f) iteratively repeating steps (c), (d) and (e) one or more times as determined by the patient, wherein one or more iterations include receiving patient data created after and resulting from the patient performing a biofeedback interaction and modifying the treatment plan based upon the resulting patient data, wherein the biofeedback interaction comprises the patient performing an action identified in the treatment plan while simultaneously (g) self-monitoring using the monitoring device and (h) observing output from the monitoring device, and the monitoring device generating the patient data during the performing of the action. 2. A method as recited in claim 1, wherein the initial treatment plan includes a biofeedback interaction component. 3. A method as recited in claim 1, wherein the monitoring device includes a Doppler vascular monitor that monitors vascular functions of the patient by attachment to locations proximate to a peripheral vascular system of the patient. 4. A method as recited in claim 1, further comprising the step of receiving data representing a selection, by the healthcare provider, of one or more parameter values that modify operation of one or more of the vascular disease analysis algorithms. 5. A method as recited in claim 1, wherein the initial treatment plan or the subsequent treatment plans comprise performing a biofeedback interaction using a client system of the patient. 6. A method as recited in claim 1, wherein step (c) further comprises displaying a multimedia vascular disease educational program to the patient. 7. A method as recited in claim 1, wherein step (c) further comprises the steps of generating and displaying, on the patient client system, a graphical representation of a vascular system anatomical element that is animated based on real-time data that is received from the patient using the monitoring device. 8. A method as recited in claim 1, wh... | Alexander F. Castellanos | A61B005/00
US07069069 | B2 | 2006-06-27 | Implantable cardiac device for and method of monitoring progression or regression of heart disease by quantifying morphological features | 2001-12-06 | An implantable cardiac device and method provides for monitoring a progression or regression in heart disease during an extended time period. A sensor generates an electrogram signal representing electrical activity of a patient's heart. From the generated electrogram signal, a processor determines morphology measurements wherein the morphology measurements indicate a progression or regression in the heart disease. The morphology measurements are stored in a memory during the extended period of time. A telemetry circuit transmits the stored morphology measurements to an external receiver for retrieval or display. The morphology measurements may be of electrogram intrinsic activity or evoked response activity characteristics. | 1. An implantable medical device for monitoring changes in heart disease, comprising: a sensing circuit that senses an intracardiac electrogram;a signal processing circuit, coupled to the sensing circuit, that determines amplitude, width and polarity from selected complexes within the intracardiac electrogram;a comparator that compares the amplitude, width and polarity of a current complex with the amplitude, width and polarity of a previous complex; anda controller that indicates relative changes between the current complex and the previous complex, wherein the relative changes correspond to changes in a progression or regression of the heart disease. | 1. An implantable medical device for monitoring changes in heart disease, comprising: a sensing circuit that senses an intracardiac electrogram;a signal processing circuit, coupled to the sensing circuit, that determines amplitude, width and polarity from selected complexes within the intracardiac electrogram;a comparator that compares the amplitude, width and polarity of a current complex with the amplitude, width and polarity of a previous complex; anda controller that indicates relative changes between the current complex and the previous complex, wherein the relative changes correspond to changes in a progression or regression of the heart disease. 2. The implantable medical device recited in claim 1, wherein: the sensing circuit senses atrial complexes within the intracardiac electrogram; andthe signal processing circuit determines amplitude, width and polarity from selected atrial complexes. 3. The implantable medical device recited in claim 1, wherein: the sensing circuit senses ventricular complexes within the intracardiac electrogram; andthe signal processing circuit determines amplitude, width and polarity from selected ventricular complexes. 4. The implantable medical device recited in claim 1, wherein: the sensing circuit senses repolarization waveforms corresponding to T-wave complexes within the intracardiac electrogram; andthe signal processing circuit determines amplitude, width and polarity from the repolarization waveforms. 5. The implantable medical device recited in claim 1, wherein: the sensing circuit senses intrinsic complexes within the intracardiac electrogram; andthe comparator compares current intrinsic waveforms with previous intrinsic waveforms. 6. The implantable medical device recited in claim 1, wherein: the sensing circuit senses evoked complexes within the intracardiac electrogram; andthe comparator compares current evoked waveforms with previous evoked waveforms. 7. The implantable medical device recited in claim 1, wherein: the controller determines a matching score based on the amplitude, width and polarity of the current complex to the previous complex, and further reports changes in the matching score, wherein the relative changes in the matching score correspond to changes in the progression or regression of the heart disease. 8. The implantable medical device recited in claim 1, wherein: the controller determines whether the current complex has exceeded an out-of-bound limit relative to the previous complex, wherein the out-of-bound limit corresponds to a change in the progression or regression of the heart disease. 9. The implantable medical device recited in claim 1, wherein: the signal processing circuit determines a template based on the amplitude, width and polarity of the selected complexes; andthe comparator compares the template of the current complex with the template of the previous complex to produce a template matching score relative to the previous complex. 10. The implantable medical... | Pacesetter, Inc. | A61B005/0472
US06856821 | B2 | 2005-02-15 | System for combined transcutaneous blood gas monitoring and vacuum assisted wound closure | 2002-02-28 | A method and apparatus for the transcutaneous monitoring of blood gases generally comprises a blood gas data acquisition device, a vacuum source and a blood gas transducer unit. The blood gas transducer unit is adapted for application to a patient's skin and administration of a local vacuum at the area of patient application. It further comprises an electrochemical blood gas transducer, well known to those of ordinary skill in the art, which is disposed entirely within the local vacuum at the area of patient application. The vacuum source is placed in fluid communication with the blood gas transducer unit, through a hydrophobic membrane filter for safety purposes, in order to induce a condition of hyperperfusion in the locality of the electrochemical blood gas transducer. Under the control of a microcontroller, or equivalent means, the blood gas acquisition device is then utilized to capture a measure of skin surface oxygen or carbon dioxide pressure. The microcontroller can then utilize this measure to arrive at an estimate of arterial partial pressure of oxygen or carbon dioxide, accordingly. Because vacuum induced perfusion produces the requisite condition of hyperperfusion without local heating and, therefore, without acceleration of the local metabolic function, the present invention results in more accurate than previously available estimates of partial pressure blood gas pressures and does so while eliminating a significant risk for injury to the patient. | 1. A transcutaneous blood gas monitoring device particularly suited for application to a site on a patient's skin under a local vacuum, comprising: a non-invasive blood gas data acquisition device; a vacuum source; an electrochemical blood gas transducer unit in fluid communication with said vacuum source; wherein said vacuum source is adapted to be in fluid communication with a wound site in proximity to said site on patient's skin for applying a negative pressure to said wound site; a screen formed to fit within said wound site, such that said screen is in fluid communication with said vacuum source; and a seal for maintaining a negative pressure within said wound site during application of said negative pressure. | 1. A transcutaneous blood gas monitoring device particularly suited for application to a site on a patient's skin under a local vacuum, comprising: a non-invasive blood gas data acquisition device; a vacuum source; an electrochemical blood gas transducer unit in fluid communication with said vacuum source; wherein said vacuum source is adapted to be in fluid communication with a wound site in proximity to said site on patient's skin for applying a negative pressure to said wound site; a screen formed to fit within said wound site, such that said screen is in fluid communication with said vacuum source; and a seal for maintaining a negative pressure within said wound site during application of said negative pressure. 2. A transcutaneous blood gas monitoring device particularly suited for application to a site on a patient's skin under a local vacuum, comprising: a non-invasive blood gas data acquisition device; a vacuum source; an electrochemical blood gas transducer unit in fluid communication with said vacuum source; wherein said vacuum source is adapted to be in fluid communication with a wound site in proximity to said site on patient's skin for applying a negative pressure to said wound site; wherein said data acquisition device monitors healing of said wound site. | KCI Licensing, Inc. | A61B005/00
US06862467 | B2 | 2005-03-01 | Imaging catheter assembly with distal end inductive coupler and embedded transmission line | 2002-08-09 | A catheter assembly includes an elongate catheter body having a proximal end and a distal end with a drive cable disposed therein, the drive cable having a proximal end and a distal end, and rotatable relative to the catheter body. A first electro-magnetic element is disposed proximate the distal end of the catheter, and a second electro-magnetic element disposed proximate the distal end of the drive cable and in electrical communication with an operative element mounted at the end of the drive cable, the first and second electro-magnetic elements forming an inductive coupler. The catheter assembly can include various other distal operative elements, which are in communication with corresponding proximal operative elements via transmission lines embedded within the wall of the catheter body. | 1. A catheter assembly, comprising: an elongate catheter body having a proximal end and a distal end; a drive cable disposed in the catheter, the drive cable having a proximal end and a distal end, and rotatable relative to the catheter body; an operative element disposed on the distal end of the drive cable, the operative element having a capacitive reactance at an operating frequency; a conductor having a proximal end and a distal end; and an inductive coupler disposed in the distal end of the catheter body and electrically coupled to the operative element and the distal end of the conductor, wherein the inductive coupler inductively couples the operative element to the conductor and the inductive coupler has an inductive reactance that is substantially equal to the capacitive reactance at the operating frequency. | 1. A catheter assembly, comprising: an elongate catheter body having a proximal end and a distal end; a drive cable disposed in the catheter, the drive cable having a proximal end and a distal end, and rotatable relative to the catheter body; an operative element disposed on the distal end of the drive cable, the operative element having a capacitive reactance at an operating frequency; a conductor having a proximal end and a distal end; and an inductive coupler disposed in the distal end of the catheter body and electrically coupled to the operative element and the distal end of the conductor, wherein the inductive coupler inductively couples the operative element to the conductor and the inductive coupler has an inductive reactance that is substantially equal to the capacitive reactance at the operating frequency. 2. The catheter assembly of claim 1, wherein the operative element comprises an ultrasonic transducer. 3. The catheter assembly of claim 1, wherein the conductor is disposed within the catheter body. 4. The catheter assembly of claim 1, wherein the inductive coupler comprises a stator fixably disposed in the catheter body and electrically coupled to the distal end of the conductor, the stator comprising a generally hollow cylinder; and a rotor mechanically coupled to the distal end of the drive cable and electrically coupled to the operative element, the rotor comprising a rod rotatably disposed in the hollow cylinder. 5. The catheter assembly of claim 4, wherein the stator further comprises a first electrically conductive coil disposed on the inner surface of the hollow cylinder, and the rotor further comprises a second electrically conductive coil disposed on the outer surface of the rod. 6. The catheter assembly of claim 5, wherein the stator and rotor having opposing surface areas, wherein the respective stator and rotor surface areas, along with the respective diameter, size and number of turns of the first and second electrically conductive coils, are selected such that the value of the inductive reactance of the inductive coupler is substantially equal to the capacitive reactance at the operating frequency. 7. The catheter assembly of claim 4, wherein the stator has a first electrically conductive coil; the rotator has a second electrically conductive coil, and the ratio of turns between the first and second electrically conductive coils is selected such that the input impedance looking into the inductive coupler from the conductor substantially matches the characteristic impedance of the conductor. 8. A catheter assembly, comprising: an elongate catheter body having a proximal end and a distal end; a drive cable disposed in the catheter, the drive cable having a proximal end and a distal end, and rotatable relative to the catheter body; first and second operative elements disposed on the distal end of the drive cable; first and second conductors, each conductor having a proximal end and a distal end; a first inductiv... | SciMed Life Systems, Inc. | A61B005/05 A61B008/12
US06942614 | B1 | 2005-09-13 | Method and system for use in treating a patient with an anticoagulant to optimize therapy and prevent an adverse drug response | 2000-08-24 | A method and system for use in treating a patient receiving an anticoagulant or a substance containing warfarin to optimize therapy and prevent an adverse drug response. This system employs surrogate markers or indicators including blood levels of the anticoagulant to determine the next required dose for a patient. Since the surrogate markers are employed as a percent change in status, virtually any indicator can be used. Surrogate markers could include any measure of the effectiveness of the anticoagulant's action. Given the effectiveness of the anticoagulant's action relative to the surrogate markers, a change in anticoagulant dose is calculated by the system. Conversely, by employing this system, one could determine the expected result of the anticoagulant dose change on the surrogate markers. | 1. A method for calculating a revised dose of an anticoagulant for a patient using said anticoagulant, comprising the steps of: accepting as a first input the patient's current anticoagulant dose; accepting as a second input a maximum dose of the anticoagulant; accepting as a third input a percent response of the patient based on one or more surrogate markers for said patient; and determining a revised dose, wherein said revised dose is a function of said current dose minus a ratio of the percent response of the patient and a ratio of said current dose to said maximum dose plus the percent of individual patient response multiplied by a response factor. | 1. A method for calculating a revised dose of an anticoagulant for a patient using said anticoagulant, comprising the steps of: accepting as a first input the patient's current anticoagulant dose; accepting as a second input a maximum dose of the anticoagulant; accepting as a third input a percent response of the patient based on one or more surrogate markers for said patient; and determining a revised dose, wherein said revised dose is a function of said current dose minus a ratio of the percent response of the patient and a ratio of said current dose to said maximum dose plus the percent of individual patient response multiplied by a response factor. 2. The method of claim 1, wherein: said determining step includes determining said revised dose based on the equation RAD=CAD−{[&lt;(PAR−100)/PAR>/&lt;1+(CAD/HIGH)>]×CAD}+LV where: LV={(RESPONSE×CAD)×[(100−RES)×0.01]}/1.3^(CAD/HIGH) and wherein: RAD=Revised Anticoagulant Dose CAD=Current Anticoagulant Dose PAR=Percent response of patient to surrogate marker RES=Percent response of patient to last dosing based on surrogate marker HIGH=The input parameter that is the high dose range for said anticoagulant RESPONSE=Percent of total dose available for individualizing patient dose 1.3^(CAD/HIGH)=1.3 raised to an exponent of (CAD/HIGH). 3. The method of claim 1, wherein: said anticoagulant is selected from a group comprising warfarin, Coumadin®, heparin, warfarin sodium salt, coumarin derivatives, indandione derivatives, dicumarol, anisindione, phenindione, ethyl bicoumacetate, bishydroxycoumarin, abcimixab, Reopro®, actilyse, alteplase, Activase®, anagrelide, Agrylin®, anistreplase, Eminase®, antithrombin III, Thrombate III®, ardeparin, Normiflo®, argatroban, clopidrogel, Plavix®, dalteparin, Fragmin®, danaparoid, Orgaran®, dipyridamole, Persantine®, dipyridamole/aspirin, Aggrenox®, duteplase, enoxaparin, Lovenox®, eptifibatide, Integrilin®, lepirudin, Refludan®, nadroparin, Fraxiparine®, oprelvekin, Neumega®, pentosan polysulfate sodium, Elmiron®, reteplase, Retavase®, reviparin, Clivarine®, saruplase, streptokinase, Kabikinase®, Streptase®, tinzaparin, Innohep®, tirofiban, Aggrastat®, unfractionated heparin, low molecular weight heparin, all antithrombotic agents, all vitamin K antagonists, and all substances derived from and/or related to the foregoing substances. 4. A method for calculating a revised dose of a anticoagulant for a patient using said anticoagulant comprising the steps of: accepting as a first input the patient's current anticoagulant dose; accepting as a second input a maximum dose of the anticoagulant; accepting as a third input one or more numerical markers indicating a response of the patient; and calculating said revised dose, wherein said revised dose is a function of said current dose minus the ratio of the change in numerical markers and the ratio of said current dose to said maximum dose plus the percent of individual patient response multiplied by a response factor. 5... | Dimensional Dosing Systems, Inc. | A61B005/00 A61B019/00 G06F017/00
US06948365 | B2 | 2005-09-27 | Dynamometer and related assessment method | 2002-09-06 | Dynamometer (1) and related method for assessing the force (F) exerted at hand level, said dynamometer comprising a fixed frame (2), a pair of levers (3, 4) apt to be taken hold of by a user, at least one lever being movable, a contrast element having a known resistance parameter, connected to the movable lever in order to contrast the displacement thereof, and displacement measuring means (10), associated to the movable lever in order to measure the displacement (θ) thereof. | 1. A method for measuring a force (F) exerted at hand level, comprising the steps of: having a user grasp, with one hand, a pair of levers (3, 4) located on a frame (2), at least one (4) thereof being movable; having the user exert force (F) against a resisting action of a contrast element (8) connected to the at least one movable lever and having a known resistance parameter (k); measuring a displacement (θ) given by the user to the at least one movable lever; and processing the measuring carried out in such a way to compute the force (F) exerted according to the measured displacement and to the resistance parameter of the contrast element; wherein said step of measuring the displacement given by the user to the at least one movable lever (4) provides the measuring of an angular displacement (θ) of said lever. | 1. A method for measuring a force (F) exerted at hand level, comprising the steps of: having a user grasp, with one hand, a pair of levers (3, 4) located on a frame (2), at least one (4) thereof being movable; having the user exert force (F) against a resisting action of a contrast element (8) connected to the at least one movable lever and having a known resistance parameter (k); measuring a displacement (θ) given by the user to the at least one movable lever; and processing the measuring carried out in such a way to compute the force (F) exerted according to the measured displacement and to the resistance parameter of the contrast element; wherein said step of measuring the displacement given by the user to the at least one movable lever (4) provides the measuring of an angular displacement (θ) of said lever. 2. The method according to claim 1, comprising the further step of computing the work (F′) performed by the user upon exerting the force (F) to be measured. 3. The method according to claim 1, wherein said step of having a user grasp a pair of levers (3, 4) provides that the thumb and the forefinger of a user's hand each grasp a respective lever of the pair of levers, thereby enabling the assessment of the force of the flexors of such fingers. 4. The method according to claim 1, wherein said step of having a user grasp a pair of levers (3, 4) provides that the thumb grasp the fixed lever (3) and that the forefinger and the middle finger grasp the movable lever (4). 5. The method according to claim 1, wherein said resisting action is an elastic return force. 6. A dynamometer (1) for assessing the force (F) exerted at hand level, comprising: a fixed frame (2); a pair of levers (3, 4) apt to be grasped by a user, at least one (4) of which is movable with respect to said frame; a contrast element (8) having a known resistance parameter (k), connected to said at least one movable lever to contrast a displacement (θ) thereof; displacement measuring means (10), associated to said at least one movable lever to measure the displacement (θ) thereof following operation by the user; and means (11) for processing the measuring carried out; wherein the displacement measuring means (10) comprises means for measuring an angular displacement of the at least one movable lever at a rotatable connection between the at least one movable lever and the frame. 7. The dynamometer (1) according to claim 6, wherein at least one lever (3, 4) of said pair has an anatomical handgrip (7). 8. The dynamometer (1) according to claim 7, wherein said anatomical handgrip comprises three bays (7), each apt to receive a finger. 9. The dynamometer (1) according to claim 7, wherein a first lever (4) of said pair has a bay (7) specifically shaped to receive a forefinger and a second lever (3) of said pair has a bay (7) specifically shaped to receive a thumb. 10. The dynamometer (1) according to claim 9, wherein said first lever (4) is movable with respect to said frame ... | K.S. Italia s.a.s. di Ambrosone Mario & Co. | A61B005/22
US06868285 | B2 | 2005-03-15 | Method and device for detecting substances in body fluids by Raman spectroscopy | 2001-05-31 | The invention relates to a method and to a device for the non-invasive detection or determination, by Raman spectroscopy, of the concentration of substances in body fluids. The aim of the invention is to facilitate a non-invasive in vivo detection of substances in body fluids that allows for very accurate, reproducible results of analysis while requiring only little time for measuring which is acceptable for the patient. To this end, the inventive method and device for carrying out a Raman spectroscopy of a body fluid in a body tissue (1) records at least two Raman spectrums under different physical conditions and compares them with each other. The result of comparison allows detection of the substance one is looking for and measurement of the concentration of said substance. | 1. A method for the non-invasive determination of a concentration of a substance in a body fluid in a body tissue by Raman spectroscopy comprising the steps: a) irradiating monochromatic primary light of a first wavelength into the body tissue, b) recording secondary light that is back scattered by the body tissue and storing the intensity of the secondary light in a first memory according to its wavelength as a first Raman spectral signal for the first primary wavelength, c) irradiating monochromatic primary light of a second wavelength into the body tissue, d) recording secondary light that is back scattered by the body tissue and storing the intensity of the secondary light in a second memory according to its wavelength as a second Raman spectral signal for the second primary wavelength, e) comparing the first Raman spectral signal in the first memory with the second Raman spectral signal in the second memory and generating a compare signal in which spectral structures caused by the body tissue are substantially eliminated, f) determining an intensity of the compare signal at at least one sample wavelength for which the intensity corresponds to the concentration of the substance in the body fluid, wherein the primary light of the first and second wavelength is irradiated with a preset irradiation frequency and each Raman spectral signal is recorded with said irradiation frequency. | 1. A method for the non-invasive determination of a concentration of a substance in a body fluid in a body tissue by Raman spectroscopy comprising the steps: a) irradiating monochromatic primary light of a first wavelength into the body tissue, b) recording secondary light that is back scattered by the body tissue and storing the intensity of the secondary light in a first memory according to its wavelength as a first Raman spectral signal for the first primary wavelength, c) irradiating monochromatic primary light of a second wavelength into the body tissue, d) recording secondary light that is back scattered by the body tissue and storing the intensity of the secondary light in a second memory according to its wavelength as a second Raman spectral signal for the second primary wavelength, e) comparing the first Raman spectral signal in the first memory with the second Raman spectral signal in the second memory and generating a compare signal in which spectral structures caused by the body tissue are substantially eliminated, f) determining an intensity of the compare signal at at least one sample wavelength for which the intensity corresponds to the concentration of the substance in the body fluid, wherein the primary light of the first and second wavelength is irradiated with a preset irradiation frequency and each Raman spectral signal is recorded with said irradiation frequency. 2. The method according to claim 1, wherein the secondary light is recorded and the intensity of the compare signal is determined in the Stokes and anti-Stokes region of the Raman spectrum. 3. The method according to claim 2, wherein an excitation light is irradiated into the body tissue which is absorbed by the substance. 4. The method according to claim 3, wherein the excitation light is pulsed with a pulse length in the picosecond region. 5. The method according to claim 4, wherein the primary light is pulsed with a pulse length in the picosecond region. 6. A method for non-invasive determination of a concentration of a substance in body tissue by Raman spectroscopy comprising the steps: a) irradiating monochromatic primary light of a first wavelength into the body tissue at a first temperature of the body tissue, b) recording secondary light that is back scattered by the body tissue and storing the intensity of the secondary light in a first memory according to its wavelength as a first Raman spectral signal for the first wavelength and for the first temperature, c) irradiating monochromatic primary light of a second wavelength into the body tissue at a first temperature of the body tissue, d) recording secondary light that is back scattered by the body tissue and storing the intensity of the secondary light in a second memory according to its wavelength as a second Raman spectral signal for the second wavelength and the first temperature, e) comparing the first Raman spectral signal in the first memory with the second Raman spectral signal in the second me... | - | A61B005/00
US06895266 | B1 | 2005-05-17 | Laser light emitter surgical site locating device and method | 2001-03-26 | An apparatus and method for guiding a surgical instrument within and around the human body. The invention employs a light emitter and an array of light sensors and a display to accurately locate an anatomical feature into which the light emitter has been placed. Laser or LED light emitted from an emitter, conveyed by a light conduit, is detected after passing through a thickness of bone and/or soft tissue by the sensor array. The signal from the sensor array is processed-and information indicating the relative direction of the emitter from the sensor is displayed. The light can be directed by a mirror or by bending the end of the light conduit. | 1. A guidance system for locating points in a living body having body passages therein comprising: a. a source of laser light for emitting light of a predetermined frequency and means for directing said laser light in a selected body passage and in a selected direction in said living body, b. a detector having a plurality of individual sensors mounted in a sensor array such that laser light from said source passing through tissue in said living body impinges on one or more of said sensor array and produces an electrical signal proportional to the laser light impinging thereon, c. a processor for processing intensity of laser light from said laser light source passing through said body tissue and falling on each sensor, respectively, in said sensor array, and d. an indicator display connected to said processor for indicating the relative intensity of laser light impinging on said sensors, respectively, as a function of the location of said laser light source in said living body. | 1. A guidance system for locating points in a living body having body passages therein comprising: a. a source of laser light for emitting light of a predetermined frequency and means for directing said laser light in a selected body passage and in a selected direction in said living body, b. a detector having a plurality of individual sensors mounted in a sensor array such that laser light from said source passing through tissue in said living body impinges on one or more of said sensor array and produces an electrical signal proportional to the laser light impinging thereon, c. a processor for processing intensity of laser light from said laser light source passing through said body tissue and falling on each sensor, respectively, in said sensor array, and d. an indicator display connected to said processor for indicating the relative intensity of laser light impinging on said sensors, respectively, as a function of the location of said laser light source in said living body. 2. The guidance system defined in claim 1 including a mirror for directing laser light in said selected direction. 3. The guidance system defined in claim 1 including a hollow laser light opaque rod having one or more apertures therein for passing laser light in said selected direction to said sensors. 4. A guidance system for locating a point in a living body having a passage therein as defined in claim 1, said source of laser light of a predetermined frequency and a light conduit having a light-receiving end and a light-emitting end means for introducing said light-emitting end into a selected passage in said living body. 5. The guidance system defined in claim 4 wherein said light-emitting end is bent and light is emitted from said bent end. 6. The guidance system defined in claim 4 including a mirror. 7. A surgical guidance method for locating a point in a living body having passages therein comprising the steps of: a. introducing a laser light source into a selected passage in said living body, b. providing a detector having a plurality of individual laser light intensity sensors mounted in a predetermined pattern and detecting laser light from said emitter passing through tissue in said living body which impinges on one or more sensors in said sensors in said predetermined pattern, and c. displaying the relative light intensity of laser light detected by each sensor in said predetermined pattern in such a manner to provide an accurate indication of the relative direction of movement to achieve alignment and orientation. | Vector Medical Inc. | A61B005/05
US06893407 | B1 | 2005-05-17 | Communication method and apparatus | 2000-05-05 | The present invention relates to a non-lingual communication method and apparatus, wherein a physical or physiological signal consciously created by a first subject (1) is detected and converted into a transmitted output signal presented to a second subject (7) in order to communicate information from the first subject (1) to the second subject (7). The invention further relates to rehabilitation of handicapped people. | 1. A non-lingual communication method comprising the steps of: selecting at least one input medium from a group of selectable input media and at least one output medium from a group of selectable output media,detecting by said at least one input medium physical signals consciously created by a first subject indicative of intentional information to be submitted to a second subject,converting said signals into a transmittable signal,transmitting said transmittable signal into a data transformer, which is functionally connected to said at least one output medium,transforming said transmittable signal to a sequence of output signals to be outputted from said at least one output medium,presenting to said second subject said sequence of output signals for communicating said intentional information from said first subject to said second subject,and calibrating the sensitivity of said at least one input medium to correspond to a specific amplitude range for said physical signals wherein said calibration step is performed according to a predetermined map;wherein said at least one input medium is selected from the group consisting of an electromagnetic radiation emitter and an electromagnetic radiation detector for detecting movements,a sound wave emitter and a reflected sound wave detector for detecting movements,a laser scanner and a light detector for detecting movements,means for motion detection;and said at least one output medium is selected from the group consisting of a sound generator,an image display,a vibration generator,a light pulse emitter,a heating device,a nerve exciting electrode, anda motor driven device. | 1. A non-lingual communication method comprising the steps of: selecting at least one input medium from a group of selectable input media and at least one output medium from a group of selectable output media,detecting by said at least one input medium physical signals consciously created by a first subject indicative of intentional information to be submitted to a second subject,converting said signals into a transmittable signal,transmitting said transmittable signal into a data transformer, which is functionally connected to said at least one output medium,transforming said transmittable signal to a sequence of output signals to be outputted from said at least one output medium,presenting to said second subject said sequence of output signals for communicating said intentional information from said first subject to said second subject,and calibrating the sensitivity of said at least one input medium to correspond to a specific amplitude range for said physical signals wherein said calibration step is performed according to a predetermined map;wherein said at least one input medium is selected from the group consisting of an electromagnetic radiation emitter and an electromagnetic radiation detector for detecting movements,a sound wave emitter and a reflected sound wave detector for detecting movements,a laser scanner and a light detector for detecting movements,means for motion detection;and said at least one output medium is selected from the group consisting of a sound generator,an image display,a vibration generator,a light pulse emitter,a heating device,a nerve exciting electrode, anda motor driven device. 2. The method of claim 1 wherein said transmittable signal comprises digital data. 3. The method of claim 1 wherein at least one of first and second subjects is a handicapped person. 4. The method of claim 1 wherein the input medium physical signals further comprise movements. 5. The method of claim 1 wherein said calibrating step is performed by an interactively learning computer including a neural network. 6. A non-lingual communication apparatus comprising: at least one input medium and at least one output medium selected from a group of selectable input and output media, said at least one input medium operable to detect physical signals consciously created by a first subject indicative of intentional information to be submitted to a second subject,a converter for converting said signals into a transmittable signal,a transmitter for transmitting said transmittable signal from said converter to a data transformer which is functionally connected to said at least one output medium, said data transformer transforming said transmittable signal into control commands for said at least one output medium, wherein said at least one output medium generates a sequence of output signals for communicating said intentional information from said first subject to said second subject, andcalibration means for calibration of the sensitivity of said... | Personics A/S | A61B005/103 A61B005/117
US06955908 | B1 | 2005-10-18 | Organism associated with nongonococcal urethritis | 2000-06-21 | The invention features a biologically pure culture of a newly identified single-celled organism Spiky Rotating Cells (SPR), methods to diagnose an SPR infection in a human patient, an instrument for collecting and detecting an SPR infection, and methods for treating an SPR infection. | 1. A biologically pure culture of a single-celled organism, Spiky Rotating Cells (SPR), having all of the characteristics of the deposited single-celled organism having the accession number ATCC PTA-2129, wherein said organism causes disease in humans and comprises the following biological characteristics: (i) provisional classification as a protozoan, (ii) a spherical shape measuring approximately 7–8 μm as a solitary single-celled organism, (iii) a refractile cell membrane, (iv) multiple circumferential spiky projections of the cell membrane, (v) rotatory motility, (vi) periodic clustering behavior to form colonies, and (vii) existence in an extracellular environment. | 1. A biologically pure culture of a single-celled organism, Spiky Rotating Cells (SPR), having all of the characteristics of the deposited single-celled organism having the accession number ATCC PTA-2129, wherein said organism causes disease in humans and comprises the following biological characteristics: (i) provisional classification as a protozoan, (ii) a spherical shape measuring approximately 7–8 μm as a solitary single-celled organism, (iii) a refractile cell membrane, (iv) multiple circumferential spiky projections of the cell membrane, (v) rotatory motility, (vi) periodic clustering behavior to form colonies, and (vii) existence in an extracellular environment. 2. The culture of claim 1, wherein said extracellular environment is skin, fluid of the genital tract, or the extracellular fluid of other organs. 3. The culture of claim 1, wherein said disease is nongonococcal urethritis. 4. The culture of claim 1, wherein said organism proliferates in Diamond's media. 5. A method of diagnosing an SPR infection in a human patient, said method comprising the steps of: a) obtaining a sample from said patient; andb) testing said sample for the presence of an organism, having all of the characteristics of the deposited single-celled organism having the accession number ATCC PTA-2129, that causes disease in humans, said organism having the following biological characteristics: (i) provisional classification as a protozoan, (ii) a spherical shape measuring approximately 7–8 μm as a solitary single-celled organism, (iii) a refractile cell membrane, (iv) multiple circumferential spiky projections of the cell membrane, (v) rotatory motility, (vi) periodic clustering behavior to form colonies, and (vii) existence in an extracellular environment. 6. The method of claim 5, wherein said method further comprises, after step a), testing the pH of said sample, wherein a pH greater than 6.0 is further indicative of the presence of said SPR infection. 7. The method of claim 5, wherein said patient is a male, and wherein step a) comprises obtaining said sample from the urethra of said male patient, wherein said sample comprises a secretion found in the urethra of said male patient. 8. The method of claim 5, wherein step b) comprises admixing said sample with saline and examining said sample by microscopy, wherein said SPR infection is confirmed by the presence of an organism comprising said biological characteristics. 9. The method of claim 8, wherein said extracellular environment is skin, fluid of the genital tract, or the extracellular fluid of an organ. 10. The method of claim 8, wherein said organism proliferates in Diamond's media. 11. The method of claim 5, wherein said patient is a female, and wherein step a) comprises obtaining a sample from the vagina of said female patient, wherein said sample comprises a cervico vaginal secretion from said female patient. 12. The method of claim 5, wherein said patient has a skin eruption or lymph node absce... | - | A61C019/04 A61B005/00
US06865415 | B2 | 2005-03-08 | Visceral fat determining device | 2001-04-16 | A visceral fat determining device is disclosed which can be used easily at home for example and can obtain information on visceral fat. The visceral fat determining device comprises: input means for inputting personal data that represents vital characteristics of the patient; a data processing unit capable of performing various calculations; and a display device capable of displaying the personal data and results of the calculations. Upon inputting of a waist size which is a girth of trunk of the patient, and a hip size which is a girth of buttocks of the patient, quantitative information on abdominal visceral fat of the patient is given through a calculation based on WHR which is a ratio of the waist size to the hip size (waist size/hip size). | 1. A visceral fat determining device comprising: input means for inputting personal data including at least an abdominal girth WL which is a circumferential length of a trunk of a patient and a gluteal girth HL which is a circumferential length of buttocks of the patient; a data processing unit for storing the personal data and for calculating quantitative information on abdominal visceral fat of the patient including an estimated value of an abdominal visceral fat cross-sectional area VA of the patient based on the personal data; and a display device for displaying the personal data and a result of the calculation performed by the data processing unit; wherein the data processing unit stores at least one regression coefficient of WHR and a regression constant, WHR being a ratio (WL/HL) of an abdominal girth WL to an gluteal girth HL, the regression coefficient and the regression constant being obtained from statistical processing based on actual measurement values of the abdominal visceral fat cross sectional area VA measured in abdominal tomography of human bodies of random samples and respective WHR values of the human samples, the data processing unit calculating the estimated value of abdominal visceral fat cross sectional area VA of the patient based on a WHR value of the patient, the regression constant of said WHR and the regression coefficient. | 1. A visceral fat determining device comprising: input means for inputting personal data including at least an abdominal girth WL which is a circumferential length of a trunk of a patient and a gluteal girth HL which is a circumferential length of buttocks of the patient; a data processing unit for storing the personal data and for calculating quantitative information on abdominal visceral fat of the patient including an estimated value of an abdominal visceral fat cross-sectional area VA of the patient based on the personal data; and a display device for displaying the personal data and a result of the calculation performed by the data processing unit; wherein the data processing unit stores at least one regression coefficient of WHR and a regression constant, WHR being a ratio (WL/HL) of an abdominal girth WL to an gluteal girth HL, the regression coefficient and the regression constant being obtained from statistical processing based on actual measurement values of the abdominal visceral fat cross sectional area VA measured in abdominal tomography of human bodies of random samples and respective WHR values of the human samples, the data processing unit calculating the estimated value of abdominal visceral fat cross sectional area VA of the patient based on a WHR value of the patient, the regression constant of said WHR and the regression coefficient. 2. The visceral fat determining device according to claim 1, wherein the quantitative information on the abdominal visceral fat further includes an amount of the abdominal visceral fat. 3. The visceral fat determining device according to claim 1, further comprising body fat ratio measuring means for measuring a bioelectrical impedance Z of the patient via electrodes contacted to end portions of the patient and for calculating a body fat ratio FAT of the patient based on the measured bioelectrical impedance Z and the inputted personal data or a portion thereof, wherein the body fat ratio FAT obtained by the body fat ratio measuring means is displayed on the display device. 4. The visceral fat determining device according to claim 1, wherein the abdominal girth WL is provided by an abdominal girth at the fourth lumbar vertebrae of the patient, and the gluteal girth HL is provided by a girth measured generally at the thickest portion on the buttocks of the patient. 5. The visceral fat determining device according to claim 1, further comprising size measuring means for measuring the abdominal girth WL and the gluteal girth HL. 6. The visceral fat determining device according to claim 5, wherein the abdominal girth WL and the gluteal girth HL measured by the size measuring means are inputted to the data processing unit. 7. The visceral fat determining device according to claim 1, wherein the calculation of the estimated value of abdominal visceral fat cross sectional area VA is performed with addition of a correction term by age and/or a correction term by sex, of the patient. 8. The visceral ... | Yamato Scale Co., Ltd. | A61B005/05
US06893396 | B2 | 2005-05-17 | Wireless internet bio-telemetry monitoring system and interface | 2001-05-18 | A system and method for monitoring patient variables in a wireless manner via a patient worn monitoring device is disclosed. The patient monitoring device is wearable and connects to a variety of sensors with at least one microphone for voice communications. The device connects to a wireless network and thence to the Internet for transmitting data to a Host for access by a medical care provider. The medical care provider communicates with the patient-wearable device via the Internet and the wireless network to send instructions to the patient-wearable monitoring unit and to communicate via voice with the patient. The medical care provider can also flexibly reconfigure the device to change collection parameters. When an alarm limit is exceeded as detected by the sensors, the data are transmitted to the Host computer for use by the medical care provider, thereby allowing full mobility to the patient wearing the device. | 1. A system for remotely monitoring patient variables comprising: at least one patient-wearable, multi-variable sensor means for acquiring medically relevant data from a patient; a patient-wearable monitoring unit connected to the sensor; said patient-wearable monitoring unit comprising a processor means for formatting sensor data from said patient into Extensible Markup Language (XML) and a wireless communication device connected to a first wireless network, wherein said first wireless network includes means adapted to send and receive communications over the Internet and a second network; a Host data archive connected to the Internet to communicate with a plurality of the patient-wearable monitoring units and archive data using XML; a plurality of medical care provider terminal means connected to the second network for communication in a bi-directional manner over said first and second network between each medical care provider and the patient-wearable monitoring unit of a patient under the medical care provider's care, and wherein said plurality of medical care provider terminals are further connected to the Internet for bi-directional XML data communication with said Host data archive. | 1. A system for remotely monitoring patient variables comprising: at least one patient-wearable, multi-variable sensor means for acquiring medically relevant data from a patient; a patient-wearable monitoring unit connected to the sensor; said patient-wearable monitoring unit comprising a processor means for formatting sensor data from said patient into Extensible Markup Language (XML) and a wireless communication device connected to a first wireless network, wherein said first wireless network includes means adapted to send and receive communications over the Internet and a second network; a Host data archive connected to the Internet to communicate with a plurality of the patient-wearable monitoring units and archive data using XML; a plurality of medical care provider terminal means connected to the second network for communication in a bi-directional manner over said first and second network between each medical care provider and the patient-wearable monitoring unit of a patient under the medical care provider's care, and wherein said plurality of medical care provider terminals are further connected to the Internet for bi-directional XML data communication with said Host data archive. 2. The system for remotely monitoring patient variables of claim 1 wherein the communication with the patient-wearable monitoring unit is bi-directional XML data communication. 3. The system for remotely monitoring patient variables of claim 2 wherein the bi-directional XML data communication further comprises instructions from the medical care provider terminal to change configurable program instructions in the processor of the patient-wearable monitoring unit. 4. The system for remotely monitoring patient variables of claim 3 wherein the instructions to change configurable program instructions comprises instructions to change alarm limits. 5. The system for remotely monitoring patient variables of claim 3 wherein the instructions to change configurable program instructions comprises instructions to change data collection parameters for the at least one patient-wearable sensor. 6. The system for remotely monitoring patient variables of claim 1 wherein the communication with the patient-wearable monitoring unit is bi-directional voice communication. 7. The system for remotely monitoring patient variables of claim 6 wherein the bi-directional voice communications comprises voice-over-IP communications with a medical care provider professional. 8. The system for remotely monitoring patient variables of claim 7 wherein the at least one sensor further comprises a microphone and a speaker for transmitting voice communications. 9. The system for remotely monitoring patient variables of claim 7 wherein voice messaging is assigned priority by the processor according to the type of message being sent. 10. The system for remotely monitoring patient variables of claim 6, further comprising a second network connected to said first wireless network and said medica... | i-Medik, Inc. | A61B005/00
US06955646 | B2 | 2005-10-18 | Method and medical system for the postdischarge surveillance of a patient | 2002-01-24 | A method and medical system for the postdischarge surveillance of a patient (1) for detecting a case of pneumonia, secondary bleeding, a wound healing problem, a pulmonary complication, a urinary tract infection or a thrombosis of the patient (1), having a data bank (13), which is arranged at a location other than the location (2) at which the patient (1) is based during the surveillance and stores data relevant for the postdischarge surveillance of the patient (1) recorded during the postdischarge surveillance. | 1. A method for monitoring a patient after the patient has been postoperatively discharged from a medical facility, the method comprising the steps of: at a location of the patient during the monitoring after the patient has been postoperatively discharged from the facility, providing the patient with a written questionnaire that gathers from the patient data suitable for detecting in the patient at least one of pneumonia, secondary bleeding, a wound healing problem, a pulmonary complication, a urinary tract infection, and a thrombosis; conveying the patient data in the written questionnaire to a data bank over a telephone connection using one of a voice transmission and a facsimile transmission; and automatically evaluating the conveyed data with an evaluating device that is associated with the data bank. | 1. A method for monitoring a patient after the patient has been postoperatively discharged from a medical facility, the method comprising the steps of: at a location of the patient during the monitoring after the patient has been postoperatively discharged from the facility, providing the patient with a written questionnaire that gathers from the patient data suitable for detecting in the patient at least one of pneumonia, secondary bleeding, a wound healing problem, a pulmonary complication, a urinary tract infection, and a thrombosis; conveying the patient data in the written questionnaire to a data bank over a telephone connection using one of a voice transmission and a facsimile transmission; and automatically evaluating the conveyed data with an evaluating device that is associated with the data bank. 2. The method of claim 1, further comprising the step of monitoring receipt of the patient data at the data bank and sending a signal to at least one of the patient and a caregiver when the patient data are not received. 3. The method of claim 1, further comprising the step of sending an alarm signal to at least one of the patient and a caregiver when the evaluating device indicates that the patient data are critical. 4. A method for monitoring a patient after the patient has been postoperatively discharged from a medical facility, the method comprising the steps of: providing on a world wide web site a questionnaire that is accessible with a computer at a location of the patient during the monitoring after the patient has been postoperatively discharged from the medical facility and that is connected to a communication network; entering patient data in the questionnaire with the computer, the patient data being suitable for detecting in the patient at least one of pneumonia, secondary bleeding, a wound healing problem, a pulmonary complication, a urinary tract infection, and a thrombosis; conveying the patient data in the questionnaire to a data bank over the communication network; and automatically evaluating the patient data with an evaluating device that is associated with the data bank. 5. The method of claim 4, further comprising the step of monitoring receipt of the patient data at the data bank and sending a signal to at least one of the patient and a caregiver when the patient data are not received. 6. The method of claim 4, further comprising the step of sending an alarm signal to at least one of the patient and a caregiver when the evaluating device indicates that the patient data are critical. 7. A system for monitoring a patient after the patient has been postoperatively discharged from a medical facility, the system comprising: a data bank at a location different from a location of the patient during the monitoring after the patient has been postoperatively discharged from the medical facility, said data bank being arranged and adapted to store patient data indicating in the patient at least one of pneumonia, secondary ble... | Siemens Aktiengesellschaft | A61B005/00
US06902534 | B2 | 2005-06-07 | Method and kit of components for delivering blood to a portable clinical analyzer | 2002-03-25 | A kit and a method are provided for facilitating point-of-care testing of bodily fluid. The kit includes a point-of-care testing cartridge and a plastic fitting for use with a syringe to deliver the fluid sample to the testing cartridge. In other embodiments, the kit includes a plastic cannula in place of a fitting. The syringe and plastic cannula can be used with an IV access kit to obtain a sample of fluid. The syringe and plastic cannula then can be used to deliver the fluid to the reservoir the testing cartridge. | 1. A kit for obtaining a sample of fluid from a patient and enabling point-of-care testing of said sample of fluid, said kit comprising: a syringe assembly for selectively drawing and expelling a sample of fluid, the syringe assembly including a distal end, a tip projecting at said distal end and having a passage communicating with a fluid receiving chamber of said syringe, said tip having selected outside cross-sectional dimensions; a point-of-care testing cartridge having an internal reservoir and an inlet port, the inlet port having cross-sectional dimensions configured for receiving the tip of the syringe; and a fitting having a proximal end configured for removable communication with the tip of said syringe, a distal end tapered to outside cross-sectional dimensions sufficiently smaller than the outside cross-sectional dimensions of the tip of the syringe for permitting removable communication in said entry port of said testing cartridge and venting of air from said entry port for delivering said sample of said fluid from said syringe to said reservoir of said testing cartridge. | 1. A kit for obtaining a sample of fluid from a patient and enabling point-of-care testing of said sample of fluid, said kit comprising: a syringe assembly for selectively drawing and expelling a sample of fluid, the syringe assembly including a distal end, a tip projecting at said distal end and having a passage communicating with a fluid receiving chamber of said syringe, said tip having selected outside cross-sectional dimensions; a point-of-care testing cartridge having an internal reservoir and an inlet port, the inlet port having cross-sectional dimensions configured for receiving the tip of the syringe; and a fitting having a proximal end configured for removable communication with the tip of said syringe, a distal end tapered to outside cross-sectional dimensions sufficiently smaller than the outside cross-sectional dimensions of the tip of the syringe for permitting removable communication in said entry port of said testing cartridge and venting of air from said entry port for delivering said sample of said fluid from said syringe to said reservoir of said testing cartridge. 2. A kit for obtaining a sample of fluid from a patient and enabling point-of-care testing of said sample of fluid, said kit comprising: a syringe assembly for selectively drawing and expelling a sample of fluid, said syringe assembly having a distal end, a tip projecting at said distal end; a point-of-care testing cartridge having an internal reservoir and an inlet port; and a fitting having a proximal end configured for removable mounting to said tip of said syringe, a distal end configured for removable communication with said entry port of said testing cartridge and a lumen extending between the ends for delivering said sample of said fluid from said syringe to said reservoir of said testing cartridge, said distal end of said fitting comprises a pair of converging triangular portions meeting at said distal end and separated from one another by transverse ports communicating with said lumen through said fitting. 3. A kit for use with a syringe to obtain a sample of fluid from a patient and to enable point-of-care testing of said sample of fluid, said kit comprising: a point-of-care testing cartridge having an internal reservoir and an inlet port; and a fitting having a proximal end configured for removable communication with said syringe, a distal end configured for removable communication with said entry port of said testing cartridge and a lumen extending between said ends for delivering said sample of said fluid from said syringe to said reservoir of said testing cartridge, said distal end of said fitting comprising a pair of converging triangular portions meeting at said distal end and separated from one another by transverse ports communicating with said lumen through said fitting. 4. A method for delivering a sample of fluid to a testing cartridge of a point-of-care blood testing system, said testing cartridge including an internal reservoir for retain... | Becton, Dickinson and Company | A61B005/00 B65D081/00
US06908619 | B1 | 2005-06-21 | Camp factor of streptococcus uberis | 1999-01-21 | The CAMP factor gene of Streptococcus uberis (S. uberis) is described, as well as the recombinant production of CAMP factor therefrom. CAMP factors can be used in vaccine compositions for the prevention and treatment of bacterial infections. | 1. An isolated nucleic acid molecule consisting of a sequence selected from the group consisting of: (a) a sequence encoding an immunogenic polypeptide having at least 90% sequence identity to the contiguous amino acid sequence shown at positions 1 through 256, inclusive, of SEQ ID NO:2; and (b) a sequence encoding an immunogenic polypeptide having at least 90% sequence identity to the contiguous amino acid sequence shown at positions 1-228, inclusive, of SEQ ID NO:5, wherein any sequence variation between the sequence of (a) and the sequence of amino acids shown at positions 1 through 256 of SEQ ID NO:2 is due to a conservative amino acid substitution, and any sequence variation between the sequence of (b) and the sequence of amino acids shown at positions 1-228 is due to a conservative amino acid substitution. | 1. An isolated nucleic acid molecule consisting of a sequence selected from the group consisting of: (a) a sequence encoding an immunogenic polypeptide having at least 90% sequence identity to the contiguous amino acid sequence shown at positions 1 through 256, inclusive, of SEQ ID NO:2; and (b) a sequence encoding an immunogenic polypeptide having at least 90% sequence identity to the contiguous amino acid sequence shown at positions 1-228, inclusive, of SEQ ID NO:5, wherein any sequence variation between the sequence of (a) and the sequence of amino acids shown at positions 1 through 256 of SEQ ID NO:2 is due to a conservative amino acid substitution, and any sequence variation between the sequence of (b) and the sequence of amino acids shown at positions 1-228 is due to a conservative amino acid substitution. 2. The nucleic acid molecule of claim 1 wherein said nucleic acid molecule encodes an immunogenic polypeptide having a sequence with at least 90% sequence identity to the contiguous amino acid sequence shown at positions 1 through 256, inclusive, of SEQ ID NO:2. 3. The nucleic acid molecule of claim 1 wherein said nucleic acid molecule encodes an immunogenic polypeptide having a sequence with at least 90% sequence identity to the contiguous amino acid sequence shown at positions 1-228, inclusive, of SEQ ID NO:5. 4. A recombinant vector comprising: (a) a nucleic acid molecule encoding an immunogenic polypeptide comprising a sequence selected from the group consisting of: (i) a sequence having at least 90% sequence identity to the contiguous amino acid sequence shown at positions 1 through 256, inclusive, of SEQ ID NO:2; and (ii) a sequence having at least 90% sequence identity to the contiguous amino acid sequence shown at positions 1-228, inclusive, of SEQ ID NO:5, wherein any sequence variation between the sequence of (i) and the sequence of amino acids shown at positions 1 through 256 of SEQ ID NO:2 is due to a conservative amino acid substitution, and any sequence variation between the sequence of (ii) and the sequence of amino acids shown at positions 1-228 is due to a conservative amino acid substitution; and (b) control elements that are operably linked to said nucleic acid molecule whereby said coding sequence can be transcribed and translated in a host cell, and at least one of said control elements is heterologous to said coding sequence. 5. A recombinant vector according to claim 4, wherein said nucleic acid molecule encodes an immunogenic polypeptide which comprises a sequence having at least 90% sequence identity to the contiguous amino acid sequence shown at positions 1 through 256, inclusive, of SEQ ID NO:2. 6. A recombinant vector according to claim 4, wherein said nucleic acid molecule encodes an immunogenic polypeptide which comprises a sequence having at least 90% sequence identity to the contiguous amino acid sequence shown at positions 1-228, inclusive, of SEQ ID NO:5. 7. A host cell transformed with the recombi... | University of Saskatchewan | A61B005/055 A01N063/00 A61K039/00 A61K039/38 A61K039/09
US06859660 | B2 | 2005-02-22 | Neuro-navigation system | 2002-02-26 | The invention relates to a Neuro-navigation system comprising a reflector referencing system including passive reflectors and a marker system with markers or landmarks wherein the reflectors as well as the markers as regards their shape, size and material selection as well as their arrangement or attachment on the parts of the body to be operatively treated and on the surgical instruments are configured so that mapping their locations is substantially facilitated or is able to take place more accurately positioned by a computer/camera unit having a graphic display terminal as well as the operative treatment with the aid of this unit. Optionally a surgical microscope, an ultrasonic diagnostic system as well as a calibration procedure may be integrated in the Neuro-navigation system in accordance with the invention. | 1. A calibration process for mapping the angular and spacing positions of referencing cameras, comprising positioning a calibration tool having two reflectors secured at predetermined positions at a known distance away from each other in the viewing range of said cameras, moving said calibration tool three-dimensionally in said viewing range, mapping several intermediate positions of said calibration tool by said referencing cameras and converting the resulting data by means of a computer unit individually into three-dimensional coordinates of the reflectors or calibration tool, and computing said angular and spacing positions of said cameras from said three-dimensional coordinates of said reflectors or calibration tool by means of said computer unit. | 1. A calibration process for mapping the angular and spacing positions of referencing cameras, comprising positioning a calibration tool having two reflectors secured at predetermined positions at a known distance away from each other in the viewing range of said cameras, moving said calibration tool three-dimensionally in said viewing range, mapping several intermediate positions of said calibration tool by said referencing cameras and converting the resulting data by means of a computer unit individually into three-dimensional coordinates of the reflectors or calibration tool, and computing said angular and spacing positions of said cameras from said three-dimensional coordinates of said reflectors or calibration tool by means of said computer unit. 2. The process set forth in claim 1, further comprising projecting on a graphic display terminal a projected relative position of said reflectors during said three-dimensional movement of said calibration tool. 3. The process set forth in claim 1, wherein a pointer provided with said removable reflectors is used as said calibration tool. 4. The process set forth in claim 1, wherein a calibration rod provided with said removable reflectors is used as said calibration tool. 5. A surgical microscope control system comprising: a microscope stand having a base, a microscope mounted to the stand for three-dimensional movement, a source of infrared radiation, at least two mapping infrared cameras, a computer unit connected to said cameras, a graphic display terminal connected to said computer unit, and at least three infrared radiation reflectors removably attached to said microscope in a characteristic arrangement, said microscope having optics and being adapted to focus the optics of said microscope on a point having known three-dimensional coordinates and provide focusing data representative of the coordinates of the point to the computer unit, said computer unit being adapted to receive said focal data from said microscope and to relate said focal data to the three-dimensional position of said microscope obtained by means of said reflectors and said cameras. 6. The control system set forth in claim 5, wherein one or more of the following control sequences are implemented by the computer unit: a) automatically tracking and focusing an instrument tip, the position of which is known to said computer unit via the use of infrared radiation reflectors and said cameras; b) automatically focusing a memorized or predetermined point of operative treatment; and c) focusing a point of operative treatment from various three-dimensional and angular positions of said microscope. | BrainLAB AG | A61B005/00
US06876878 | B2 | 2005-04-05 | Medical broad band electromagnetic holographic imaging | 2001-06-06 | A method of imaging an object, such as a diseased human heart or bone or malignant tumor, in a nontransparent medium, such as the human body, involves placing an array of transmitters and receivers in operational association with the medium. The transmitters generate a broad band harmonic (frequency domain) or pulse (time domain) primary electromagnetic field (EM) field, including the lower frequency portions of the EM spectrum, whose propagation is typically characterized by the diffusion phenomena, or by the combination of the diffusion and wave phenomena. The primary field propagates through the examined medium and interacts with the object to produce a scattered field, which is recorded by the receivers. The scattered EM field components measured by the receivers are applied as an artificial EM field to generate a backscattering EM field. Cross power spectra of the primary and backscattering fields (in the frequency domain) or cross correlation between these fields (in the time domain) produce a numerical reconstruction of an EM hologram. The desired properties of the medium, such as conductivity or dielectric permittivity, are then derived from this hologram. | 1. A method for imaging an anomalous region located within a nontransparent medium of an organism, said method comprising the steps of: a. placing one or more electromagnetic transmitters in transmission contact with said medium; b. placing one or more electromagnetic receivers at receiving positions with respect to said medium; c. operating said transmitters to generate a broad band electromagnetic field, comprising a frequency domain and/or time domain electromagnetic field, whereby said generated electromagnetic field propagates through said medium to interact with said anomalous region, resulting in a scattered electromagnetic field; d. measuring said scattered electromagnetic field with said receivers; e. obtaining a background field {Eb, Hb} representative of a background medium equivalent to said medium without the presence of said anomalous region; f. obtaining a backscattering anomalous field {Eas, Has} equivalent to that obtainable by illuminating said background medium with said scattered electromagnetic field transmitted from the receiving positions of said receivers; and g. producing a broad band holographic image of said anomalous region by calculating cross power spectra of said background and said backscattering fields, or calculating cross correlation functions between said background and said backscattering fields. | 1. A method for imaging an anomalous region located within a nontransparent medium of an organism, said method comprising the steps of: a. placing one or more electromagnetic transmitters in transmission contact with said medium; b. placing one or more electromagnetic receivers at receiving positions with respect to said medium; c. operating said transmitters to generate a broad band electromagnetic field, comprising a frequency domain and/or time domain electromagnetic field, whereby said generated electromagnetic field propagates through said medium to interact with said anomalous region, resulting in a scattered electromagnetic field; d. measuring said scattered electromagnetic field with said receivers; e. obtaining a background field {Eb, Hb} representative of a background medium equivalent to said medium without the presence of said anomalous region; f. obtaining a backscattering anomalous field {Eas, Has} equivalent to that obtainable by illuminating said background medium with said scattered electromagnetic field transmitted from the receiving positions of said receivers; and g. producing a broad band holographic image of said anomalous region by calculating cross power spectra of said background and said backscattering fields, or calculating cross correlation functions between said background and said backscattering fields. 2. The method of claim 1, wherein said produced image of said anomalous region identifies a medical condition of said anomalous region. 3. The method of claim 1, wherein said produced image of said anomalous region identifies an abnormal portion of said anomalous region. 4. The method of claim 1, wherein operating said transmitters to generate a broad band electromagnetic field comprises operating said transmitters to generate a broad band electromagnetic field with a frequency between approximately one megahertz and approximately ten megahertz. 5. The method of claim 1, wherein operating said transmitters to generate a broad band electromagnetic field comprises operating said transmitters to generate a broad band electromagnetic field with a frequency between approximately ten megahertz and approximately one hundred megahertz. 6. The method of claim 1, wherein operating said transmitters to generate a broad band electromagnetic field comprises operating said transmitters to generate a broad band electromagnetic field with a frequency of between approximately one megahertz and approximately one hundred megahertz. 7. The method of claim 1, wherein said broad band electromagnetic field comprises a portion of EM spectrum, wherein propagation of said portion is characterized by diffusion phenomena. 8. The method of claim 1, wherein said broad band electromagnetic field comprises a portion of EM spectrum, wherein propagation of said portion is characterized by a combination of diffusion phenomena and wave phenomena. 9. A system of imaging a body part of a human or animal, comprising: an imaging ring including one ... | University of Utah Research Foundation | A61B005/00
US07096055 | B1 | 2006-08-22 | Method to control delivery of radiation therapy | 1998-06-24 | A method is disclosed for controlling the delivery of radiation therapy to a tumor of a patient from one or more beams of ionizing radiation, to conform to a prescribed dosage of radiation for each of predetermined plural respectively shaped portions of the tumor according to the shape and other characteristics of the tumor. A radiation beam is selectively generated from different directions with respect to travel of the beam along a plurality of spatial paths including oscillating and arcuate movements. Parameters of the beam are calculated from conditions of distribution of a target dose, and the cross-section of the beam is adjusted so as to deliver the prescribed dosage of radiation to each of the respectively shaped portions of the tumor on which the beam impinges. The cross-section of the beam is constantly adjusted according to a predetermined area of the tumor which is to receive radiation therapy, and adjusted cross-sections of the beam are moved along selected individual ones of the spatial paths at least one time. Also, movements of the beam along individual spatial paths are split according to the plural portions of the tumor which are to receive different radiation doses. Travel of the beam is controlled along the selected individual paths so as to deliver radiation therapy within the prescribed dosage to each of the plural portions of the tumor in a minimum amount of time. In one embodiment, a micro multi-leaf collimator is placed between the beam and the tumor, and the collimator leaves are adjusted to change the cross-section of the beam impinging on a specified portion of the tumor according to the shape of the specified portion. | 1. A radiation therapy process for controlling the delivery of a beam of radiation of adaptable cross-section from any of a plurality of different directions to impinge on a preselected tumor region in a patient's body in a treatment session, said process comprising measuring the boundaries of said tumor region based on a selected fixed position of the patient's body; adjusting said boundaries to compensate for shifting of said tumor region from movement during patient breathing and other natural movements; subdividing the area occupied by said tumor region into subareas based upon differences in radiation dosage to be delivered thereto from said beam; specifying a maximum radiation dose to be delivered to each respective subarea in the treatment session; determining the dose rate applicable for each of said subareas for a plurality of discrete directions of impingement of said beam along a selected travel path thereof in a predetermined volume of space onto the respective subareas of said tumor region according to the distribution of said differences in radiation dosage for the respective maximum dose to be delivered thereto; calculating a minimum interval of time for said treatment session in which to dispense a radiation dosage not more than said maximum dose for all said subareas of said tumor region; and controlling the intensity, cross-section and direction of impingement of said radiation beam on said subareas, to deliver the respective radiation dosages throughout said tumor region in substantially said minimum interval of time. | 1. A radiation therapy process for controlling the delivery of a beam of radiation of adaptable cross-section from any of a plurality of different directions to impinge on a preselected tumor region in a patient's body in a treatment session, said process comprising measuring the boundaries of said tumor region based on a selected fixed position of the patient's body; adjusting said boundaries to compensate for shifting of said tumor region from movement during patient breathing and other natural movements; subdividing the area occupied by said tumor region into subareas based upon differences in radiation dosage to be delivered thereto from said beam; specifying a maximum radiation dose to be delivered to each respective subarea in the treatment session; determining the dose rate applicable for each of said subareas for a plurality of discrete directions of impingement of said beam along a selected travel path thereof in a predetermined volume of space onto the respective subareas of said tumor region according to the distribution of said differences in radiation dosage for the respective maximum dose to be delivered thereto; calculating a minimum interval of time for said treatment session in which to dispense a radiation dosage not more than said maximum dose for all said subareas of said tumor region; and controlling the intensity, cross-section and direction of impingement of said radiation beam on said subareas, to deliver the respective radiation dosages throughout said tumor region in substantially said minimum interval of time. 2. The radiation therapy process of claim 1, wherein the step of controlling said cross-section of said radiation beam includes adjusting the positions of leaves of a micro-multileaf collimator on which said beam impinges before impinging on said tumor region. 3. A method for controlling the delivery of radiation therapy to a tumor of a patient from one or more beams of ionizing radiation, to conform to a prescribed dosage of radiation for each of predetermined plural respectively shaped portions of the tumor according to the shape and other characteristics of the tumor, said method comprising the steps of: generating a radiation beam selectively from different directions with respect to travel of the beam along a plurality of spatial paths including oscillating and arcuate movements, to impinge on said plural portions of said tumor, calculating parameters of said beam from conditions of distribution of a target dose, and adjusting the cross-section of said beam so as to deliver said prescribed dosage of radiation to each of said predetermined plural respectively shaped portions of the tumor. 4. The method of claim 3, including the step of constantly adjusting said cross-section of said beam according to a predetermined area of aid tumor which is to receive said radiation therapy. 5. The method of claim 4, including the step of moving adjusted cross-sections of said beam along selected individual ones of said... | - | A61B005/05
US07005517 | B2 | 2006-02-28 | Paramagnetic metal-phthalocyanine complex compounds and contrast agent using the same | 2002-01-04 | The present invention provides a novel paramagnetic metal-phthalocyanine complexes and pharmaceutically acceptable salts thereof, which are useful as contrast agents for MRI(Magnetic Resonance Imaging), diagnostic X-ray imaging and computed tomography(CT).;The present invention also provides contrast agents for imaging, comprising the new paramagnetic metal-phthalocyanine complexes. The new contrast agents of the present invention show high imaging enhancement effects at lower concentration and are safer than the previously reported contrast agents. | 1. A compound having the structural formula: and the pharmaceutically acceptable salts thereof, wherein M is Gd(III) or Mn(II and III), and R and R1 are respectively OCH2CH2OCH2CH2OCH2CH2OCH3. | 1. A compound having the structural formula: and the pharmaceutically acceptable salts thereof, wherein M is Gd(III) or Mn(II and III), and R and R1 are respectively OCH2CH2OCH2CH2OCH2CH2OCH3. 2. A compound in which the compound of claim 1 is bonded to a macromolecule selected from the group consisting of amino acid, polypeptide, protein, and polysaccharide. 3. A contrast agent for MRI, which comprises a compound of claim 2. 4. A contrast agent for diagnostic X-ray imaging or computer tomography, which comprises a compound of claim 2. 5. A composition of a contrast agent, which comprises a compound of claim 2, and pharmaceutically acceptable carrier, adjuvant, or vehicle. 6. The composition of a contrast agent according to claim 5, which further comprises free organic ligands, or their pharmaceutically acceptable salts additionally. 7. The composition of a contrast agent according to claim 5, which further comprises free organic ligands, or calcium meglumin, sodium meglumin, or their complex salts additionally. 8. A method of enhancing tissue-specific contrast of MR images of organs and tissues of a mammal, which comprises the step of administrating a diagnostically effective amount of a compound of claim 2. 9. A contrast agent for MRI, which comprises a compound of claim 1. 10. A contrast agent for diagnostic X-ray imaging or computer tomography, which comprises a compound of claim 1. 11. A composition of a contrast agent, which comprises a compound of claim 1, and pharmaceutically acceptable carrier, adjuvant, or vehicle. 12. The composition of a contrast agent according to claim 11, which further comprises free organic ligands, or their pharmaceutically acceptable salts additionally. 13. The composition of a contrast agent according to claim 11, which further comprises free organic ligands, or calcium meglumin, sodium meglumin, or their complex salts additionally. 14. A method of enhancing tissue-specific contrast of MR images of organs and tissues of a mammal, which comprises the step of administrating a diagnostically effective amount of a compound of claim 1. | - | C09B047/04 C07D487/22 A61K049/04 A61B005/055 A01N055/02
US07300408 | B2 | 2007-11-27 | Spectroscopic breath analysis | 2002-08-16 | Methods and apparatus for the analysis of exhaled breath by spectroscopy are disclosed. An optical cavity containing the exhaled breath, typically comprising a pair of opposing high reflectivity mirrors, is used to implement a cavity enhanced absorbtion technique. Pairs of 12CO2 and 13CO2 absorbtion lines suitable for use in spectroscopic breath analysis are also disclosed. | 1. Exhaled breath analysis apparatus for quantifying the presence of one or more target substances in exhaled breath comprising: a cavity enhanced absorption assembly comprising an optical cavity coupled to an optical source operable to emit radiation and an optical detector configured to generate a signal representative of the intensity of radiation in said optical cavity, wherein the optical cavity is arranged to allow radiation emitted from the optical source to be repeatedly reflected and retrace its path to excite a plurality of cavity modes;a breath collection assembly arranged to pass at least a portion of said exhaled breath into said optical cavity for illumination by said radiation; anda data processor connected to said optical detector and adapted to quantify the presence of said one or more target substances in said optical cavity by measuring a reduction in the intensity of radiation in said optical cavity caused by variations in cavity ringdown time caused by absorption of said radiation by said target substance. | 1. Exhaled breath analysis apparatus for quantifying the presence of one or more target substances in exhaled breath comprising: a cavity enhanced absorption assembly comprising an optical cavity coupled to an optical source operable to emit radiation and an optical detector configured to generate a signal representative of the intensity of radiation in said optical cavity, wherein the optical cavity is arranged to allow radiation emitted from the optical source to be repeatedly reflected and retrace its path to excite a plurality of cavity modes;a breath collection assembly arranged to pass at least a portion of said exhaled breath into said optical cavity for illumination by said radiation; anda data processor connected to said optical detector and adapted to quantify the presence of said one or more target substances in said optical cavity by measuring a reduction in the intensity of radiation in said optical cavity caused by variations in cavity ringdown time caused by absorption of said radiation by said target substance. 2. The apparatus of claim 1 wherein said optical source is a tunable laser. 3. The apparatus of claim 2 further comprising a frequency controller connected to said tunable laser, the frequency controller being configured to cause the optical source to repeatedly change the frequency of said radiation within a predefined range comprising one or more absorption lines of said target substances. 4. The apparatus of claim 3 wherein the frequency controller is a sweep generator adapted to cause the optical source to repeatedly sweep the frequency of said radiation across a predefined range comprising one or more absorption lines of said target substances. 5. The apparatus of claim 3 wherein said absorption lines comprise at least a first line characteristic of 12CO2 and at least a second line characteristic of 13CO2. 6. The apparatus of claim 5 wherein said first line has a wavelength of about 1607.634 nm and said second line has a wavelength of about 1607.501 mm. 7. The apparatus of claim 1 further comprising a blind detector having substantially the same noise characteristics as said optical detector, said blind detector being optically isolated from said radiation and connected to said data processor. 8. The apparatus of claim 7 wherein said optical detector and said blind detector are connected to said data processor by a Dicke switch, and said data processor comprises a lock-in amplifier. 9. The apparatus of claim 1 wherein the optical cavity comprises at least two opposed mirrors. 10. The apparatus of claim 1 wherein said optical cavity is configured so as to yield an effective path length of said radiation within the optical cavity of at least 10 meters. 11. A method of quantifying the presence of one or more target substances in breath exhaled by a subject, the method comprising the steps of: collecting said exhaled breath;passing at least a portion of said exhaled breath into an optical cavity of a cavity enhan... | Isis Innovation Limited | G01N033/497 A61B005/08
US06997880 | B2 | 2006-02-14 | Waveform interpreter for respiratory analysis | 2002-02-28 | A capnograph which performs an analysis of the breath waveforms measured by the carbon dioxide sensor, interprets the results of this analysis, and outputs to the operator diagnostic information about the respiratory status of the patient, or about the adequacy of the breathing support provided to the patient. The instrument compares a number of parameters characteristic of the waveforms of the patient's breath with an internal library of the values of those parameters expected from normal waveforms stored in its memory. These parameters may either relate to specific features of the waveform in the time domain, or may characterize spectral components of the waveform in the frequency domain. The capnographic waveform analysis may be combined with further non-invasive measurements in order to provide an indication of the deviation of the value of EtCO2 from PaCO2. | 1. Apparatus for computerized breath analysis comprising: a breath receiver adapted to be placed in fluid communication with a subject;a breath analyzer coupled to said breath receiver which analyzes at least one breath of the subject, wherein said apparatus provides an indication of the discrepancy between the end tidal carbon dioxide partial pressure in breath of the subject and the arterial carbon dioxide partial pressure of the subject; anda corrector unit connected to said breath analyzer and operative to provide an indication of the arterial carbon dioxide partial pressure of the subject based on said discrepancy. | 1. Apparatus for computerized breath analysis comprising: a breath receiver adapted to be placed in fluid communication with a subject;a breath analyzer coupled to said breath receiver which analyzes at least one breath of the subject, wherein said apparatus provides an indication of the discrepancy between the end tidal carbon dioxide partial pressure in breath of the subject and the arterial carbon dioxide partial pressure of the subject; anda corrector unit connected to said breath analyzer and operative to provide an indication of the arterial carbon dioxide partial pressure of the subject based on said discrepancy. 2. Apparatus for computerized breath analysis according to claim 1 and wherein said breath analyzer comprises a carbon dioxide analyzer and an oxygen analyzer. 3. Apparatus for computerized breath analysis according to claim 2 said apparatus also comprising a computational unit, and wherein: said oxygen analyzer provides values of the partial pressures of inspired and expired oxygen in said at least one breath of the subject; andsaid computational unit utilizes the difference between said values to provide said indication of the discrepancy between said end tidal carbon dioxide partial pressure in breath of the subject and said arterial carbon dioxide partial pressure of the subject. 4. Apparatus for computerized breath analysis according to claim 3 and wherein said discrepancy is utilized to provide an indication of the arterial carbon dioxide partial pressure of the subject. 5. Apparatus for computerized breath analysis according to claim 4 and also comprising a respiration diagnosis generator providing an indication of the respiratory status of the subject based on said indication of the arterial carbon dioxide partial pressure of the subject. 6. Apparatus for computerized breath analysis according to claim 4 and wherein said breath analyzer performs analysis of at least one breath waveform. 7. Apparatus for computerized breath analysis according to claim 1 and also comprising a respiration diagnosis generator providing an indication of the respiratory status of the subject based on said indication of said discrepancy between said end tidal carbon dioxide partial pressure in breath of the subject and said arterial carbon dioxide partial pressure of the subject. 8. Apparatus for computerized breath analysis according to claim 1 and wherein said breath analyzer performs analysis of at least one breath waveform. 9. Apparatus for computerized breath analysis according to claim 1 and also comprising a pulmonary volume meter, and wherein respiratory volume measured by said meter is utilized to provide said indication of the discrepancy between the end tidal carbon dioxide partial pressure in breath of the subject and the arterial carbon dioxide partial pressure of the subject. 10. Apparatus for computerized breath analysis according to claim 1 and also comprising a pulmonary volume meter, said pulmonary volume meter providin... | Oridion Medical 1987, Ltd. | A61B005/08
US07044911 | B2 | 2006-05-16 | Gateway platform for biological monitoring and delivery of therapeutic compounds | 2001-10-29 | The invention relates to methods and devices for remote or distributed continuous monitoring of physiologically relevant states. The invention provides for methods to automatically detect deviations or other states in physiological parameters and automatically alert a measured subject, user or other authorized party. The device provides for a universal platform for sensors, and further provides for the automatic compensation or distribution of devices or bioactive agents at appropriate levels and/or intervals in response to deviations or other states sensed in various physiological parameters. | 1. A system for monitoring physiological status of a mammalian subject, comprising: a. one or more biointerface heads (BIH), each comprising one or both of a sensor for measuring a physiological parameter and a device for therapeutic compound delivery, at least one of said biointerface heads being implanted subdermally and configured to communicate data and a BIH identifier;b. at least one control and communication module (CCM) attached to an external surface of said subject and storing a CCM identifier and which is linked to at least one biointerface head to receive said data and said BIH identifier, wherein said control and communication module processes data from at least one biointerface head;c. at least one data collection unit (DCU) which receives data and said identifiers from the control and communication module; andd. a remote database management system which receives data from the at least one data collection unit and processes said data. | 1. A system for monitoring physiological status of a mammalian subject, comprising: a. one or more biointerface heads (BIH), each comprising one or both of a sensor for measuring a physiological parameter and a device for therapeutic compound delivery, at least one of said biointerface heads being implanted subdermally and configured to communicate data and a BIH identifier;b. at least one control and communication module (CCM) attached to an external surface of said subject and storing a CCM identifier and which is linked to at least one biointerface head to receive said data and said BIH identifier, wherein said control and communication module processes data from at least one biointerface head;c. at least one data collection unit (DCU) which receives data and said identifiers from the control and communication module; andd. a remote database management system which receives data from the at least one data collection unit and processes said data. 2. The system of claim 1, wherein said biointerface head is attached to said mammalian subject with at least one adhesive. 3. The system of claim 2, wherein said adhesive contains one or more of the groups comprising: growth factors, adherence molecules, adherence attractants or factors which promote cutaneous wound-healing mechanisms and formation of an epithelial-like structure around a portion of said system. 4. The system of claim 1, wherein said system comprises a transdermal conduit and a biofluid access port, and wherein the transdermal conduit and biofluid access port are coated with a hydrogel material. 5. The system of claims 4, wherein the hycirogel material contains preservatives, anti-inflammatory agents, antibiotics or antimicrobial agents. 6. The system of claims 4, wherein the hydrogel material contains a chemical, compound or molecule for calibration of the sensor. 7. The system of claim 4, wherein the transdermal conduit comprises a fluid material containing preservatives, anti-inflammatory agents, antibiotics or antimicrobial agents. 8. The system of claim 1, wherein said system comprises a chamber which releases one or more therapeutic agents. 9. A system comprising: a subdermal physiological parameter sensor to measure a physiological parameter of a mammalian subject and to generate measurement information based on the measurements;a mounting structure that anchors said system to a dermal layer;a flexible transdermal conduit attached to said mounting structure at or near a first end, wherein said transdermal conduit is in contact with a sensor assembly;a sensor mounting head, wherein said sensor mounting head is attached to a second end of said transdermal conduit;a biofluid access port within said sensor mounting head, the biofluid access port further comprising microstructures allowing biofluid flow into the transdermal conduit to contact the sensor assembly and block transmission of external pathogens into a subject;a control and communication module in data communicatio... | PhiloMetron, Inc. | A61B005/00
US07104963 | B2 | 2006-09-12 | Method and apparatus for monitoring intravenous (IV) drug concentration using exhaled breath | 2002-01-22 | A method and system is provided for detecting the depth of anesthesia wherein at least one anesthetic agent is absorbed in a patient's bloodstream during the administration of anesthesia, which includes sampling a patient's expired breath; analyzing the breath for concentration of at least one substance indicative of the anesthetic agent using sensor technology such as free (unmetabolized) anesthetic agent or its metabolites; determining the effect of the agent based on that concentration; and determining depth of anesthesia based thereon. The method also detects endogenous compounds such as ketones and ammonia in exhaled breath as well as other pathologic organisms. | 1. A method for determining the blood level concentration of at least one agent selected from the group consisting of anesthetics, analgesics, muscle relaxants, sedatives, and anxiolytics, wherein the agent is administered into a patient's bloodstream, comprising: sampling a patient's expired breath; analyzing the breath for concentration of at least one substance indicative of the agent using sensor technology; determining at least one blood level concentration based on the concentration of at least one substrance indicative of the at least one agent; and using a flow sensor to detect starting and completion of exhalation during said sampling step. | 1. A method for determining the blood level concentration of at least one agent selected from the group consisting of anesthetics, analgesics, muscle relaxants, sedatives, and anxiolytics, wherein the agent is administered into a patient's bloodstream, comprising: sampling a patient's expired breath; analyzing the breath for concentration of at least one substance indicative of the agent using sensor technology; determining at least one blood level concentration based on the concentration of at least one substrance indicative of the at least one agent; and using a flow sensor to detect starting and completion of exhalation during said sampling step. 2. The method of claim 1 wherein the breath is analyzed after a predetermined period of time. 3. The method of claim 1 further comprising the step of controlling an infusion pump for delivering the agent intravenously based on the determined blood level concentration. 4. The method of claim 1 wherein the agent is delivered by a delivery method selected from the group comprising: intravenous continuous delivery, parenteral delivery, sublingual delivery, transdermal delivery, and intravenous bolus delivery. 5. The method of claim 1 wherein the agent is delivered by continuous infusion. 6. The method of claim 1 wherein the agent is delivered by an infusion pump. 7. The method of claim 1 wherein the agent is selected from the group comprising Remifentanil and Propofol. 8. The method of claim 1 wherein the steps are repeated periodically to monitor trending over time. 9. The method of claim 1 wherein the agent is for amnesia. 10. The method of claim 1 wherein the agent is for analgesia. 11. The method of claim 1 wherein the agent is for muscle relaxation. 12. The method of claim 1 wherein the agent is for sedation. 13. The method of claim 1 wherein a combination of agents is administered. 14. The method of claim 1 wherein the determined blood level concentration is measured to determine anesthetic blood concentration. 15. The method of claim 1 wherein the determined blood level concentration is measured to determine analgesic blood concentration. 16. The method of claim 1 wherein the determined blood level concentration is measured for alevel indicative of recovery. 17. The method of claim 1 wherein the sampling is continuous. 18. The method of claim 1 wherein the sampling is periodic. 19. The method of claim 1 wherein the patient's breath is analyzed by sensor technology selected from semiconductor gas sensor technology, conductive polymer gas sensor technology, or surface acoustic wave gas sensor technology. 20. The method of claim 19 wherein the sensor technology produces a unique electronic fingerprint to characterize the concentration of said at least one substance. 21. The method of claim 1 further comprising the step of recording data resulting from analysis of the patient's breath. 22. The method of claim 1 further comprising the step of transmitting data resulting from analys... | University of Florida Research Foundation, Inc. | A61B005/08 A61M015/00 A61M016/00 A61M016/10
US07214187 | B2 | 2007-05-08 | System and method for identification of false statements | 2002-12-18 | A method of detecting lie detection including the steps of first conditioning the subject to produce an involuntary physiological response triggered when the subject provides false statements; and secondly asking the subject to state whether each of series of statements presented to the subject are true. As the subject answers, the subject is monitored for the conditioned involuntary response. Detecting whether the subject is providing false statements is based on whether the involuntary response is observed during these answers. A system (19) for lie detection includes a computing device (21), a controller (40), and at least one conditioning interface (42, 44). | 1. A system for detecting false statements made by a subject, the system comprising: (a) a programmed computing device, the programmed computing device including a software module operative to; (i) condition a subject to produce an involuntary physiological response pattern which is not a naturally occurring physiological response, when the subject states a false statement, and an opposite response pattern to a true statement; and to A. present the subject with a plurality of questions, the answer to some of which are known in advance;B. record the subject's answers to the plurality of questions while contemporaneously recording data related to the involuntary physiological responses; and(ii) analyze the recorded data to determine whether the subject is providing false statements by comparing conditioned responses to the statements to the distribution of conditioned responses shown to known true and known false statements;(b) an output device configured to receive data from the programmed computing device;(c) a controller configured to be operated by the programmed computing device;(d) at least one conditioning interface controllably operated by the controller, each conditioning interface further comprising: i. an attachment structure for attaching the conditioning interface to a selected body part of the subject; andii. a physiological measurement device. | 1. A system for detecting false statements made by a subject, the system comprising: (a) a programmed computing device, the programmed computing device including a software module operative to; (i) condition a subject to produce an involuntary physiological response pattern which is not a naturally occurring physiological response, when the subject states a false statement, and an opposite response pattern to a true statement; and to A. present the subject with a plurality of questions, the answer to some of which are known in advance;B. record the subject's answers to the plurality of questions while contemporaneously recording data related to the involuntary physiological responses; and(ii) analyze the recorded data to determine whether the subject is providing false statements by comparing conditioned responses to the statements to the distribution of conditioned responses shown to known true and known false statements;(b) an output device configured to receive data from the programmed computing device;(c) a controller configured to be operated by the programmed computing device;(d) at least one conditioning interface controllably operated by the controller, each conditioning interface further comprising: i. an attachment structure for attaching the conditioning interface to a selected body part of the subject; andii. a physiological measurement device. 2. A method of using a system for detecting false statements made by a subject, the system comprising a programmed computing device, an output device, a controller, and at least one controlling conditioning interface, the method comprising the steps of: (a) attaching at least one conditioning interface to the subject;(b) conditioning the subject to produce an involuntary physiological response which is not a naturally occurring physiological response pattern when the subject states a false statement, and an opposite response pattern to a true statement;(c) subjecting the subject to a testing stage comprising the steps of: i. presenting the subject with a plurality of questions which are answered by statements, the answers to some of which are known in advance;ii. recording the subject's answers to the plurality of questions while contemporaneously recording data related to the involuntary physiological response; and(d) analyzing the recorded data to determine whether the subject is providing false statements by comparing conditioned responses to the statements, to the distribution of conditioned responses shown to known true and known false statements. 3. The method of claim 2, wherein the step of conditioning the subject to produce an involuntary physiological response is comprised of the following steps: (a) exposing the subject to a statement that is a true statement to the subject;(b) stimulating a first body part of the subject with a first stimulus and contemporaneously stimulating a second body part of the subject with a second stimulus to obtain a first physiological response, the fir... | The University of Chicago | A61B005/00
US06988993 | B2 | 2006-01-24 | Biophysical sensor | 2001-06-20 | A biophysical sensor including a sound vibration sensing element (12) to produce a sensed output, and two or more separate signal processing paths to process the sensed outputs to produce output signals. The output signals may reflect specific frequency bands within the raw biological signal. In a further aspect the invention concerns an electronic stethoscope in which the sound vibration sensing element (12) is a PVDF membrane, the signal processing involves an operational amplifier (OP amp) (10) connected to the PVDF membrane, and a unity gain buffer amplifier is connected to the output of the OP amp to allow the connection of headphones directly to the sensor output. | 1. A biophysical sensor including a sound vibration sensing element using a PVDF membrane to produce a sensed output, and at least two separate signal processing paths to process the sensed output to produce at least two output signals; where one signal processing path includes an operational amplifier (OP Amp) connected to the PVDF membrane and for amplifying an audible frequency range of the sensed output to an audible amplitude range, and a unity gain buffer amplifier is connected to the output of the OP Amp to allow connection of headphones directly to the sensor output. | 1. A biophysical sensor including a sound vibration sensing element using a PVDF membrane to produce a sensed output, and at least two separate signal processing paths to process the sensed output to produce at least two output signals; where one signal processing path includes an operational amplifier (OP Amp) connected to the PVDF membrane and for amplifying an audible frequency range of the sensed output to an audible amplitude range, and a unity gain buffer amplifier is connected to the output of the OP Amp to allow connection of headphones directly to the sensor output. 2. A biophysical sensor according to claim 1, further including recording and display apparatus to record and display the output signals. 3. A biophysical sensor according to claim 1 or 2, where the output signals are two, or more, of heart sounds, heart movement, breathing movement and breathing sounds. 4. A biophysical sensor according to claim 1, where a silver coating on the PVDF membrane is used to take off electrical signals. 5. A biophysical sensor according to claim 1, where the signal processing includes floating amplifiers mounted on a printed circuit board inside the sensor housing. 6. A biophysical sensor according to claim 1, where the OPAmp uses the intrinsic capacitance of the PVDF together with a resistance to create a high pass filter. 7. A biophysical sensor according to claim 6, where alternatively, or in addition, an OPAmp is AC coupled directly to the PVDF membrane. 8. An electronic stethoscope, comprising a biophysical sensor including a sound vibration sensing element which uses a PVDF membrane to produce a sensed output, signal processing involving an operational amplifier (OPAmp) connected to the PVDF membrane and for amplifying an audible frequency range of the sensed output to an audible amplitude range, and a unity gain buffer amplifier connected to the output of the OPAmp to allow connection of headphones directly to the sensor output. | Australian Centre for Advanced Medical Technology Ltd. | A61B005/02
USRE038728 | E1 | 2005-04-19 | Breath test analyzer | 2002-04-16 | A breath test analyzer, which analyzes exhaled breadths of a patient for isotope labeled products generated in the patient's body after ingestion by the patient of an isotope labeled substance, where the presence of these isotope labeled products provide an indication of a medical condition in the patient. The analyzer uses a very sensitive infra-red spectrophotometer, which enables it to continuously collect and analyze multiple samples of the patient's breath, and process the outputs in real time, while the patient is still connected to the analyzer, such that a definitive result is obtained within a short time, such as the order of a few minutes. The breath test analyzer is sufficiently small in that it can be easily accommodated in the office of a physician. The breath test analyzer can be utilized for a number of diagnostic breath tests, according to the isotope labeled substance ingested by the patient and the gases detected in the patient's breath. | 1. A method of breath testing in which samples collected from exhaled breaths of a subject are analyzed by means of an analyzer, for a product of an isotopically labeled substance ingested by said subject, and comprising the steps of: performing a first analyzing of a first sample collected from at least a first exhaled breath of said subject; performing a second analyzing of a second sample collected from at least a second exhaled breath of said subject, at least said second sample being collected following the subject's ingestion of said substance; and providing an indication of a change in the level of said isotope labeled product in said second sample, while said subject is coupled to said analyzer. | 1. A method of breath testing in which samples collected from exhaled breaths of a subject are analyzed by means of an analyzer, for a product of an isotopically labeled substance ingested by said subject, and comprising the steps of: performing a first analyzing of a first sample collected from at least a first exhaled breath of said subject; performing a second analyzing of a second sample collected from at least a second exhaled breath of said subject, at least said second sample being collected following the subject's ingestion of said substance; and providing an indication of a change in the level of said isotope labeled product in said second sample, while said subject is coupled to said analyzer. 2. The method of claim 1 wherein at least one of said first sample collected from at least a first exhaled breath of a subject and said sample collected from at least a second exhaled breath of said subject is at least one essentially complete exhaled breath of said subject. 3. The method of claim 1, wherein the first sample is collected prior to the subject's ingestion of the isotopically labeled substance, and the second sample is collected following the subject's ingestion of said isotopically labeled substance. 4. The method of claim 1, wherein both the first and second samples are collected following the subject's ingestion of said isotopically labeled substance. 5. The method of claim 1, wherein said second sample is exhaled following analysis of said first sample. 6. The method of claim 1, and further comprising the step of coupling the subject to the analyzer substantially continuously from the analyzing of said first sample to the analyzing of said second sample. 7. The method of claim 1, and further comprising the step of coupling the subject to the analyzer substantially continuously from the collection of said first sample to the collection of said second sample. 8. The method of claim 1, wherein the subject is coupled to said analyzer by means of a nasal cannula. 9. The method of claim 1, wherein the first analyzing of said subject's exhaled breath takes place prior to said product being detectable in said subject's breath and the second analyzing of said subject's exhaled breath takes place once said product could be detectable in said subject's breath. 10. The method of claim 1, and further comprising the step of determining that said test has a clinically significant outcome in accordance with the results of the ongoing analyses of said samples. 11. The method of claim 1, characterized in that said samples are collected by means of a nasal cannula. 12. The method of claim 1 and further comprising the steps of: fitting the results of the analyzing of successive samples to a curve; and determining from said curve whether said test has a clinically significant outcome. 13. The method of claim 12 wherein the derivative of said curve is used for said determining. 14. The method of claim 1, wherein the collection of said s... | Oridion Medical, LTD | A61B005/08
US06917829 | B2 | 2005-07-12 | Method and system for a distributed analytical and diagnostic software over the intranet and internet environment | 2001-08-09 | An apparatus for monitoring the medical condition of a patient, including a monitoring device which during use monitors one or more clinical features of the patient; a predictive instrument arranged to receive output from the monitoring device and programmed to compute a probability of a medical outcome or diagnosis based on the monitored one or more clinical features, the predictive instrument further programmed to compute that probability by executing an algorithm which models the medical outcome or diagnosis; and an interface through which a user enters information characterizing the patient, wherein the predictive instrument is further programmed to request that algorithm from a remote location. | 1. An apparatus for monitoring the medical condition of a patient, said apparatus comprising: a monitoring device which during use monitors one or more clinical features of the patient; a predictive instrument arranged to receive output from the monitoring device and programmed to compute a probability of a medical outcome or diagnosis based on the monitored one or more clinical features, said predictive instrument further programmed to compute said probability by executing an algorithm which models said medical outcome or diagnosis; and a user interface through which a user enters information characterizing the patient, wherein said predictive instrument is further programmed to request said algorithm from a remote location. | 1. An apparatus for monitoring the medical condition of a patient, said apparatus comprising: a monitoring device which during use monitors one or more clinical features of the patient; a predictive instrument arranged to receive output from the monitoring device and programmed to compute a probability of a medical outcome or diagnosis based on the monitored one or more clinical features, said predictive instrument further programmed to compute said probability by executing an algorithm which models said medical outcome or diagnosis; and a user interface through which a user enters information characterizing the patient, wherein said predictive instrument is further programmed to request said algorithm from a remote location. 2. The apparatus of claim 1 wherein the monitoring device is an electrocardiograph. 3. The apparatus of claim 2 wherein the algorithm enables the predictive instrument to compute said probability using a regression equation. 4. The apparatus of claim 3 wherein the regression equation is of the form: P=100⁡[1-11+ⅇb0+∑bi⁢xi]wherein P is the probability of medical outcome or diagnosis, b0 is a constant, the xi's are explanatory variables, and the bi's are coefficients of corresponding explanatory variables. 5. The apparatus of claim 4 wherein P is a probability of acute cardiac ischemia. 6. The apparatus of claim 1 wherein said predictive instrument is programmed to request said algorithm from a server at the remote location. 7. The apparatus of claim 6 wherein said apparatus further comprises a communication interface for communicating with the server over a network and wherein the predictive instrument is programmed to send a request for said algorithm through the communication interface. 8. The apparatus of claim 7 wherein the predictive instrument is further programmed to send a patient profile along with the request for said algorithm and to receive said algorithm from the server through communication interface. 9. The apparatus of claim 8 wherein the patient profile characterizes a specific population to which the patient belongs and said algorithm is tailored for use with members of that specific population. 10. The apparatus of claim 7 wherein the predictive instrument is further programmed to specify in the request an analysis type which corresponds to the medical outcome or diagnosis for which the probability will be computed. 11. The apparatus of claim 1 wherein the predictive instrument includes a local repository for storing code which is to be executed by the predictive instrument and wherein the predictive instrument is programmed to request said algorithm from the remote location for storage in the local repository. 12. An apparatus for enabling a remotely located predictive instrument to compute a probability of a medical outcome or diagnosis based on monitored one or more clinical features of a patient, said apparatus comprising: a data storage area which stores a plurality of different algorithms... | Clinical Care Systems, Inc. | A61B005/0402
US06875418 | B2 | 2005-04-05 | Method of early detection of Duchenne muscular dystrophy and other neuromuscular disease | 2002-06-19 | The mdx mouse is a model of Duchenne muscular dystrophy. The present invention describes that mdx mice exhibited clinically relevant cardiac phenotypes. A non-invasive method of recording electrocardiograms (ECGs) was used to a study mdx mice (n=15) and control mice (n=15). The mdx mice had significant tachycardia, consistent with observations in patients with muscular dystrophy. Heart-rate was nearly 15% faster in mdx mice than control mice (P | 1. A method of diagnosing or aiding in diagnosing Duchenne muscular dystrophy in a subject, the method comprising the steps of: a) administering to said subject, an effective amount of a cholinergic agent or a derivative thereof; and, b) determining whether the cholinergic agent or a derivative thereof, induces a decrease in the heart-rate of the subject in response to the cholinergic agent or a derivative thereof: wherein a decrease in said heart rate is indicative that said subject has Duchenne muscular dystrophy. | 1. A method of diagnosing or aiding in diagnosing Duchenne muscular dystrophy in a subject, the method comprising the steps of: a) administering to said subject, an effective amount of a cholinergic agent or a derivative thereof; and, b) determining whether the cholinergic agent or a derivative thereof, induces a decrease in the heart-rate of the subject in response to the cholinergic agent or a derivative thereof: wherein a decrease in said heart rate is indicative that said subject has Duchenne muscular dystrophy. 2. The method of claim 1, wherein the decrease in the heart rate is determined using a detector supplying data associated with electrocardiogram signals and providing an electrocardiogram using an apparatus for interpreting the data associated with electrocardiogram signals of the subject. 3. The method of claim 1, wherein the decrease in heart rate activity is detected by a computing apparatus. 4. The method of claim 3, wherein the computing apparatus further comprises a digital processor, a working memory, and a database, which records changes in heart rate in the subject. 5. The method of claim 1, wherein the perturbation is caused by an alteration in property and function of an acetylcholine receptor. 6. The method of claim 1, wherein the perturbation is caused by an alteration in property and function of a cholinergic receptor. 7. The method of claim 6, wherein the cholinergic receptor is a muscarinic receptor. 8. The method of claim 1, wherein the perturbation is caused by a perturbation in acetylcholine concentration. 9. The method of claim 1, wherein the perturbation is caused by perturbation in acetylcholinesterase concentration. 10. The method of claim 1, wherein the perturbation is caused by a deficiency in dystrophin or a dystrophin-associated protein. 11. The method according to claim 1, wherein the cholinergic agent or a derivative thereof, is a muscarinic agent. 12. The method claim 11, wherein the muscarinic agent is a muscarinic antagonist or a derivative thereof. 13. The method of claim 12, wherein the muscarinic antagonist is selected from a group consisting of; an atropine, a latrotoxin, a scopolamine, a propantheline, a pirezapine, and a derivative thereof. 14. The method of claim 1, wherein the agent is administered intravenously, intraperitonially, ophthalmologically, buccal, intracoronary, intramuscularly, topically, intranasally, rectally, sublingually, orally, subcutaneously, by patch, or inhalation. 15. The method of claim 1, wherein the subject is a human. 16. The method of claim 1, wherein the heart rate of said subject is compared to a population of aged-matched healthy individuals. 17. The method of claim 12, wherein the result of the administration of a muscarinic antagonist provides relief of a heart disorder in subjects who demonstrate a reduction in heart rate after administration of the cholinergic agent or a derivative thereof. 18. The method of claim 13, wherein the muscarinic an... | Mouse Specifics, Inc. | A61K049/00 A61B005/00
US06860866 | B1 | 2005-03-01 | Method of determining the efficiency of a dialyzer of a dialysis machine and a dialysis machine for carrying out this method | 2000-06-22 | A method and a device for determining the efficiency of a dialyzer of a dialysis machine during a dialysis treatment wherein the dialyzer is divided by a semipermeable membrane into a blood chamber and a dialysis fluid chamber and wherein the blood flows at a predetermined flow rate through the blood chamber and dialysis fluid flows at a predetermined flow rate through the dialysis fluid chamber. In use, the dialysance or clearance for a predetermined dialysis fluid flow rate and/or blood flow rate is measured, and then the dialysance or clearance for any dialysis fluid flow rate and blood flow rate is calculated. | 1. A dialysis machine, comprising: a dialyzer; a semipermeable membrane dividing the dialyzer into a blood chamber and a dialysis fluid chamber, the blood chamber having an inlet and an outlet, and the dialysis fluid chamber having an inlet and an outlet; an arterial blood line connected to the inlet of the blood chamber; a venous blood line connected to the outlet of the blood chamber; a dialysis fluid inlet line connected to the inlet of the dialysis fluid chamber; a dialysis fluid outlet line connected to the outlet of the dialysis fluid chamber; a first device for determining at least a first measured or predetermined flow rate and a second measured or predetermined flow rate selected from the group of a blood flow rate Qb through the blood chamber, a dialysis fluid flow rate Qd through the dialysis fluid chamber, and an ultrafiltration rate Qf; a computer unit configured to: establish at least one of a clearance and a dialysance at the first flow rate; and determine the at least one of the clearance and the dialysance at the second flow rate on the basis of the at least one of the clearance and a dialysance established at the first flow rate. | 1. A dialysis machine, comprising: a dialyzer; a semipermeable membrane dividing the dialyzer into a blood chamber and a dialysis fluid chamber, the blood chamber having an inlet and an outlet, and the dialysis fluid chamber having an inlet and an outlet; an arterial blood line connected to the inlet of the blood chamber; a venous blood line connected to the outlet of the blood chamber; a dialysis fluid inlet line connected to the inlet of the dialysis fluid chamber; a dialysis fluid outlet line connected to the outlet of the dialysis fluid chamber; a first device for determining at least a first measured or predetermined flow rate and a second measured or predetermined flow rate selected from the group of a blood flow rate Qb through the blood chamber, a dialysis fluid flow rate Qd through the dialysis fluid chamber, and an ultrafiltration rate Qf; a computer unit configured to: establish at least one of a clearance and a dialysance at the first flow rate; and determine the at least one of the clearance and the dialysance at the second flow rate on the basis of the at least one of the clearance and a dialysance established at the first flow rate. 2. The dialysis machine of claim 1, wherein the computer unit calculates a diffusive component D1 of the dialysance or the clearance at the predetermined dialysis fluid flow rate Qd1, the predetermined blood flow rate Qb1 , and the predetermined ultrafiltration rate Qf1, from the clearance K1 or the dialysance at the predetermined dialysis fluid flow rate Qd1, the predetermined blood flow rate Qb1, and the predetermined ultrafiltration rate Qf1 according to the following equation: D1=K1-Qf11-Qf1/Qe1,wherein an effective blood flow Qe is calculated from the blood flow rate Qb corresponding to a delivery rate of a blood pump according to the following equation: Qe=Qb⁡(1-HCT100)⁢Fp,wherein HCT is the hematocrit (%) and Fp is the plasma water fraction. 3. The dialysis machine of claim 2, wherein the computer unit calculates a diffusive dialysance D(Qd(t), Qe(t)) for any dialysis fluid flow rate Qd(t), blood flow rate Qe(t) and ultrafiltration rate Qf(t), from the clearance K1 or the dialysance at the predetermined dialysis fluid flow rate Qd1, the predetermined blood flow rate Qb1, and the predetermined ultrafiltration rate Qf1 according to the following equations: D⁡(Qd⁡(t),Qe⁡(t))=Qe⁡(t)·(1-exp⁡(Qd⁡(t)Qd1⁢ln⁡(1-DQecorrQd⁡(t))))wherein DQecorr=Qd1⁡(1-exp⁡(Qe⁡(t)Qe1⁢ln⁡(1-D1Qd1))). 4. The dialysis machine of claim 3, wherein the computer unit calculates the sum of the diffusive and convective dialysance K(Qd(t), Qe(t), Qf(t)) or the clearance from the diffusive dialysance D(Qd(t), Qe(t)) according to the following equation: K⁡(Qd⁡(t),Qe⁡(t),Qf⁡(t))=D⁡(Qd⁡(t),Qe⁡(t))⁢(1-Qf⁡(t)Qe⁡(t))+Qf⁡(t). 5. The dialysis machine of claim 1, wherein the first device is configured so that at least one of the clearance and the dialysance can be determined for different times t during a dialysis treatment and effec... | Fresenius Medical Care Deutschland GmbH | A61M037/00 B01D063/00 B01D035/00 C02F001/44 A61B005/00
US06842637 | B2 | 2005-01-11 | Magnetic field measurement apparatus | 2001-09-10 | A magnetic field measurement apparatus includes a plurality of magnetometers each having SQUID's and three detection coils, one of which detects each of three orthogonal directional magnetic field components (Bx, By, Bz) of a magnetic field generated from a subject to be inspected, a display which displays time variation of waveforms of the magnitude (√(Bx2+By2+Bz2)) of magnetic field synthesized by square sum of each of the three orthogonal directional magnetic field components of the magnetic field generated from the subject to be inspected, a holder for holding a Dewar's vessel for arranging magnetometers therein, and a controller for controlling a positional relationship between the subject to be inspected and the Dewar's vessel. Accurate time variation of the magnetic field generated from the subject to be inspected can be detected without influence of positional change of the subject to be inspected, by simultaneously measuring each of the three orthogonal directional magnetic field components (Bx, By, Bz) of the magnetic field generated from current sources in the subject to be inspected. | 1. A magnetic field measurement apparatus comprising: a magnetometer comprising SQUIDs and three detection coils each of which detects each of three orthogonal directional magnetic field components of a magnetic field generated from a fetus in a living body; a computer which synthesizes the magnitude of the magnetic field by a square sum of said three orthogonal directional magnetic field components of the magnitude field generated from the fetus in the living body; and a display which displays the time variation of a waveform of said magnitude of the synthesized magnetic field; wherein the signal of each of the three orthogonal directional magnetic field components detected by the magnetometer is divided into two signals, the magnitude of the magnetic field is synthesized by using one of the two signals, and the time variation of waveform of the magnitude of the synthesized magnetic field after passing through a high pass filter is displayed on the display. | 1. A magnetic field measurement apparatus comprising: a magnetometer comprising SQUIDs and three detection coils each of which detects each of three orthogonal directional magnetic field components of a magnetic field generated from a fetus in a living body; a computer which synthesizes the magnitude of the magnetic field by a square sum of said three orthogonal directional magnetic field components of the magnitude field generated from the fetus in the living body; and a display which displays the time variation of a waveform of said magnitude of the synthesized magnetic field; wherein the signal of each of the three orthogonal directional magnetic field components detected by the magnetometer is divided into two signals, the magnitude of the magnetic field is synthesized by using one of the two signals, and the time variation of waveform of the magnitude of the synthesized magnetic field after passing through a high pass filter is displayed on the display. 2. A magnetic field measurement apparatus according to claim 1, further comprising: a coil bobbin which comprises an insulator material and holds the three detection coils in three orthogonal directions, respectively; wherein the three detection coils have planes orthogonal to one other, two of the three detection coils pass through a center axis of the coil bobbin, a plane formed by one of the three detection coils is perpendicular to the center axis of the coil bobbin, and two planes formed by the other two of the three detection coils are orthogonal to each other and are parallel to the center axis of the coil bobbin. 3. A magnetic field measurement apparatus comprising: a magnetometer comprising SQUIDs and three detection coils each of which detects each of three orthogonal directional magnetic field components of a magnetic field respectively, which are generated from a fetus in a living body; a computer which synthesizes the magnitude of the magnetic field by a square sum of said three orthogonal directional magnetic field components of the magnitude of the magnetic field generated from the fetus in the living body; and a display which displays the time variation of a waveform of said magnitude of the synthesized magnetic field; wherein the signal of each of the three orthogonal directional magnetic field components detected by the magnetometer is divided into two signals, the magnitude of the magnetic field is synthesized by using one of the two signals after passing said one of the two signals through a high pass filter, and the time variation of waveform of the magnitude of the synthesized magnetic field is displayed on the display. 4. A magnetic field measurement apparatus according to claim 3, further comprising: a coil bobbin which comprises an insulator material and holds the three detection coils in three orthogonal directions, respectively; wherein the three detection coils have planes orthogonal to one other, two of the three detection coils pass through a center axis of... | Hitachi, Ltd. | A61B005/05
US06918882 | B2 | 2005-07-19 | Guidewire with stiffness blending connection | 2001-10-05 | Alternative designs and materials suitable for connecting different guidewire sections together. More particularly, connecting two portions of a guidewire having different material compositions with a connector having a third composition which is readily weldable to each of the dissimilar portions of the guidewire. A transition area may be designed to provide a region of desired flexibility characteristics. | 1. A guidewire, comprising: a proximal section having a distal portion, the proximal section comprising a first material, wherein the first material comprises stainless steel; a distal section having a proximal portion, the distal section comprising a second material, wherein the second material comprises a nickel-titanium alloy; and a connector disposed adjacent and welded to the distal portion of the proximal section and the proximal portion of the distal section, the connector comprising a third material different from the first material and the second material, the third material comprising nickel-chromium-molybdenum alloy, C276, or B2 alloy that is compatible for welding to both stainless steel and nickel-titanium alloy. | 1. A guidewire, comprising: a proximal section having a distal portion, the proximal section comprising a first material, wherein the first material comprises stainless steel; a distal section having a proximal portion, the distal section comprising a second material, wherein the second material comprises a nickel-titanium alloy; and a connector disposed adjacent and welded to the distal portion of the proximal section and the proximal portion of the distal section, the connector comprising a third material different from the first material and the second material, the third material comprising nickel-chromium-molybdenum alloy, C276, or B2 alloy that is compatible for welding to both stainless steel and nickel-titanium alloy. 2. A guidewire as in claim 1, wherein the proximal section has a first flexibility and the distal section has a second flexibility, and wherein the distal portion of the proximal section and the proximal portion of the distal section overlap to define a region that blends the first flexibility with the second flexibility. 3. A guidewire as in claim 1, wherein the distal portion of the proximal section has a reduced size portion and the proximal portion of the distal section has a reduced size portion. 4. A guidewire as in claim 3, wherein the reduced size portions have a taper. 5. A guidewire as in claim 3, wherein the reduced size portions have an interlocking shape. 6. A guidewire as in claim 3, wherein at least a portion of the connector is disposed between the reduced size distal portion of the proximal section and the reduced sized proximal portion of the distal section. 7. A guidewire as in claim 1, wherein the third material comprises a nickel-chromium-iron alloy. 8. A guidewire as in claim 1, wherein the connector comprises a tubular member disposed about the distal portion of the proximal section and the proximal portion of the distal section. 9. A guidewire as in claim 8, wherein the tubular connector has an outside diameter that is the same as an outside diameter of the proximal section. 10. A guidewire as in claim 9, wherein the tubular connector has an outside diameter that is the same as an outside diameter of the distal section. 11. A guidewire, comprising: a proximal section having a distal portion, the proximal section comprising stainless steel; a distal section having a proximal portion, the distal section comprising a nickel-titanium alloy; and a connector disposed adjacent and welded to the distal portion of the proximal section and the proximal portion of the distal section, the connector comprising a nickel-chromium-molybdenum alloy. 12. A guidewire as in claim 11, wherein the proximal section has a first flexibility transition region having a first flexibility and the distal section has a second flexibility transition region having a second flexibility, and wherein flexibility transition regions overlap to define a region that blends the first flexibility with the second flexibility. 13.... | SciMed Life Systems, Inc. | A61B005/00 A61M025/00
US06882880 | B2 | 2005-04-19 | Arthroscopic impedance probe to detect cartilage degeneration | 2002-12-19 | The change in tissue impedance due to the change in the extracellular matrix that results from the degradation of cartilage is utilized to detect degradation of articular cartilage. A probe includes electrodes that apply a current to the articular cartilage which results in a current distribution and electric field within the cartilage, along with an associated voltage drop across the electrodes. The amplitude of this voltage drop is then measured and divided by the current applied to determine the tissue impedance. By measuring the impedance of patient tissue and comparing the detected patient impedance to a normal value for the tissue from clinically normal tissue, a determination of whether the patient tissue is degraded, and a determination of the extent of degradation is possible. Preferably, the impedance is measured using a probe with interdigitated electrodes. By changing which electrodes are utilized, the wavelength of the current distribution changes, allowing the probe to image depth dependent focal lesions. | 1. A system for detecting degeneration in mammalian cartilage by determining an impedance of the cartilage, said system comprising: a self contained surface probe having a plurality of excitation electrodes, said plurality of electrodes including a pair of excitation electrodes to provide contact with two surface points of the cartilage; a current source, operatively connected to said pair of excitation electrodes, to apply, through said pair of excitation electrodes, current to the cartilage; and a computing device, operatively connected to said pair of excitation electrodes, through which said current from said current source is applied to the cartilage, to measure an amplitude of a voltage drop between said pair of excitation electrodes, through which said current from said current source is applied to the cartilage, being induced by the application of said current and normalizing said measured amplitude to said applied current to determine said impedance of the cartilage. | 1. A system for detecting degeneration in mammalian cartilage by determining an impedance of the cartilage, said system comprising: a self contained surface probe having a plurality of excitation electrodes, said plurality of electrodes including a pair of excitation electrodes to provide contact with two surface points of the cartilage; a current source, operatively connected to said pair of excitation electrodes, to apply, through said pair of excitation electrodes, current to the cartilage; and a computing device, operatively connected to said pair of excitation electrodes, through which said current from said current source is applied to the cartilage, to measure an amplitude of a voltage drop between said pair of excitation electrodes, through which said current from said current source is applied to the cartilage, being induced by the application of said current and normalizing said measured amplitude to said applied current to determine said impedance of the cartilage. 2. The system for detecting degeneration in mammalian cartilage, as per claim 1, wherein said plurality of excitation electrodes comprises a plurality of selectably addressable interdigitated electrodes. 3. The system for detecting degeneration in mammalian cartilage, as per claim 1, wherein said excitation electrodes comprise a metal salt. 4. The system for detecting degeneration in mammalian cartilage, as per claim 3, wherein said excitation electrodes comprise silver chloride. 5. The system for detecting degeneration in mammalian cartilage, as per claim 1, said self contained surface probe further having a stress sensor for detecting a current generated stress induced in the cartilage by said applied current and transducing said current generated stress into one of a voltage or current, said stress sensor comprising a set of electrodes for transmitting said voltage or current to said computing device, said computing device performing processing on said received voltage or current to determine an electrokinetic parameter of the cartilage. 6. The system for detecting degeneration in mammalian cartilage, as per claim 1, said system further comprising: a waveform generator, said waveform generator operatively connected to said current source and providing said current source with a time varying voltage waveform, said time varying voltage waveform causing said current source to generate a time varying current for application to the cartilage via said excitation electrodes. 7. A system for detecting degeneration in mammalian cartilage by determining an impedance of the cartilage, said system comprising: a self contained surface probe having a set of excitation electrodes, said set of electrodes providing contact with two surface points of the cartilage; a current source, operatively connected to said electrodes, to apply, through said set of electrodes, current to the cartilage; a computing device, operatively connected to said electrodes, to measure an amplitude of a v... | Massachusetts Institute of Technology | A61B005/05
US06875183 | B2 | 2005-04-05 | Biopsy needle | 2001-09-19 | A biopsy needle for sampling bone marrow tissue, comprising a handle, and tissue sampling means comprising a sampling tube with a bore therein to receive a tissue sample, the outer surface of the sampling tube having an abrading formation extending in an axial direction along the tube, for abrading the sampled tissue, to permit the cannula tip to be laterally displaced whilst the tube is inserted into the bone marrow tissue. A biopsy needle for sampling bone marrow tissue, comprising a handle, tissue sampling means comprising a sampling tube with a bore therein to receive a tissue sample, and coupling means, separable from the needle, for coupling the needle to a rotary motor drive, whereby the needle is adaptable to both manual insertion and motor-assisted insertion. | 1. A biopsy needle instrument for sampling bone marrow tissue, said instrument comprising: a handle for inserting said needle into the bone marrow tissue; and said needle further comprising a single hollow tube, configured for both cutting and receiving a bone marrow tissue sample, coupled to said handle; said tube having (a) a tube bore defining a tissue-receiving space for the bone marrow tissue sample, (b) a substantially rigid tip, and (c) an outer wall configured to contact the bone marrow tissue, said outer wall being provided with an abrading formation comprising at least one slot cut into said outer wall, said slot having at least one sharp edge, said slot having a width of less than about 1 mm, said abrading formation extending in an axial direction along said tube to abrade the bone marrow tissue and thereby to allow said tip of said single hollow tube to be laterally displaced within the bone marrow tissue to facilitate retrieval of the bone marrow tissue sample. | 1. A biopsy needle instrument for sampling bone marrow tissue, said instrument comprising: a handle for inserting said needle into the bone marrow tissue; and said needle further comprising a single hollow tube, configured for both cutting and receiving a bone marrow tissue sample, coupled to said handle; said tube having (a) a tube bore defining a tissue-receiving space for the bone marrow tissue sample, (b) a substantially rigid tip, and (c) an outer wall configured to contact the bone marrow tissue, said outer wall being provided with an abrading formation comprising at least one slot cut into said outer wall, said slot having at least one sharp edge, said slot having a width of less than about 1 mm, said abrading formation extending in an axial direction along said tube to abrade the bone marrow tissue and thereby to allow said tip of said single hollow tube to be laterally displaced within the bone marrow tissue to facilitate retrieval of the bone marrow tissue sample. 2. A biopsy needle as claimed in claim 1, wherein two outer edges of said slot are defined by the intersection of said slot with said outer wall, and wherein both said outer edges are sharp. 3. A biopsy needle as claimed in claim 1, wherein said slot extends through said sampling tube between said outer wall and said bore. 4. A biopsy needle as claimed in claim 1, wherein said abrading formation comprises a plurality of said slots spaced circumferentially about said tube. 5. A biopsy needle as claimed in claim 4, wherein said slots of said abrading formation are spaced from said tip. 6. A biopsy needle as claimed in claim 1, wherein said slot extends, in an axial direction, at least 1 cm from said tip. 7. A biopsy needle as claimed in claim 1, wherein said tube further comprises a sample detacher at said tip to assist in detaching a base portion of said tissue sample from adjoining tissue. 8. A biopsy needle as claimed in claim 7, wherein said sample detacher comprises a sample detacher slot cut into a wall of said tip. 9. A biopsy needle as claimed in claim 8, wherein said sample detacher slot extends through said wall of said tip. 10. A biopsy needle as claimed in claim 8, wherein said sample detacher comprises a plurality of said sample detacher slots circumferentially about said tip. 11. A biopsy needle as claimed in claim 8, wherein said sample detacher slot has a length of between 1 mm and 2 mm. 12. A biopsy needle as claimed in claim 7, wherein at least a part of said sample detacher is disposed on said tube substantially opposite at least a part of said abrading formation. 13. A biopsy needle as claimed in claim 1 further comprising a stop to inhibit over-insertion of said tube into the bone marrow tissue being sampled. 14. A biopsy needle as claimed in claim 13, wherein at least a portion of said stop is integrally formed with at least a portion of said handle. 15. A biopsy needle as claimed in claim 1, wherein said tube bore extends through said handle... | - | A61B005/00 B65D081/00
US06896660 | B2 | 2005-05-24 | Therapeutic decisions systems and method using stochastic techniques | 2002-06-19 | Patients are evaluated in terms of various physiological parameters such as hemodynamic parameters. Covariate information indicative of the patient's individual characteristics may also be determined. A state variable indicative of the patient's characteristics is determined. This state variable is used to compare the patient to other patients in the database. The other patients in the database are evaluated in terms of similar parameters and their probability of survival as well as different treatments and the probability of survival based on those different treatments. The patient may be simply evaluated based on the database or a proper course of treatments may be determined based on the database. | 1. A system comprising: a monitoring device monitoring a plurality of physiological signals of a current patient; a database, which stores records indicative of a plurality of patients including at least some of said physiological signals for said plurality of patients and prognosis information for said plurality of patients; a processor, comparing said physiological signals to said stored records in a database, to determine mathematically patients in the database which are most similar to a current patient; and to make available information about treatment and results for said patients who are most similar to the current patient; wherein said physiological signals include hemodynamic information; and wherein said processor also carries out interpolation to place a plurality of said signals on a common time grid. | 1. A system comprising: a monitoring device monitoring a plurality of physiological signals of a current patient; a database, which stores records indicative of a plurality of patients including at least some of said physiological signals for said plurality of patients and prognosis information for said plurality of patients; a processor, comparing said physiological signals to said stored records in a database, to determine mathematically patients in the database which are most similar to a current patient; and to make available information about treatment and results for said patients who are most similar to the current patient; wherein said physiological signals include hemodynamic information; and wherein said processor also carries out interpolation to place a plurality of said signals on a common time grid. 2. A method, comprising: monitoring a plurality of physiological signals of a current patient; comparing said physiological signals of the current patient to stored physiological signals for previous patients stored in a database; determining patients in said database who are most similar to the current patient; determining information about treatments and results for said patients who are most similar to the current patient, wherein said physiological signals include signals indicative of hemodynamic information; obtaining covariate information indicative of a patient's individual characteristics, and wherein the steps of determining most similar patients and of determining information are also based on said covariate information; and wherein said determining most similar patients comprises mathematically determining a state variable indicative of the current patient based on at least said hemodynamic information and said covariate information. 3. A method, comprising: monitoring a plurality of physiological signals of a current patient; comparing said physiological signals of the current patient to stored physiological signals for previous patients stored in a database; determining patients in said database who are most similar to the current patient; determining information about treatments and results for said patients who are most similar to the current patient; wherein said physiological signals include signals indicative of hemodynamic information; obtaining covariate information indicative of a patient's individual characteristics, and wherein the steps of determining most similar patients and of determining information are also based on said covariate information; wherein said determining most similar patients comprises mathematically determining a state variable indicative of the current patient based on said hemodynamic information and said covasiate information; and wherein said determining a state variable comprises investigating derivatives of the hemodynamic information and integrals of the hemodynamic information. 4. A method, comprising: monitoring a plurality of physiological signals of a current patient; comparin... | University of Southern California | A61B005/02
US06847490 | B1 | 2005-01-25 | Optical probe accessory device for use in vivo diagnostic procedures | 2000-06-09 | The present invention recognizes that optical probes function both as medical access devices and as instruments which collect complex optical data. The invention provides an optical probe accessory device which can access luminal spaces within the body of a patient without sacrificing the quality of optical data obtained. The accessory device further comprises either, singly, or in combination, selectable features or options which optimize light transmission, maximize patient comfort, and provide single-use capabilities. | 1. A system for detecting features of a tissue sample, comprising: an optical probe; and an accessory device for attachment to the optical probe, wherein at least one of the probe and the accessory device includes an element for providing encoded information relating to at least one of the probe and the accessory device, wherein the element comprises a bar code for storing the encoded information. | 1. A system for detecting features of a tissue sample, comprising: an optical probe; and an accessory device for attachment to the optical probe, wherein at least one of the probe and the accessory device includes an element for providing encoded information relating to at least one of the probe and the accessory device, wherein the element comprises a bar code for storing the encoded information. 2. The system of claim 1, wherein the element further comprises an electrical element and the encoded information is stored therein. 3. The system of claim 1, further comprising an element reader for accessing the encoded information in the element. 4. The system of claim 1, wherein the encoded information includes identification information. 5. The system of claim 1, wherein the encoded information enables particular operating modes of the device. 6. The system of claim 2, wherein the electrical element is remotely programmable. 7. The system of claim 2, wherein the electrical element includes an RFID chip. 8. The system of claim 3, wherein the element reader further comprises a processor including a memory, and wherein the processor compares identification information encoded in the element to identification information located within the memory. 9. The system of claim 8, wherein the processor transmits instructions based on whether a match is found between the identification information encoded in the element and the identification information encoded in the memory. 10. The system of claim 9, wherein if no match is found, the identification information encoded in the element is added to the memory. 11. The system of claim 9, wherein the instructions include an instruction that permits the optical probe to function if no match is found. 12. The system of claim 9, wherein the instructions include an instruction that prevents the optical probe from functioning if no match is found. 13. The system of claim 8, wherein the processor controls transmission of light by the probe. 14. The system of claim 3, wherein the element reader is attached to the probe. 15. The system of claim 3, wherein the element reader is separate from the accessory device. 16. The system of claim 3, wherein the element reader is removably attached to the probe. 17. The system of claim 1, wherein the accessory device includes a flexible portion for conforming to a body space. 18. The system of claim 1 wherein the accessory device includes an integral lens. 19. The system of claim 1 wherein the accessory device includes a body and an attachment element for attaching the accessory device to the probe, the attachment element detaching from the body of the accessory device when the accessory device is removed from the probe, thereby preventing re-use of the accessory device. 20. A system for detecting features of a tissue sample, comprising: an optical probe; and an accessory device for attachment to the optical probe, wherein at least one of the probe and the accesso... | MediSpectra, Inc. | G02B003/00 A61B005/05 A61B006/00
US06896655 | B2 | 2005-05-24 | System and method for conditioning the psychological state of a subject using an adaptive autostereoscopic display | 2002-08-05 | An adaptive autostereoscopic display system (10) provides an apparatus for conditioning the psychological state, physiological state, or behavior of a subject (12) by displaying a stereoscopic virtual image at a left viewing pupil (14l) and a right viewing pupil (14r). A first set of images (100) is displayed and physiological response measurements are obtained from the subject (12). Based on the response of the subject (12) a personalized image response profile is obtained. Then, in order to condition the psychological state, physiological state, or behavior of the subject (12), a second set of images (102), based on the personalized image response profile is displayed. | 1. A system for conditioning the psychological state of a subject, the system comprising: (a) an image display for providing an autostereoscopic image to the subject; (b) at least one feedback sensor for providing a physiological measurement from the subject; (c) a control logic processor obtaining said physiological measurement from said at least one feedback sensor, and controlling the selection and processing of said autostereoscopic image by said image display based on said physiological measurement; and wherein said autostereoscopic image is selected from a personalized preferred image response profile. | 1. A system for conditioning the psychological state of a subject, the system comprising: (a) an image display for providing an autostereoscopic image to the subject; (b) at least one feedback sensor for providing a physiological measurement from the subject; (c) a control logic processor obtaining said physiological measurement from said at least one feedback sensor, and controlling the selection and processing of said autostereoscopic image by said image display based on said physiological measurement; and wherein said autostereoscopic image is selected from a personalized preferred image response profile. 2. A system for conditioning the psychological state of a subject according to claim 1 wherein said autostereoscopic image is a virtual image viewable at a right viewing pupil and a left viewing pupil. 3. A system for conditioning the psychological state of a subject according to claim 2 wherein said right viewing pupil is formed as the conjugate image of a ball lens. 4. A system for conditioning the psychological state of a subject according to claim 2 wherein said control logic processor modifies said left eye image and said right eye image in order to adjust the stereoscopic viewing relationship between said left viewing pupil and said right viewing pupil for displaying correct parallax information as the subject moves. 5. A system for conditioning the psychological state of a subject according to claim 1 wherein said response profile, is representatively selected to manage a psychological or physical state of said subject. 6. A system for conditioning the psychological state of a subject according to claim 1 wherein said physiological measurement is selected from a group consisting of EMG, EEG, galvanic skin response, skin temperature, heart rate, blood pressure, eye movement, head movement, and pupil dilation. 7. A system for conditioning the psychological state of a subject according to claim 1 wherein said image display forms said autostereoscopic image using a curved mirror. 8. A system for conditioning the psychological state of a subject according to claim 1 wherein said autostereoscopic image is a motion picture image. 9. A system for conditioning the psychological state of a subject according to claim 1 further comprising an image library accessed by said control logic processor. 10. A system for conditioning the psychological state of a subject according to claim 1 wherein said system dynamically modifies the display of said autostereoscopic image according to said physiological measurement. 11. A system for conditioning the psychological state of a subject according to claim 1 wherein said autostereoscopic image originates from film. 12. A system for conditioning the psychological state of a subject according to claim 1 wherein said autostereoscopic image originates from video. 13. A system for conditioning the psychological state of a subject according to claim 1 wherein said autostereoscopic image originates from ... | Eastman Kodak Company | A61B005/00 A61B003/02 A61F002/00
US06855115 | B2 | 2005-02-15 | Implantable wireless sensor for pressure measurement within the heart | 2002-01-22 | The progress of a endovascular cardiac repair can be monitored by inserting a pressure transducer sensor using a catheter into a chamber of the heart during endovascular repair and then using a small, hand-held read out device to measure pressure easily, safely, inexpensively and accurately. In one aspect a sensor is introduced into the body by the steps of folding or rolling the sensor into a cylinder, loading it into a catheter, and deploying into the heart chamber by allowing it to unroll or unfold, either by itself or facilitated by the incorporation of a super-elastic alloy component. | 1. A sensor delivery system comprising: a sensor comprising a self-contained resonant circuit comprising a capacitor and an inductor, wherein the circuit is variable in response to the physical property of the patient, and wherein the sensor is sufficiently flexible to be folded for delivery percutaneously; an outer catheter having at least one lumen and a distal open end, and an inner cylindrical member, wherein the sensor is folded within the distal end of the outer catheter and the inner cylindrical member pushes the folded sensor out at a desired location. | 1. A sensor delivery system comprising: a sensor comprising a self-contained resonant circuit comprising a capacitor and an inductor, wherein the circuit is variable in response to the physical property of the patient, and wherein the sensor is sufficiently flexible to be folded for delivery percutaneously; an outer catheter having at least one lumen and a distal open end, and an inner cylindrical member, wherein the sensor is folded within the distal end of the outer catheter and the inner cylindrical member pushes the folded sensor out at a desired location. 2. The delivery system of claim 1, wherein the inner catheter has a longitudinally extending lumen so that the delivery system can be slidably positioned over a guidewire. 3. The delivery system of claim 1, wherein the sensor has a safety wire attached thereto and said safety wire extends proximally in a longitudinally extending groove or lumen. 4. A flexible sensor for wirelessly determining a physical property in a patient's heart chamber, which sensor comprises a self-contained resonant circuit comprising a capacitor and an inductor, wherein the circuit is variable in response to the physical property of the patient, wherein the sensor is sufficiently flexible to be folded for delivery percutaneously, and wherein a safety wire is attached to one surface of the sensor. 5. The sensor of claim 4, wherein the safety wire has a sheath. 6. The sensor of claim 5, wherein the sheath can be slid distally to free the safety wire from the sensor. 7. The sensor of claim 4, wherein the safety wire is attached to the sensor at an adhesive point. 8. The sensor of claim 7, wherein the adhesive point comprises an epoxy or a cyanoacrylate material. | Cardiomems, Inc. | A61B005/0215
US06896874 | B2 | 2005-05-24 | MR-signal emitting coatings | 2002-05-09 | The present invention provides a coating that emits magnetic resonance signals and a method for coating medical devices therewith. The coating includes a paramagnetic metal ion-containing polymer complex that facilitates diagnostic and therapeutic techniques by readily visualizing medical devices coated with the complex. The present invention also provides methods by which pre-existing polymers and medical devices may be made MR-imageable. The invention also provides methods of improving MR-imageability of polymers and medical devices by encapsulating the polymers and medical devices with hydrogels. | 1. A method of making a medical device magnetic-resonance imageable, the method comprising: providing a coating on the medical device in which a paramagnetic-metal ion/chelate complex is encapsulated by a first hydrogel, a chelate of the paramagnetic-metal-ion/chelate complex being linked to a functional group, and the functional group being an amine group or a carboxyl group. | 1. A method of making a medical device magnetic-resonance imageable, the method comprising: providing a coating on the medical device in which a paramagnetic-metal ion/chelate complex is encapsulated by a first hydrogel, a chelate of the paramagnetic-metal-ion/chelate complex being linked to a functional group, and the functional group being an amine group or a carboxyl group. 2. The method of claim 1, wherein at least a portion of the medical device is made from a solid-base polymer, and the method further comprises treating the solid-base polymer to yield the functional group thereon, the complex accordingly being covalently linked to the medical device. 3. The method of claim 2, wherein treating the solid-base polymer comprises plasma treating the solid-base polymer with a plasma gas which is hydrazine, ammonia, a chemical moiety of a nitrogen-hydrogen combination or combinations thereof, and wherein the resulting plasma-treated functional group is an amine group. 4. The method of claim 2, wherein treating the solid-base polymer comprises plasma treating the solid-base polymer with a plasma gas which is carbon dioxide or oxygen, and wherein the resulting plasma treated functional group is a carboxyl group. 5. The method of claim 2, wherein treating the solid-base polymer comprises melt coating with a hydrophilic polymer or precoating with a hydrophilic polymer containing primary amine groups. 6. The method of claim 2, wherein the chelate is covalently linked to the functional group by an amide linkage. 7. The method of claim 2, wherein the polymer is selected from the group consisting of polyethylenes, polypropylenes, polyesters, polyamides, polytetrafluoroethylene, polyurethanes, polyamino undecanoic acid, polydimethylsiloxane, polyglycols, polyoxyethylenes, polysorbate 60, stearate and palmitate esters of sorbitol copolymerized with ethylene glycol, polyvinyl acetate phthalate, polyvinyl alcohol and polystyrene sulfonate. 8. The method of claim 2, wherein the paramagnetic-metal ion is designated as Mn+, and M is a lanthanide or a transition metal which is iron, manganese, chromium, cobalt or nickel, and n is an integer that is 2 or greater. 9. The method of claim 8, wherein M is a lanthanide and the lanthanide is gadolinium. 10. The method of claim 2, wherein the chelate is selected from the group consisting of diethylenetriaminepentaacetic acid (DTPA), 1,4,7,10-tetracyclododecane-N,N′,N″,N′″-tetraacetic acid (DOTA) and 1,4,8,11-tetraazacyclotradecane-N,N′,N″,N′″-tetraacetic acid (TETA), diethylenetriaminepentaacetic acid-N,N′-bis(methylamide) (DTPA-BMA), diethylenetriaminepentaacetic acid-N,N′-bis(methoxyethylamide) (DTPA-BMEA), s-4-(4-ethoxybenzyl)-3,6,9-tris[(carboxylatomethyl)]-3,6,9-triazaundecanedionic acid (EOB-DTPA), benzyloxypropionictetraacetate(BOPTA), (4R)-4-[bis(carboxymethylamino]-3,6,9-triazaundecanedionic acid (MS-325), 1,4,7-tris(carboxymethyl)-10-(2′-hydroxypropyl)-1,4,7,10-tetraazacyclododecane (HP-DO3A), and DO3... | Wisconsin Alumni Research Foundation | A61B005/055
US06962568 | B1 | 2005-11-08 | Measuring device for determining an uneven loading of the human body in the upright position | 2000-03-05 | The present invention relates to a measuring device for determining an uneven loading of the human body in the upright position, comprising a base element with two balances, and an adjustment unit for positioning individual points or regions of the body of the person to be measured. The adjustment unit comprises two vertical struts fixed to a base element, on which measuring and fixing elements are arranged in a vertically displaceable manner. A first measuring and fixing element is provided for measuring and fixing the knee with a transverse edge for the sides of the feet, together with a second measuring and fixing element for measuring and fixing the hips, and a third measuring and fixing agent element for measuring and fixing the shoulder region, are provided. | 1. A measuring device for determining an incorrect posture and/or uneven loading of a human body in an upright posture, comprising: (a) a base element comprising two weight balances, the base element being provided with at least one abutment for positioning heels of a person to be measured, wherein the at least one abutment defines a rear frontal plane; and(b) an adjustment unit for positioning individual body points or regions of a person to be measured, the adjustment unit comprising: (i) at least two vertical struts which define a positioning axis for the definition of a middle body line of the person to be measured; and(ii) a first vertically displaceable fixation element and a second vertically displaceable fixation element positioned on the at least two vertical struts for fixing a hip region and a shoulder region, wherein each of said first and second fixation elements comprises a first and a second displaceable transverse strut and a first and a second sagittal rod displaceably attached to said first and second transverse struts, and wherein a first and a second pivotable and displaceable measuring bar are provided on said first and second saggital rods,whereby the position of a body point to be measured may be determined in a transverse plane defined by the first and second transverse struts and the first and second sagittal rods. | 1. A measuring device for determining an incorrect posture and/or uneven loading of a human body in an upright posture, comprising: (a) a base element comprising two weight balances, the base element being provided with at least one abutment for positioning heels of a person to be measured, wherein the at least one abutment defines a rear frontal plane; and(b) an adjustment unit for positioning individual body points or regions of a person to be measured, the adjustment unit comprising: (i) at least two vertical struts which define a positioning axis for the definition of a middle body line of the person to be measured; and(ii) a first vertically displaceable fixation element and a second vertically displaceable fixation element positioned on the at least two vertical struts for fixing a hip region and a shoulder region, wherein each of said first and second fixation elements comprises a first and a second displaceable transverse strut and a first and a second sagittal rod displaceably attached to said first and second transverse struts, and wherein a first and a second pivotable and displaceable measuring bar are provided on said first and second saggital rods,whereby the position of a body point to be measured may be determined in a transverse plane defined by the first and second transverse struts and the first and second sagittal rods. 2. A measuring device according to claim 1, wherein the adjustment unit further comprises a third vertically displaceable fixation element positioned on the at least two vertical struts for fixing and measuring a knee region, the third fixation element comprising at least one transverse strut, at least one sagittal rod displaceably fastened attached to said transverse strut, a frontal rod attached to said sagittal rod, and at least one displaceable measuring bar attached to the frontal rod in a pivotably movable manner. 3. A measuring device according to any one of claims 1 and 2, wherein each of said transverse struts is held displaceably on said vertical struts by a base block and each of said sagittal rods is held displaceably on said transverse struts by an angle element. 4. A measuring device according to claim 1, wherein each of said measuring bars is provided with at least one terminal bore for receiving the saggital rod. 5. A measuring device according to claim 2, wherein the at least one measuring bar is provided with a terminal bore for receiving the saggital rod. 6. A measuring device according to any one of claims 4 and 5, wherein the at least one terminal bore is surrounded by an angle scale, and wherein each of said saggital rods is provided with a reference line. 7. A measuring device according to claim 2, wherein a length scale is provided on at least one of the vertical struts, at least one of the saggital rods and the frontal rod. 8. A measuring device according to claim 7, wherein the length scale may be read visually. 9. A measuring device according to claim 7, wherein the length sca... | - | A61B005/103 A61B005/117
US06950689 | B1 | 2005-09-27 | Dynamically alterable three-dimensional graphical model of a body region | 1998-08-03 | The present invention is a system and method for graphically displaying a three-dimensional model of a region located within a living body. A three-dimensional model of a region of interest is displayed on a graphical display. The location in three-dimensional space of a physical characteristic (e.g. a structure, wall or space) in the region of interest is determined using at least one probe positioned within the living body. The graphical display of the model is deformed to approximately reflect the determined three-dimensional location of the physical characteristic. Preferably, the probe or probes are moved throughout the region of interest so as to gather multiple data points that can be used to increase the conformity between the graphical display and the actual region of interest within the patient. | 1. A method of generating a three-dimensional graphical model of a region located within a living body, comprising the steps of: (a) generating a three-dimensional model of a region of interest; (b) determining a three-dimensional location of a physical characteristic in a portion of the region of interest using at least one roving probe positioned within the living body; (c) deforming the model to at least approximately incorporate the physical characteristic at the determined three-dimensional location of the physical characteristic; (d) displaying the model on a graphical display; and (e) repeating steps (b) through (d) multiple times for different roving probe locations to increase the conformity of the model to the region of interest. | 1. A method of generating a three-dimensional graphical model of a region located within a living body, comprising the steps of: (a) generating a three-dimensional model of a region of interest; (b) determining a three-dimensional location of a physical characteristic in a portion of the region of interest using at least one roving probe positioned within the living body; (c) deforming the model to at least approximately incorporate the physical characteristic at the determined three-dimensional location of the physical characteristic; (d) displaying the model on a graphical display; and (e) repeating steps (b) through (d) multiple times for different roving probe locations to increase the conformity of the model to the region of interest. 2. The method of claim 1 further including the step of (e) determining a relative three-dimensional location of a medical device positioned in the region of interest and graphically representing the medical device superimposed on the model at the determined three-dimensional location of the medical device. 3. The method of claim 2 wherein the medical device is the at least one probe. 4. The method of claim 2 wherein: the medical device includes at least one mapping electrode; step (e) includes determining a relative three-dimensional location of the at least one mapping electrode; and the method further includes the steps of: (f) detecting electrical activity using the at least one mapping electrode; and (g) generating a map of the detected electrical activity and superimposing the map on the model at a three-dimensional location corresponding to the determined three-dimensional location of the at least one mapping electrode. 5. The method of claim 2 wherein: the medical device includes at least one ablation electrode; step (e) includes determining a relative three-dimensional location of the at least one ablation electrode; and the method further includes the steps of: (f) forming a lesion using the at least one ablation electrode; and (g) generating a visual representation of the lesion and superimposing the visual representation on the model at a three-dimensional location corresponding to the determined three-dimensional location of the at least one ablation electrode. 6. The method of claim 1 wherein the model generated in step (a) includes an anatomical feature known to exist in the region of interest, wherein the physical characteristic is the anatomical feature. 7. The method of claim 6 wherein step (c) includes scaling the model to correlate a corresponding anatomical feature on the model with the determined three-dimensional location of the anatomical feature. 8. The method of claim 6 wherein step (c) includes orienting the model to correlate a corresponding anatomical feature on the model with the determined three-dimensional location of the anatomical feature. 9. The method of claim 6 wherein step (c) includes performing a rigid body transformation. 10. The method of claim 9 wherein the ... | Boston Scientific Scimed, Inc. | A61B005/00
US06905464 | B2 | 2005-06-14 | Visceral fat meter provided with tonometer | 2002-02-01 | A visceral fat scale equipped with a sphygmomanometer is described, with which a subject can keep track of blood pressure values and grasp the state of obesity to attain more accurate, comprehensive and diversified diagnoses and disease prevention. The visceral fat scale is designed such that after exerting pressure to an artery of a subject by inflation of a cuff, a pulse wave signal is detected in the course of gradual cuff deflation and blood pressure is obtained based on the detected pulse wave signal. The scale comprises an increment/decrement key and display mode selector key for inputting personal data of the subject; an arithmetic operation unit for calculating the amount of visceral fat of the subject based on data input by these keys; and a display unit for displaying a result of the calculation performed by the arithmetic operation unit. | 1. A visceral fat scale equipped with a sphygmomanometer comprising a cuff, wherein after exerting the cuff exerts pressure to an artery of a subject by inflation of the cuff, the sphygmomanometer obtains the blood pressure of the subject based on a detected pulse wave signal, and the visceral fat scale comprising: an input device for inputting personal data of the subject; an arithmetic operation unit for calculating the amount of visceral fat of the subject based on the data input by the input device, the amount of visceral fat being the cross-sectional area of abdominal visceral fat of the subject; and a display unit for displaying a result of the calculation performed by the arithmetic operation unit. | 1. A visceral fat scale equipped with a sphygmomanometer comprising a cuff, wherein after exerting the cuff exerts pressure to an artery of a subject by inflation of the cuff, the sphygmomanometer obtains the blood pressure of the subject based on a detected pulse wave signal, and the visceral fat scale comprising: an input device for inputting personal data of the subject; an arithmetic operation unit for calculating the amount of visceral fat of the subject based on the data input by the input device, the amount of visceral fat being the cross-sectional area of abdominal visceral fat of the subject; and a display unit for displaying a result of the calculation performed by the arithmetic operation unit. 2. The visceral fat scale equipped with a sphygmomanometer according to claim 1, wherein the personal data input by the input device includes the height, weight and waist size of the subject and wherein the arithmetic operation unit calculates the body mass index (BMI) of the subject in addition to the amount of visceral fat of the subject based on the data input by the input device. 3. The visceral fat scale equipped with a sphygmomanometer according to claim 2, wherein the display unit displays the respective ranks of the visceral fat amount and BMI of the subject, each of said indices having a plurality of ranks. 4. The visceral fat scale equipped with a sphygmomanometer according to claim 2 or 3, wherein the arithmetic operation unit makes an assessment of obesity by comparing the amount of visceral fat and BMI of the subject with their respective reference values for assessment which have been input beforehand, and wherein the display unit displays a result of the assessment of obesity. 5. The visceral fat scale equipped with a sphygmomanometer according to claim 1, wherein the display unit displays blood pressure values and the amount of visceral fat at the same time. 6. The visceral fat scale equipped with a sphygmomanometer according to claim 1, wherein the arithmetic operation unit calculates the amount of change from a previous measurement result and the ratio of the present measurement result to the previous measurement result for blood pressure values and visceral fat amount of the subject, and wherein the display unit displays a result of the amount of change and ratio calculation. 7. The visceral fat scale equipped with a sphygmomanometer according to claim 1, wherein a target inflation value for the cuff is determined by a systolic blood pressure estimated based on a pulse wave generated during cuff inflation and by a result of the assessment of obesity. 8. The visceral fat scale equipped with a sphygmomanometer according to claim 1, wherein a target inflation value for the cuff is determined by a systolic blood pressure estimated based on a pulse wave generated during cuff inflation and by the cross-sectional area of abdominal visceral fat. 9. A visceral fat scale equipped with a sphygmomanometer comprising a cuff, where... | Yamato Scale Co., Ltd. | A61B005/00
US06839583 | B1 | 2005-01-04 | Disposable tissue probe tip | 2000-06-01 | A disposable tip adapted for use in connection with an optical probe of an instrument. The tip includes a one-piece molded elastomeric base member having a boot and an opaque tissue-engaging surface. The boot releasably mates with and provides an interference/friction fit to the probe. The tissue-engaging surface surrounds and extends from the boot. A window in the base member extends into the boot to transmit light between the probe and tissue. A layer of optically transparent material covers the window on the tissue-engaging surface. Adhesive on the tissue-engaging surface of the base member releasably secures the probe to the tissue being analyzed. | 1. A disposable tip adapted for use in connection with an optical probe of an instrument, including: a one piece elastomeric base member having an extended tissue-engaging surface and structure for releasably engaging the optical probe of the instrument, the structure for releasably engaging includes a cavity for engaging a projection on the probe, the cavity having outwardly sloping inner walls with increasing distance toward the tissue-facing surface for mating with the projection on the probe; and a window in the base member for transmitting light between the probe and a tissue. | 1. A disposable tip adapted for use in connection with an optical probe of an instrument, including: a one piece elastomeric base member having an extended tissue-engaging surface and structure for releasably engaging the optical probe of the instrument, the structure for releasably engaging includes a cavity for engaging a projection on the probe, the cavity having outwardly sloping inner walls with increasing distance toward the tissue-facing surface for mating with the projection on the probe; and a window in the base member for transmitting light between the probe and a tissue. 2. The tip of claim 1 and further including adhesive on the tissue-enganging surface of the base member for releaseably securing the tip to a patient. 3. The tip of claim 2 and further including a release liner covering the adhesive. 4. The tip of claim 1 wherein the cavity surrounds the window. 5. The tip of claim 1 wherein the structure for engaging the optical probe surrounds the window. 6. The tip of claim 1 wherein the window includes an optically transparent member mounted to the base member. 7. The tip of claim 1 and further including adhesive on an optically clear substrate which covers the window and tissue-engaging surface of the base member. 8. The tip of claim 1 wherein the tip and window are a one-piece member. 9. The tip of claim 1 wherein the structure for releasably engaging the probe provides an interference/friction fit to the probe. 10. The tip of claim 1 wherein the tissue-engaging surface of the base member is optically opaque. 11. A disposable tip adapted for use in connection with an optical probe of an instrument, including: a one-piece molded elastomeric base member including: a boot having a cavity with outwardly sloping interior walls with increasing distance toward the tissue-facing surface for releasably mating with and providing an interference/friction fit to the probe; and a tissue-engaging surface surrounding and extending from the boot; a window in the base member extending into the cavity of the boot, for transmitting light between the probe and tissue; a layer of optically transparent material over the window on the tissue-engaging surface; and adhesive on the tissue-engaging surface of the base member, for releasably securing the probe to tissue being analyzed. 12. The disposable tip of claim 11 wherein the tissue-engaging surface is optically opaque. 13. The disposable tip of claim 12 and further including a release liner covering the adhesive. 14. The disposable tip of claim 11 wherein at least a portion of the optically transparent window extends to an edge of the base member. | Hutchinson Technology Corporation | A61B005/00
US07207952 | B2 | 2007-04-24 | Body fluid composition measuring apparatus | 2001-07-26 | A body fluid component measuring apparatus capable of accurately measuring a specific component in a body fluid for a short time includes, in a main body, a measuring device for detecting the sampling of a body fluid and measuring a component of the sampled body fluid, and a pump and an electromagnetic valve which constitute an evacuating mechanism. When the sampling of blood is detected, the pump is stopped to release the evacuation state. Another apparatus according to a second aspect includes a pressure detecting means and a notifying device. A puncturing device is operated only when it is decided on the basis of a result of detection by a sensor, that the housing is in an evacuation state, and after the sampling is detected, the evacuation is released and information is notified. A further apparatus according to a third aspect includes a pressure adjusting means for fluctuating the pressure. | 1. A combination of a body fluid component measuring apparatus and a chip adapted to be mounted thereto, said chip comprising a chip housing, a puncture needle and a strip of test paper disposed on said chip housing, said chip housing possessing a contact portion at a tip side thereof, and said contact portion internally possessing a tip opening, said body fluid component measuring apparatus comprising: a touch portion to which an epidermis to be punctured is touched;a housing space for holding said chip;puncturing means for operating said puncture needle when said chip is mounted to said apparatus so as to puncture the epidermis touched to said touch portion;evacuating means for evacuating a puncture site of the epidermis to be punctured by said puncture needle as well as the housing space;measuring means for measuring an amount of a specific component of a body fluid sampled from the puncture site and diffused in said test paper;body fluid sampling detecting means for detecting the sampling of the body fluid; andevacuation releasing means for releasing or relieving the evacuation state of at least said housing space;wherein when the sampling of the body fluid is detected by said body fluid sampling detecting means, the evacuation state of at least said housing space is released or relieved, and then the amount of the specific component in the sampled body fluid is measured by said measuring means. | 1. A combination of a body fluid component measuring apparatus and a chip adapted to be mounted thereto, said chip comprising a chip housing, a puncture needle and a strip of test paper disposed on said chip housing, said chip housing possessing a contact portion at a tip side thereof, and said contact portion internally possessing a tip opening, said body fluid component measuring apparatus comprising: a touch portion to which an epidermis to be punctured is touched;a housing space for holding said chip;puncturing means for operating said puncture needle when said chip is mounted to said apparatus so as to puncture the epidermis touched to said touch portion;evacuating means for evacuating a puncture site of the epidermis to be punctured by said puncture needle as well as the housing space;measuring means for measuring an amount of a specific component of a body fluid sampled from the puncture site and diffused in said test paper;body fluid sampling detecting means for detecting the sampling of the body fluid; andevacuation releasing means for releasing or relieving the evacuation state of at least said housing space;wherein when the sampling of the body fluid is detected by said body fluid sampling detecting means, the evacuation state of at least said housing space is released or relieved, and then the amount of the specific component in the sampled body fluid is measured by said measuring means. 2. The combination of the body fluid component measuring apparatus and the chip according to claim 1, wherein the measurement of the amount of the specific component in the body fluid requires a specific component in atmospheric air. 3. The combination of the body fluid component measuring apparatus and the chip according to claim 1, wherein said evacuation releasing means has a flow passage for communicating said housing space to the outside and a valve for opening/closing said flow passage. 4. The combination of the body fluid component measuring apparatus and the chip according to claim 1, wherein said evacuation releasing means has a flow passage for communicating said housing space to the outside, and at least part of said flow passage has a portion in which an air passing resistance is relatively high. 5. The combination of the body fluid component measuring apparatus and the chip according to claim 1, wherein a component forming at least part of said measuring means also forms at least part of said body fluid sampling detecting means. 6. The combination of the body fluid component measuring apparatus and the chip according to claim 1, the body fluid component measuring apparatus further comprising an apparatus housing which houses said puncturing means, said housing space being provided at said apparatus housing wherein said evacuating means brings said housing space in said housing into an evacuation state. 7. The combination of the body fluid component measuring apparatus and the chip according to claim 6, wherein said evacuation releas... | Terumo Kabushiki Kaisha | A61B005/00 B65D081/00
US06996429 | B2 | 2006-02-07 | Apparatus and method for detecting flinch response to nociceptive agents | 2001-01-17 | An apparatus and method for measuring spatial displacement of an animal's paw injected with irritant. This apparatus comprises an electromagnetic detecting assembly having (1) a transmitting oscillator for generating electrical current (2) an electromagnetic transmitter coil coupled to the oscillator for generating an electromagnetic field (3) an electromagnetic receiving coil placed in axial plane directly below the transmitter coil (4) a first, receiving amplifier connected to the receiving coil (5) an amplitude detector connected to the receiving amplifier (6) a second amplifier connected to the amplitude detector (7) a metal object attached to the animal's paw (8) and a cylindrical observation chamber of a diameter not greater than the diameter of the generated magnetic field. | 1. An automated flinch-detection apparatus for measuring spatial displacement of an animal's paw injected with irritant, comprising: an electromagnetic detecting assembly having a transmitting oscillator for generating electrical current;an electromagnetic transmitter coil coupled to the oscillator for generating an electromagnetic field;an electromagnetic receiving coil placed in a linear plane directly below the transmitter coil; a first, receiving amplifier connected to the receiving coil;an amplitude detector connected to the receiving amplifier;a second amplifier connected to the amplitude detector;a metal object adapted to be attached to the animal's paw; anda cylindrical observation chamber of a diameter not greater than the diameter of the generated magnetic field, said chamber placed directly over the receiving/transmitting coil assembly, wherein the current generated by the transmitting oscillator circulates in the transmitter coil, creating an electromagnetic field that penetrates the metal object, creating eddy currents perturbing the electromagnetic field, said fluctuating perturbations being picked up by the receiving coil, amplified by the receiving amplifier, detected by the amplitude detector and further amplified, filtered and digitized to produce a measured spatial displacement of the animal's paw injected with the irritant. | 1. An automated flinch-detection apparatus for measuring spatial displacement of an animal's paw injected with irritant, comprising: an electromagnetic detecting assembly having a transmitting oscillator for generating electrical current;an electromagnetic transmitter coil coupled to the oscillator for generating an electromagnetic field;an electromagnetic receiving coil placed in a linear plane directly below the transmitter coil; a first, receiving amplifier connected to the receiving coil;an amplitude detector connected to the receiving amplifier;a second amplifier connected to the amplitude detector;a metal object adapted to be attached to the animal's paw; anda cylindrical observation chamber of a diameter not greater than the diameter of the generated magnetic field, said chamber placed directly over the receiving/transmitting coil assembly, wherein the current generated by the transmitting oscillator circulates in the transmitter coil, creating an electromagnetic field that penetrates the metal object, creating eddy currents perturbing the electromagnetic field, said fluctuating perturbations being picked up by the receiving coil, amplified by the receiving amplifier, detected by the amplitude detector and further amplified, filtered and digitized to produce a measured spatial displacement of the animal's paw injected with the irritant. 2. The apparatus according to claim 1, wherein the current passing through the transmitter coil generates an electromagnetic field in the 6 to 8 kilohertz range with a signal strength on the order of 5 to 8 milliwatts. 3. The apparatus according to claim 1, wherein the metal object is a small metal annular collar. 4. The apparatus according to claim 1, wherein the metal object is a small metal “C” collar in incomplete annular form. 5. The apparatus according to claim 1, wherein the metal object comprises a ferrous metal. 6. The apparatus according to claim 1, wherein the metal object comprises a non-ferrous metal. 7. The apparatus according to claim 1, wherein the observation chamber is a transparent cylindrical container, insuring that the animal will remain inside the boundaries of the electromagnetic field generated by the coil assembly. 8. The apparatus according to claim 1, wherein the observation chamber has individual compartments permitting testing of a plurality of animals. 9. The apparatus according to claim 1, wherein the detection assembly below the observation chamber has multiple independent detection units. 10. The apparatus according to claim 1, wherein the observation chamber is constructed of any rigid transparent plastic. 11. A method for measuring a flinch response by an animal whose paw has been subjected to an irritant, comprising: attaching a metal object to the animal's paw;placing the animal in an observation chamber situated directly over a detection assembly having a transmitting oscillator for generating electrical current,electromagnetic transmitter coil coupled to the... | Regents of the Univeristy of California, San Diego | A61B005/05
US07076283 | B2 | 2006-07-11 | Device for sensing cardiac activity in an implantable medical device in the presence of magnetic resonance imaging interference | 2001-10-31 | A method is provided, the method comprising detecting a magnetic resonance imaging (MRI) interference signal and enabling at least one preventive measure to protect an implantable medical device from interference by the magnetic resonance imaging (MRI) interference signal. The method also comprises switching from a first sensing mode more affected by the magnetic resonance imaging (MRI) interference signal to a second sensing mode less affected by the magnetic resonance imaging (MRI) interference signal. | 1. A device adapted to perform a cardiac sensing-mode switch so to sense cardiac activity in the presence of magnetic resonance imaging (MRI) interference that exceeds a threshold of about 0.20 Tesla but not to perform said mode switch in the presence of electromagnetic interference (EMI) that does not exceed a threshold of about 0.20 Tesla, comprising: means for detecting a magnetic field consistent with the characteristics of an MRI scanning device and providing an MRI interference signal related to the detection of the magnetic field; andswitching means coupled to the means for detecting, for switching from a first cardiac activity sensing mode that is relatively more affected by the MRI interference signal to a second cardiac activity sensing mode that is relatively less affected by the magnetic field in response to receipt of the MRI interference signal, wherein in the event that a detected EMI field strength falls below about 0.20 Tesla then no switching of the first cardiac activity sensing mode occurs;wherein the second cardiac activity sensing mode employs at least one of the following cardiac activity sensing circuitry disposed within an implantable medical device; a can-based accelerometer, a pressure sensor on a lead, an accelerometer on a lead, an accelerometer coupled to a connector block a flow sensor, a heart motion sensor based on time-of-flight, a temperature sensor, an impedance-based sensor, an oxygen sensor. | 1. A device adapted to perform a cardiac sensing-mode switch so to sense cardiac activity in the presence of magnetic resonance imaging (MRI) interference that exceeds a threshold of about 0.20 Tesla but not to perform said mode switch in the presence of electromagnetic interference (EMI) that does not exceed a threshold of about 0.20 Tesla, comprising: means for detecting a magnetic field consistent with the characteristics of an MRI scanning device and providing an MRI interference signal related to the detection of the magnetic field; andswitching means coupled to the means for detecting, for switching from a first cardiac activity sensing mode that is relatively more affected by the MRI interference signal to a second cardiac activity sensing mode that is relatively less affected by the magnetic field in response to receipt of the MRI interference signal, wherein in the event that a detected EMI field strength falls below about 0.20 Tesla then no switching of the first cardiac activity sensing mode occurs;wherein the second cardiac activity sensing mode employs at least one of the following cardiac activity sensing circuitry disposed within an implantable medical device; a can-based accelerometer, a pressure sensor on a lead, an accelerometer on a lead, an accelerometer coupled to a connector block a flow sensor, a heart motion sensor based on time-of-flight, a temperature sensor, an impedance-based sensor, an oxygen sensor. 2. The device of claim 1, wherein the means for detecting is capable of detecting the magnetic field by detecting a high magnetic field having a magnetic field strength above a predetermined threshold other than about 0.17 Tesla. 3. The device of claim 2, wherein the predetermined threshold is about 0.20 Tesla. 4. The device of claim 2, wherein the means for detecting is capable of detecting the magnetic field by detecting the high magnetic field using a Hall Effect sensor in communication with the implantable medical device. 5. The device of claim 4, further comprising means for opening a case switch for the implantable medical device in response to receipt of the MRI interference signal. 6. The device of claim 5, further comprising means for electrically separating one or more leads for the implantable medical device from a portion of a housing for the implantable medical device in response to receipt of the MRI interference signal. 7. A device according to claim 1, wherein means for detecting the magnetic field comprises detecting a high magnetic field having a magnetic field strength of about 0.2 Tesla (2000 Gauss) to about 10 Tesla (100,000 Gauss). 8. A device according to claim 1, wherein the means for detecting the magnetic field comprises detecting a high magnetic field having one of: a static gradient magnetic field, a variable gradient magnetic field with a frequency of about 5 KHz, a radio-frequency pulses of up about 10 MHz to about 50 MHz, a variable magnetic field having a frequency of about 64 Hz. | Medtronic, Inc. | A61B005/05
US06998841 | B1 | 2006-02-14 | Method and system which forms an isotropic, high-resolution, three-dimensional diagnostic image of a subject from two-dimensional image data scans | 2000-03-31 | MRI scans typically have higher resolution within a slice than between slices. To improve the resolution, two MRI scans are taken in different, preferably orthogonal, directions. The scans are registered by maximizing a correlation between their gradients and then fused to form a high-resolution image. Multiple receiving coils can be used. When the images are multispectral, the number of spectral bands is reduced by transformation of the spectral bands in order of image contrast and using the transformed spectral bands with the highest contrast. | 1. A method of forming an isotropic, high-resolution, three-dimensional diagnostic image of a subject from two-dimensional image data, the method comprising: (a) scanning the subject with an imaging scanner device, in a first direction relative to the subject, in order to take image data of a first plurality of slices, the image data of the first plurality of slices being two-dimensional image data and having a low resolution in the first direction and a high resolution in directions orthogonal to the first direction; (b) scanning the subject with said imaging scanner device in a second direction relative to the subject, which is different from the first direction in order to take image data of a second plurality of slices, the image data of the second plurality of slices being two-dimensional image data having at least one dimension substantially in common with the image data of the first plurality of slices and having a low resolution in the second direction and a high resolution in directions orthogonal to the second direction; (c) registering the first plurality of slices with the second plurality of slices to define a matrix of isotropic, high-resolution voxels in image space, wherein the matrix has unknown high-resolution voxel values; and (d) solving for the unknown high-resolution voxel values in the matrix defined in step (c) in accordance with the image data taken in steps (a) and (b) in order to form the isotropic high-resolution three-dimensional diagnostic image in the image space. | 1. A method of forming an isotropic, high-resolution, three-dimensional diagnostic image of a subject from two-dimensional image data, the method comprising: (a) scanning the subject with an imaging scanner device, in a first direction relative to the subject, in order to take image data of a first plurality of slices, the image data of the first plurality of slices being two-dimensional image data and having a low resolution in the first direction and a high resolution in directions orthogonal to the first direction; (b) scanning the subject with said imaging scanner device in a second direction relative to the subject, which is different from the first direction in order to take image data of a second plurality of slices, the image data of the second plurality of slices being two-dimensional image data having at least one dimension substantially in common with the image data of the first plurality of slices and having a low resolution in the second direction and a high resolution in directions orthogonal to the second direction; (c) registering the first plurality of slices with the second plurality of slices to define a matrix of isotropic, high-resolution voxels in image space, wherein the matrix has unknown high-resolution voxel values; and (d) solving for the unknown high-resolution voxel values in the matrix defined in step (c) in accordance with the image data taken in steps (a) and (b) in order to form the isotropic high-resolution three-dimensional diagnostic image in the image space. 2. The method of claim 1, wherein the second direction is orthogonal to the first direction. 3. The method of claim 1, wherein step (d) comprises treating the image as a linear combination of at least two low-resolution functions and deriving the unknown high-resolution voxel values from the image data of the first and second pluralities of slices. 4. The method of claim 3, wherein the functions are derived through an iterative process using the image data of the first and second pluralities of slices as initial assumptions for the functions. 5. The method of claim 1, wherein step (c) comprises maximizing a correlation based on the directionally scanned image data of the first and second pluralities of slices. 6. The method of claim 5, wherein the correlation is maximized through a hill-climbing technique in which successively larger values of the correlation are located until a local maximum value of the correlation is reached. 7. The method of claim 6, wherein the hill-climbing technique is a multiresolution hill-climbing technique. 8. The method of claim 7, wherein: steps (a) and (b) are performed with an imaging scanner device having an in-plane resolution; and the multiresolution hill-climbing technique is performed with a plurality of resolutions including: (i) a maximum resolution which is twice the in-plane resolution; and (ii) a minimum resolution which is one-quarter of the in-plane resolution. 9. The method of claim 6, wherein the hill... | Virtualscopics, LLC | G01R033/20 G01V003/00 A61B005/055 G06K009/00
US07089051 | B2 | 2006-08-08 | Implantable medical device with valve opening detector | 2001-06-19 | A valve opening detector for use in an implantable medical device detects opening of at least one valve in the left side of heart, and has an impedance measuring unit for measuring electrical impedance between measuring poles and for generating an impedance signal corresponding thereto, at least one of the poles being arranged in the coronary sinus and/or in the great cardiac vein of the heart. The valve opening detector also includes an impedance signal processor which processes the impedance signal to detect the opening of one or both valves in the left side of the heart. | 1. An implantable medical device comprising: an aortic valve opening detector for detecting opening of the aortic valve of a heart;said aortic valve opening detector including an impedance measuring unit for measuring an electrical impedance between measurement poles and generating an impedance signal representing said electrical impedance, at least one of said measurement poles being disposed at a location selected from the group consisting of in the coronary sinus and in the great cardiac vein of said heart; andsaid aortic valve opening detector further including an impedance signal processor, supplied with said impedance signal, for processing said impedance signal and detecting opening of said aortic valve therefrom. | 1. An implantable medical device comprising: an aortic valve opening detector for detecting opening of the aortic valve of a heart;said aortic valve opening detector including an impedance measuring unit for measuring an electrical impedance between measurement poles and generating an impedance signal representing said electrical impedance, at least one of said measurement poles being disposed at a location selected from the group consisting of in the coronary sinus and in the great cardiac vein of said heart; andsaid aortic valve opening detector further including an impedance signal processor, supplied with said impedance signal, for processing said impedance signal and detecting opening of said aortic valve therefrom. 2. An implantable medical device as claimed in claim 1 wherein said impedance signal processor in said aortic valve opening detector processes said impedance signal according to a predetermined signal processing for detecting said opening of said aortic valve. 3. An implantable medical device as claimed in claim 2 wherein said impedance signal processor comprises a filter having filter characteristics for amplifying and filtering said impedance signal to detect said opening of said aortic valve, as said signal processing. 4. An implantable medical device as claimed in claim 3 wherein said filter has bandpass filter characteristics with a band range of between 10 Hz and 30 Hz. 5. An implantable medical device as claimed in claim 4 wherein said filter has a lower limit frequency and an upper limit frequency defining said band range, each of said upper and lower limit frequencies being variable. 6. An implantable medical device as claimed in claim 1 further comprising a housing containing said valve opening detector, and wherein said impedance measuring unit comprises an electrode lead carrying one pole, and an indifferent electrode on said housing, and wherein said impedance measurement unit measures said electrical impedance between said one pole on said electrode lead and said indifferent electrode. 7. An implantable medical device as claimed in claim 1 further comprising a housing containing said valve opening detector, and wherein said impedance measuring unit comprises an electrode lead carrying two poles, and an indifferent electrode on said housing, and wherein said impedance measurement unit measures said electrical impedance between said two poles on said electrode lead and said indifferent electrode. 8. An implantable medical device as claimed in claim 1 further comprising a housing containing said valve opening detector, and wherein said impedance measuring unit comprises an electrode lead carrying three poles, and an indifferent electrode on said housing, and wherein said impedance measurement unit measures said electrical impedance between said three poles on said electrode lead and said indifferent electrode. 9. An implantable medical device comprising: a mitral valve opening detector for detecting opening of ... | St. Jude Medical AB | A61B005/053
US06997882 | B1 | 2006-02-14 | 6-DOF subject-monitoring device and method | 2002-12-21 | The present invention comprises novel methods and devices for monitoring a subject by acquiring reliable and accurate 6-DOF data regarding the subject, and by using those data to obtain information about the subject's movements in three-dimensional space. Information regarding the subject's movements is, optionally, combined with information regarding the subject's physiological status so that comprehensive knowledge regarding the subject may be acquired by those monitoring the subject. | 1. A method of monitoring a subject during a monitoring period, said method comprising the steps of: (a) attaching at least one array of accelerometer modules to at least one body-segment of the subject, the accelerometer modules having an aggregate of at least six axes of measurement;(b) positioning the subject in at least three substantially stationary calibration poses, the calibration poses being sufficiently different from one another to produce appreciable differences in acceleration signals produced by the array of step (a);(c) acquiring acceleration signals from the array of step (a) for each of the calibration poses of step (b), said acceleration signals representing the accelerations of the body-segment relative to each of the x, y, and z-axes of an anatomical reference frame;(d) averaging the acceleration signals acquired at step (c) to minimize the effects of noise and involuntary motion;(e) acquiring a gravity vector value from the averaged acceleration signals obtained at step (d);(f) applying the gravity vector value acquired at step (e) to the acceleration signals obtained at step (c) to obtain a correction factor for correcting the orientation of each axis of measurement relative to the anatomical reference-frame;(g) acquiring from the array of step (a) acceleration signals representing the accelerations of the body-segment movements relative to each of the x, y, and z—axes of an anatomical reference-frame;(h) processing the acceleration signals acquired at step (g) to obtain 6-DOF data;(i) processing the 6-DOF data obtained at step (h) by incorporating the correction factor obtained at step (f) to obtain information regarding the body-segment movement with respect to the x, y, and z-axes of an inertial reference-frame; and(j)) displaying the body-segment movement information obtained at step (i). | 1. A method of monitoring a subject during a monitoring period, said method comprising the steps of: (a) attaching at least one array of accelerometer modules to at least one body-segment of the subject, the accelerometer modules having an aggregate of at least six axes of measurement;(b) positioning the subject in at least three substantially stationary calibration poses, the calibration poses being sufficiently different from one another to produce appreciable differences in acceleration signals produced by the array of step (a);(c) acquiring acceleration signals from the array of step (a) for each of the calibration poses of step (b), said acceleration signals representing the accelerations of the body-segment relative to each of the x, y, and z-axes of an anatomical reference frame;(d) averaging the acceleration signals acquired at step (c) to minimize the effects of noise and involuntary motion;(e) acquiring a gravity vector value from the averaged acceleration signals obtained at step (d);(f) applying the gravity vector value acquired at step (e) to the acceleration signals obtained at step (c) to obtain a correction factor for correcting the orientation of each axis of measurement relative to the anatomical reference-frame;(g) acquiring from the array of step (a) acceleration signals representing the accelerations of the body-segment movements relative to each of the x, y, and z—axes of an anatomical reference-frame;(h) processing the acceleration signals acquired at step (g) to obtain 6-DOF data;(i) processing the 6-DOF data obtained at step (h) by incorporating the correction factor obtained at step (f) to obtain information regarding the body-segment movement with respect to the x, y, and z-axes of an inertial reference-frame; and(j)) displaying the body-segment movement information obtained at step (i). 2. The method of claim 1 further comprising adaptively calibrating the accelerometer modules using recursive prediction-error analysis by carrying out the steps of: (k) constructing a model of the subject's movements based on the 6-DOF data obtained at step (h);(l) guessing a value with respect to some chosen future point in time for at least one parameter used to construct the model;(m) measuring the value for the parameter when the future point in time arrives;(n) obtaining a correction factor by refining the guess made at step (k) with respect to a next future point in time and applying a nonlinear optimization algorithm that minimizes the size of the error between the value guessed at step (l) and value measured at step (m);(o) applying the correction factor obtained at step (n) to the 6-DOF data obtained in step (h), whereby the corrected 6-DOF data more accurately represent the movements of the body-segment with respect to the inertial reference-frame; and,(p) repeating steps (k) through (o) during the monitoring period. 3. A method of monitoring respiratory functions of a subject, said method comprising the steps of: (a) attach... | Barron Associates, Inc. | A61B005/08
US07022071 | B2 | 2006-04-04 | Method for measuring the concentration of substances in living organisms using microdialysis and a device for carrying out said method | 2002-07-05 | The invention relates to a method for measuring the concentration of substances in living organisms using microdialysis, whereby bodily fluid is mixed with a perfusate and passed through a membrane and a substance value is determined for the bodily fluid with the aid of a sensor unit. According to the invention, to allow for the limited transport through the membrane, so that the measured substance value differs from the actual substance value in the organism, a reference substance present in said organism with a known value is also measured after passing through the membrane and mixing with the perfusate and the actual substance value is determined, based on the quotient from the known reference substance value divided by the measured reference substance value, multiplied by the measured substance value. | 1. Method for measuring the concentration of substances in living organisms using microdialysis, wherein bodily fluid is mixed with a perfusate and passed through a membrane, and a substance concentration of the bodily fluid is determined using a sensor unit, characterized in that in order to take into consideration the restricted transport of the substance through the membrane, so that the measured substance concentration differs from the actual substance concentration in the organism, an endogenous substance used as a reference substance, having a known value, is measured after passing through the membrane and mixing with the perfusate, and the actual substance concentration is determined on the basis of the quotient of the known reference substance value divided by the measured reference substance value, multiplied by the measured substance concentration. | 1. Method for measuring the concentration of substances in living organisms using microdialysis, wherein bodily fluid is mixed with a perfusate and passed through a membrane, and a substance concentration of the bodily fluid is determined using a sensor unit, characterized in that in order to take into consideration the restricted transport of the substance through the membrane, so that the measured substance concentration differs from the actual substance concentration in the organism, an endogenous substance used as a reference substance, having a known value, is measured after passing through the membrane and mixing with the perfusate, and the actual substance concentration is determined on the basis of the quotient of the known reference substance value divided by the measured reference substance value, multiplied by the measured substance concentration. 2. Method according to claim 1, characterized in that any different transport behavior of the reference substance and the substance to be measured is taken into consideration by means of a correction factor that is determined in advance. 3. Method according to claim 2, characterized in that the correction factor is determined in advance, in vitro, by mixing a solution with ions and the substance to be determined, and passing it through a membrane. 4. Method according to claim 2, characterized in that the correction factor is determined in advance, in vivo, by measuring the substance value in the blood. 5. Method according to claim 1, characterized in that the measurement of the value to be determined is performed on the basis of measuring the electrical conductivity, corresponding to the sum of ions. 6. Method according to claim 1, characterized in that in order to achieve high flow-through rates, the perfusate or the mixture of perfusate/bodily fluid is pumped into and out of the body. 7. Device for measuring the concentration of substances in living organisms using microdialysis, having a microdialysis catheter that has a membrane, through which bodily fluid is mixed with a perfusate that is supplied from a perfusate reservoir, and having a sensor unit for determining the substance concentration of the bodily fluid, characterized in that the membrane is also set up for passage of an endogenous substance used as a reference substance, having a known value, and this reference substance is measured after passing through the membrane and mixing with the perfusate, using the sensor unit, and that the sensor unit is connected with an electronic evaluation unit, which is set up to determine the actual substance concentration on the basis of the quotient of the known reference substance value divided by the measured reference substance value, multiplied by the measured substance concentration. 8. Device according to claim 7, wherein a first pump unit is arranged on the inflow side of the microdialysis catheter, and a second pump unit is arrancred on the outflow side of the microdialysis cath... | - | A61B005/00
US07225012 | B1 | 2007-05-29 | Methods and systems for image-guided surgical interventions | 2000-09-18 | Featured is a system to determine the three-dimensional position and orientation of an effector (a needle, probe, or other medical instrument) relative to a subject using cross-sectional images (e.g. from a CT or MRI scanner). Also provided is a method for image guided effector placement that requires no immobilization of the patient or fiducial implantation. A localization module (fiducial object) is integrated or associated with the effector allowing for the localization of the effector in the image space using a single cross-sectional image. | 1. An imaging system for invasive therapy of a patient, the system comprising: an imaging apparatus that can provide a cross-sectional image of a patient; anda medical instrument comprising a fiducial object that can be simultaneously imaged in the same image as a targeted site of the patient wherein the fiducial object comprises three N-shaped fiducial motifs each on a coplanar space, the coplanar space of each N-shaped fiducial motif being non-coplanar. | 1. An imaging system for invasive therapy of a patient, the system comprising: an imaging apparatus that can provide a cross-sectional image of a patient; anda medical instrument comprising a fiducial object that can be simultaneously imaged in the same image as a targeted site of the patient wherein the fiducial object comprises three N-shaped fiducial motifs each on a coplanar space, the coplanar space of each N-shaped fiducial motif being non-coplanar. 2. The system of claim 1 wherein the fiducial object representation in the image is unique for the pose of the instrument within the therapeutic field or range of motion of the instrument. 3. The system of claim 1 or 2 wherein the image can produce three identifiable points to coordinate pose of the instrument and the targeted site of the patient. 4. The system of claim 1 wherein the instrument pose is directly manipulated in reference to the medical image. 5. The system of claim 1 wherein the relative position and orientation of the medical instrument and target site of the patient can be determined from the information contained in a single cross-sectional image produced by the imaging apparatus. 6. The system of claim 1 wherein the system comprises a control apparatus that can register the instrument in a detected image space and calculate instrument movement. 7. The system of claim 6 wherein the control apparatus calculates the instrument pose in the image space by generating at least three corresponding points. 8. The system of claim 1 wherein the three N-shaped fiducial motifs are arranged orthogonally in a U-shape with one fiducial motif forming the bottom and two fiducial motifs forming the sides. 9. The system of claim 1 wherein the medical instrument is manipulated manually. 10. The system of claim 1, wherein the system further comprises a robotic apparatus capable of positioning the medical instrument. 11. The system of claim 10 wherein the instrument is positioned by the robot in the desired pose relative to the patient. 12. The system of claim 1 wherein the imaging device is a CT, MRI or ultrasound device. 13. The system of claim 1 wherein the fiducial object is affixed to the instrument. 14. The system of claim 1 wherein the fiducial object is integral to the instrument. 15. A method for guiding invasive therapy in a patient, comprising: a) providing a system that comprises an imaging apparatus and a medical instrument comprising a fiducial object that can be simultaneously imaged in the same image as a targeted site of the patient, the fiducial object comprising three N-shaped fiducial motifs each on a coplanar space, the coplanar space of each N-shaped fiducial motif being non-coplanar;b) obtaining a cross-sectional image that comprises both the fiducial object and the targeted site of the patient; andc) manipulating the instrument with respect to the patient using information derived from the image. 16. The method of claim 15 wherein the relative position and orie... | The Johns Hopkins University | A61B005/00
US06992477 | B2 | 2006-01-31 | Medical device with position sensor having core with high permeability material for determining location coordinates of a portion of the medical device | 2001-06-15 | A medical location system includes a medical device having a body and a position sensor at a portion of the body. The position sensor has a core made of a Wiegand effect material and a winding circumferentially positioned around the core. The position sensor provides signals that are used to determine temperature at the position sensor and location information of the portion of the body of the medical device. A signal processor is coupled to receive the signals from the position sensor and the signal processor determines the temperature at the position sensor and location information of the portion of the body of the medical device based on the signals received from the position sensor. | 1. A medical location system comprising: a medical device having a body;a position sensor at a portion of the body, the position sensor having a core made of a Wiegand effect material; and a winding circumferentially positioned around the core, the position sensor providing signals that are used to determine temperature at the position sensor and are used to determine location information of the portion of the body of the medical device; anda signal processor having a temperature sensitivity algorithm and a location algorithm and, which is coupled to receive the signals from the position sensor, the signal processor determining a temperature at the position sensor using the temperature sensitivity algorithm and location information of the portion of the body of the medical device based on the signals received from the position sensor using the location algorithm. | 1. A medical location system comprising: a medical device having a body;a position sensor at a portion of the body, the position sensor having a core made of a Wiegand effect material; and a winding circumferentially positioned around the core, the position sensor providing signals that are used to determine temperature at the position sensor and are used to determine location information of the portion of the body of the medical device; anda signal processor having a temperature sensitivity algorithm and a location algorithm and, which is coupled to receive the signals from the position sensor, the signal processor determining a temperature at the position sensor using the temperature sensitivity algorithm and location information of the portion of the body of the medical device based on the signals received from the position sensor using the location algorithm. 2. The medical location system according to claim 1, wherein the location information comprises position coordinates. 3. The medical location system according to claim 2, wherein the location information further comprises orientation coordinates. 4. The medical location system according to claim 3, wherein the position sensor has an accuracy within approximately 0.5 mm. 5. The medical location system according to claim 1, wherein the signal processor determines the location information with accuracy of ≦1 mm at temperatures greater than 75° C. 6. The medical location system according to claim 5, wherein the signal processor determines the location information with accuracy of ≦1 mm at temperatures at approximately 80° C. 7. The medical location system according to claim 1, wherein the core has an outer diameter less than approximately 0.3 mm. 8. The medical location system according to claim 7, wherein the core has an outer diameter of about 0.25 mm. 9. The medical location system according to claim 8, wherein the winding is attached to the core. 10. The medical location system according to claim 9, wherein a combination of the core and the wire winding has an outer diameter less than approximately 0.5 mm. 11. The medical location system according to claim 10, wherein the combination of the core and the winding have an outer diameter of about 0.4 mm. 12. The medical location system according to claim 11, wherein the material of the core comprises cobalt. 13. The medical location system according to claim 12, wherein the material of the core further comprises vanadium. 14. The medical location system according to claim 13, wherein the material of the core further comprises iron. 15. The medical location system according to claim 14, wherein the material of the core comprises approximately 20%–80% cobalt. 16. The medical location system according to claim 14, wherein the material of the core comprises approximately 2%–20% vanadium. 17. The medical location system according to claim 14, wherein the material of the core comprises approximately 25%–50% iron. 18. The medical l... | Biosense, Inc. | G01B007/004 G01R033/00 A61B005/05
US07120481 | B2 | 2006-10-10 | Probe and apparatus for measuring cerebral hemodynamics and oxygenation | 2001-07-19 | A probe and an apparatus for cerebral diagnostics and therapy, in particular for measuring absolute values of regional cerebral flow (CBF) and cerebral oxygenation. The probe is inserted through a burr hole in the skull and comprises illuminating device, light receiving device and a coating encapsulating said illuminating device and said light receiving means. The coating has a longitudinal shape and is adapted to fit through a burr hole in the skull. The coating is further adapted to slide between the skull and the dura, to be inserted into the ventricular system, and/or to be inserted into the cerebral tissue. | 1. A probe for cerebral diagnostics and/or therapy, in particular for measuring characteristics of cerebral hemodynamics and oxygenation by optical reflectance, comprising illuminating means comprising first optical transmission means including at least one first optical fiber, and first deflection means coupled to the first optical transmission means for deflection of transmitted light into a direction other than the direction of light transmission within the first optical transmission means;light receiving means comprising second optical transmission means including at least one second optical fiber, and second deflection means coupled to the second optical transmission means for deflection of light into the second optical transmission means, the light coming from a direction other than the direction of light transmission within the second optical transmission means, wherein the first and second deflection means are located at a distance (D1) from each other of 20 to 50 mm;a coating encapsulating said illuminating means and said light receiving means, said coating having a longitudinal shape and being adapted to fit through a burr hole in the skull, said coating further being adapted to at least one of the following: sliding between the skull and the dura, being inserted into the ventricular system, being inserted into the cerebral tissue. | 1. A probe for cerebral diagnostics and/or therapy, in particular for measuring characteristics of cerebral hemodynamics and oxygenation by optical reflectance, comprising illuminating means comprising first optical transmission means including at least one first optical fiber, and first deflection means coupled to the first optical transmission means for deflection of transmitted light into a direction other than the direction of light transmission within the first optical transmission means;light receiving means comprising second optical transmission means including at least one second optical fiber, and second deflection means coupled to the second optical transmission means for deflection of light into the second optical transmission means, the light coming from a direction other than the direction of light transmission within the second optical transmission means, wherein the first and second deflection means are located at a distance (D1) from each other of 20 to 50 mm;a coating encapsulating said illuminating means and said light receiving means, said coating having a longitudinal shape and being adapted to fit through a burr hole in the skull, said coating further being adapted to at least one of the following: sliding between the skull and the dura, being inserted into the ventricular system, being inserted into the cerebral tissue. 2. The probe according to claim 1, wherein the coating is made of silicone rubber or polyurethane. 3. The probe according to claim 1, having a width (W) less than about 20 mm, and a thickness (T) less than about 5 mm. 4. The probe according to claim 1, further comprising a pressure sensor and signal transfer means connected to said pressure sensor for transmitting signals containing pressure information. 5. The probe according to claim 1, further comprising means for the transfer and release of a substance into the cerebral tissue and/or into the ventricular system. 6. The probe according to claim 1, wherein the first and second deflection means deflect light into/from a direction (B) substantially vertical to the direction (A) of light propagation in the transmission means. 7. The probe according to claim 1, wherein the distance (D1) between the first and second deflection means is 30 to 40 mm. 8. The probe according to claim 1, wherein the first and second deflection means include a mirror oriented at approximately 45° with respect to the direction of the first and second optical transmission means. 9. The probe according to claim 8, wherein the mirror is a prison. 10. The probe according to claim 1, wherein the face of the at least one first or second fiber at the distal end of the first and second transmission means is oriented at approximately 45° with respect to the direction of light propagation (A) within the first respectively second optical transmission means. 11. The probe according to claim 1, wherein the first and second optical transmission means each include a plurality of first and s... | Universitat Zurich | A61B005/00
US06986745 | B2 | 2006-01-17 | Device for measuring inspiratory strength | 2002-07-03 | A device having variable inspiratory resistance suitable for measuring the inspiratory strength of a patient which comprises a chamber in communication with a mouthpiece, said chamber (a) being provided with a calibration system for measuring the inhalation flow rate during inhalation through the mouthpiece and (b) being provided with a variable resistance feature by which the resistance to airflow through the device may be altered. | 1. A device having variable inspiratory resistance suitable for measuring the inspiratory strength of a patient comprising: an elongate chamber having first and second ends defining a pathway and adapted for air to enter the second end wherein the first end of the chamber is in communication with a mouthpiece,located in the chamber, a calibration system adapted to measure an inhalation flow rate, the calibration system including sliding member adapted to slide and be retained within the chamber, and,a variable resistance system adapted to vary the area through which air leaves the chamber at the first end or enters the chamber at the second end, the variable resistance system adapted to include at least one specified setting that provides the device with a resistance matching that of an inhalation device. | 1. A device having variable inspiratory resistance suitable for measuring the inspiratory strength of a patient comprising: an elongate chamber having first and second ends defining a pathway and adapted for air to enter the second end wherein the first end of the chamber is in communication with a mouthpiece,located in the chamber, a calibration system adapted to measure an inhalation flow rate, the calibration system including sliding member adapted to slide and be retained within the chamber, and,a variable resistance system adapted to vary the area through which air leaves the chamber at the first end or enters the chamber at the second end, the variable resistance system adapted to include at least one specified setting that provides the device with a resistance matching that of an inhalation device. 2. The device of claim 1, wherein the chamber has a circular cross sectional area. 3. The device of claim 1, wherein the calibration system includes calibrations along the length of the chamber, and wherein a distance of travel by the sliding member from a rest position indicates a patient's maximum inspiratory flow rate. 4. The device of claim 1, wherein the calibration system includes calibration along the length of the sliding member, and wherein a distance of travel by the sliding member from a rest position indicates a patient's maximum inspiratory flow rate. 5. The device of claim 1, wherein the sliding member includes a means for guiding movement. 6. The device of claim 5, wherein the guiding means is a rod located along a chamber axis. 7. The device of claim 1, further comprising a means for restoring the sliding member to rest position. 8. The device of claim 7, wherein the restoring means is a weight. 9. The device of claim 1, wherein variable resistance system includes an iris. 10. The device of claim 9, wherein the iris includes two or more defined aperture size settings. 11. The device of claim 1, wherein the variable resistance system includes first and second co-operating aperture-bearing plates adapted to vary the area through which air leaves the chamber by orienting the plates. 12. The device of claim 11, wherein the first and second plates include at least two apertures. 13. The device of claim 12, including two or more orientations each defining a specific area. 14. The device of claim 12, including at least one orientation defining an area with an associated resistance corresponding to the resistance of an inhalation device. 15. A device having variable inspiratory resistance suitable for measuring the inspiratory strength of a patient comprising: an elongate chamber having first and second ends and defining a pathway and adapted for air to enter the second end wherein the first end of the chamber is in communication with a mouthpiece,located in the chamber, a calibration system adapted to measure an inhalation flow rate, the calibration system including a sliding member adapted to slide and be retained within ... | Smithkline Beecham Corp. | A61B005/08
USRE038712 | E1 | 2005-03-22 | Method and apparatus for storing signals in an implantable medical device | 2002-11-15 | A method of storing signals having a course over time in an implantable medical device, and/or transmitting the signals out of the device and out of the patient's body, with the course over time being detected in the device, particularly at predetermined time intervals with a constant sampling interval, and a sequence of signal samples being obtained therefrom, of which a partial quantity is stored and/or transmitted as the result of a selection made with a predetermined selection criterion, with the selection criterion being the first derivation of the course over time of the signals according to time, and with the selection being made such that the first derivation of the straight-line connection between an nth and the (n−1)th signal sample is determined and compared to the first derivation of the straight-line connection between the (n−1)thnth signal sample and the last-stored and/or last-transmitted signal sample, and the nth(n−1)th signal sample is then stored and/or transmitted as the new last signal sample exactly when these first derivations differ by more than a predetermined value. | 1. Method of storing signals having a course over time in an implantable medical device and/or transmitting the signals out of the device and out of a patient's body, said method comprising the steps of: sensing physiological signals of a patient's body; detecting a course over time associated with said sensed physiological signals in the device, said course over time being detected at predetermined time intervals with a constant sampling interval; obtaining a quantity of signal samples from the detected course over time; selecting a partial quantity of said obtained signal samples based on predetermining selection criterion wherein the selection criterion is the first derivation of the course over time of the signals according to time and said selecting step includes determining an initial first derivation of the straight-line connection between an nth and the (n−1)th signal sample; and comparing the initial first derivation to an earlier first derivation of the straight-line connection between (n−1)thnth signal sample and the last-stored and/or last-transmitted signal sample; and storing and/or transmitting the nth(n−1)th signal sample as the new last signal when the initial first derivation and the earlier first derivation differ by more than a predetermined value. | 1. Method of storing signals having a course over time in an implantable medical device and/or transmitting the signals out of the device and out of a patient's body, said method comprising the steps of: sensing physiological signals of a patient's body; detecting a course over time associated with said sensed physiological signals in the device, said course over time being detected at predetermined time intervals with a constant sampling interval; obtaining a quantity of signal samples from the detected course over time; selecting a partial quantity of said obtained signal samples based on predetermining selection criterion wherein the selection criterion is the first derivation of the course over time of the signals according to time and said selecting step includes determining an initial first derivation of the straight-line connection between an nth and the (n−1)th signal sample; and comparing the initial first derivation to an earlier first derivation of the straight-line connection between (n−1)thnth signal sample and the last-stored and/or last-transmitted signal sample; and storing and/or transmitting the nth(n−1)th signal sample as the new last signal when the initial first derivation and the earlier first derivation differ by more than a predetermined value. 2. Method according to claim 1, wherein the storing and/or transmitting step further includes storing or transmitting, for each stored and/or transmitted signal sample, the operational sign, the amplitude value and the duration of the time intervals that have passed since the last storage of a signal sample are stored or transmitted in a single data word. 3. Method according to claim 1, further comprising the step of holding the last-transmitted signal sample in intermediate storage in the case of time-near transmission of the signal samples of the partial sequence. 4. Method according to claim 1, wherein the step of obtaining signal samples comprises digitizing amplitude values. 5. Method according to claim 4, wherein the digitizing of smaller amplitude values is effected with a higher resolution than that of large amplitude values. 6. Method according to claim 1, wherein the physiological signals are cardiac-action signals, and further comprising the step of transmitting an intracardially-recorded EKG, or model or comparison signals representative thereof from outside of the patient's body. 7. Apparatus for storing signals having a course over time in an implantable device, said apparatus comprising: a timer for establishing sampling intervals, a signal recorder for recording the signal-amplitude value in the sampling intervals, a processing unit that is connected to the timer and at least indirectly to the signal recorder for processing the recorded signal-amplitude values corresponding to a predetermined algorithm, and for outputting a memory-control signal as a result of the processing, a counter that is connected to the output of the timer and, by way of a reset input, ... | Biotronik Mess- und Therapiegeraete GmbH & Co. | A61B005/04 G01D001/12
USRE038879 | E1 | 2005-11-15 | Extravasation detection technique | 2001-09-07 | The technique for detecting extravasation during the injection of fluid into a patient involves the establishment of a baseline representing impedance at the zone of the injection prior to the injection starting. Extravasation is signaled when at least two characteristics appear. First is that the impedance varies from the baseline more than a predetermined amount in more than a predetermined number of discreet time slots called epochs herein. Second is that, the rate of change of the impedance, which is called the slope herein, is consistently greater than a predetermined amount. | 1. The method of detecting extravasation that may occur when a needle is inserted into a patient for the purpose of delivering fluid into the patient's vascular system comprising the steps of: prior to the delivery of fluid, establishing an impedance baseline for patient tissue impedance near the tip of the needle; determining an impedance slope value based on deviations from said baseline for each of a plurality of time based epochs during the delivery of fluid, and signaling extravasation when said slope values are outside a first predetermined threshold with a consistency that meets a predetermined consistency criterion. | 1. The method of detecting extravasation that may occur when a needle is inserted into a patient for the purpose of delivering fluid into the patient's vascular system comprising the steps of: prior to the delivery of fluid, establishing an impedance baseline for patient tissue impedance near the tip of the needle; determining an impedance slope value based on deviations from said baseline for each of a plurality of time based epochs during the delivery of fluid, and signaling extravasation when said slope values are outside a first predetermined threshold with a consistency that meets a predetermined consistency criterion. 2. The method of claim 1 wherein: said predetermined consistency criterion includes the combination of (a) a first predetermined number of out of threshold slopes, and (b) a second predetermined number of consecutive out of threshold slopes subsequent to the infusion of a predetermined volume of said fluid. 3. The method of claim 2 wherein said first predetermined number is of consecutive out of threshold slopes. 4. The method of claim 2 wherein said first predetermined number is inversely proportional to the rate at which fluid is delivered. 5. The method of claim 2 wherein said first predetermined number is greater than said second predetermined number. 6. The method of claim 2 wherein said step of determining a slope value comprises determining a best fit slope value over a predetermined number of epochs adjacent to the epoch for which said slope value is determined. 7. The method of claim 2 wherein said second predetermined number of out of threshold slopes commences after said first predetermined number of out of threshold slopes is determined. 8. The method of claim 2 wherein: a sliding window of a plurality of epochs is employed to provide a base for the slope value determination, each slope value determination being based on a set of epochs that include epochs on which the preceding slope value determination was made. 9. The method of claim 1 wherein said step of determining a slope value comprises determining a best fit slope value over a predetermined number of epochs adjacent to the epoch for which said slope value is determined. 10. The method of claim 1 wherein: a sliding window of a plurality of epochs is employed to provide a base for the slope value determination, each slope value determination being based on a set of epochs that include epochs on which the preceding slope value determination was made. 11. The method of claim 1 further comprising the steps of: establishing said impedance baseline based on a first sliding window of epochs to provide an epoch impedance baseline, establishing a noise exclusion gate around said epoch impedance baseline, and during the delivery of fluid, providing a count of the number of consecutive epoch impedance averages which lie outside said gate, said step of signaling extravasation further requiring that said count of consecutive epoch impedance averages outsid... | E-Z-EM, Inc. | A61B005/05
US06898456 | B2 | 2005-05-24 | Method for determining a current lung filling extent and method for assisting radiation therapy during respiratory shifting of the radiation target | 2001-05-17 | The invention concerns a method for determining the filling of a lung, wherein the movement of an anatomical structure which moves with breathing, or one or more points on the moving anatomical structure whose movement trajectory is highly correlated with lung filling, is detected with respect to the location of at least one anatomical structure which is not spatially affected by breathing, and wherein each distance between the structures is assigned a particular lung filling value. It further concerns a method for assisting in radiotherapy during movement of the radiation target due to breathing, wherein the association of lung filling values with the distance of the moving structure which is identifiable in an x-ray image and the structure which is not spatially affected by breathing is determined by means of an above-named method, the current position of the radiation target is detected on the basis of the lung filling value, and wherein radiation exposure is carried out, assisted by the known current position of the radiation target. | 1. A method for determining the filling of a lung using image data to detect the movement of an anatomical structure that moves with the patient's breathing cycle, the method comprising the steps of, producing an image of a patient's body, determining the position of a first anatomical structure of the patient's body that moves with the patient's breathing cycle with respect to the position of a second anatomical structure of the patient's body that is not spacially affected by the patient's breathing cycle, determining the distance between the first and second anatomical structures, and assigning to this distance a lung filling value. | 1. A method for determining the filling of a lung using image data to detect the movement of an anatomical structure that moves with the patient's breathing cycle, the method comprising the steps of, producing an image of a patient's body, determining the position of a first anatomical structure of the patient's body that moves with the patient's breathing cycle with respect to the position of a second anatomical structure of the patient's body that is not spacially affected by the patient's breathing cycle, determining the distance between the first and second anatomical structures, and assigning to this distance a lung filling value. 2. The method as set forth in claim 1, wherein at least two images are produced by transillumination and at least two respective lung filling values are assigned and at least one additional lung filling value is determined by interpolation. 3. The method as set forth in claim 1, wherein at least two images are produced from tomographic imaging methods and at least two respective lung filling values are assigned and at least one additional lung filling value is determined by interpolation. 4. The method as set forth in claim 3, wherein digitally reconstructed x-ray images, obtained from a data set of a tomographic imaging method, are used as the images. 5. The method as set forth in claim 1, wherein the anatomical structure which moves with breathing is a structure which is identifiable on an x-ray image. 6.The method as set forth in claim 1, wherein the anatomical structure which is not spatially affected by breathing is a vertebral body or a part of a vertebral body. 7. The method as set forth in claim 1, wherein at least two images are produced by transillumination and at least two respective transillumination lung filling values are assigned and at least two images are produced from tomographic imaging methods and at least two respective tomographic lung filling values are assigned and the transillumination and the tomographic lung filling values are compared and/or aligned. 8. A method for assisting in radiotherapy during movement of a radiation target due to a patient's breathing cycle, said method comprising the steps of: acquiring at least two planning images of the patient's body during different times in the patient's breathing cycle, using the planning image data to determine the position of a first anatomical structure of the patient's body that moves with the patient's breathing cycle with respect to the position of a second anatomical structure of the patient's body that is not spacially affected by the patient's breathing cycle, determining the distance between the first and second anatomical structures and assigning to this distance a particular lung filling value, determining a relationship between the position of the radiation target and the lung filling value, acquiring at least two transillumination images of the patient's body during different times in the patient's breathing cycle wherei... | BrainLAB AG | A61B005/00
US06878119 | B2 | 2005-04-12 | Method for reducing artifacts in a spatial measurement | 2002-07-17 | A method for artifact reduction in sonomicrometer obtained intracranial pressure measurements generally comprises isolating a component of a sonomicrometer waveform attributable solely to changes in intracranial volume by using a neural network or other nonlinear engine to extract a heartbeat component from the sonomicrometer output. Because the heartbeat is so characteristic, no actual measurement of the heartbeat as the forcing function is required to isolate the resulting changes in distance from the artifact induced changes in distance. The neural network then be utilized to directly map measured changes in skull distance over time to changes in intracranial pressure over a volume change, the inverse of compliance. The method is generally extendable to use with other volumetric based measurement techniques. | 1. A method of artifact reduction in intracranial pressure measurements comprising isolating the change in distance attributable to changes in intracranial volume by extracting a heartbeat component from the measurements utilizing a neural network to map measured changes in skull distance over time. | 1. A method of artifact reduction in intracranial pressure measurements comprising isolating the change in distance attributable to changes in intracranial volume by extracting a heartbeat component from the measurements utilizing a neural network to map measured changes in skull distance over time. 2. The method according to claim 1 wherein changes in intracranial pressure over time are compared to changes in skull distance. 3. A method of artifact reduction in a spatial measurement of a cranium, comprising the steps of: isolating a first component of the spatial measurement from a second, artifact-induced component of the spatial measurement by extracting components of the spatial measurement that correspond to a heart beat waveform; and mapping measured changes in skull distance over time to changes in intracranial pressure over a volume change. 4. The method of claim 3, wherein a neural network is used to isolate the first component from the second, artifact-induced component. | KCI Licensing, Inc. | A61B005/00
US06875185 | B2 | 2005-04-05 | Integrated confirmation sample in a body fluid test device and method of using | 2002-07-29 | A lateral flow immunoassay contact chemical test device and method integrates sample collection, prescreen testing, and confirmation sample collecting and storing with a single device and a minimum of steps. While particularly advantageously used with an oral fluid sample absorbed directly from a person's mouth, (in which constant monitoring of a collection process is made possible), the test device may be used with any of a variety of sample fluids. Prescreen testing and confirmation testing are performed on a single sample. Prescreen testing is performed by a lateral flow assay. The sample collection pad is subsequently separated from the wicking path to prevent continued migration from, and backflow into the sample collection pad, so that the confirmation sample is preserved in the sample collection pad. A multitude of antigens can be detected with a single device. | 1. A test device for determining a presence of an analyte in a body fluid, the device integrating a prescreen test and a confirmation test sample collection with the device, the device comprising: a sample collection pad, the sample collection pad having a first end and a second end, the sample collection pad being adapted for absorbing the body fluid; a holder having a front-end, a middle, and a rear end, the front end of the holder removably holding the second end of the sample collection pad with the first end of the sample collection pad protruding from the holder for absorbing the body fluid; a flexible cap, the cap having a first end and a second end, the second end of the flexible cap supported on the holder; wherein the sample collection pad is movable relative to the holder between a first configuration in which the sample collection pad has not been moved and a first portion of the cap is spaced from the sample collection pad, and a second configuration in which the sample collection pad has been pulled and moved relative to the holder; and wherein the sample collection pad is gripped between inner walls of the cap by a pinching action and moved into the second configuration, the sample collection pad remaining in the cap with the cap replaced on the holder for protection of the sample collection pad in the second configuration. | 1. A test device for determining a presence of an analyte in a body fluid, the device integrating a prescreen test and a confirmation test sample collection with the device, the device comprising: a sample collection pad, the sample collection pad having a first end and a second end, the sample collection pad being adapted for absorbing the body fluid; a holder having a front-end, a middle, and a rear end, the front end of the holder removably holding the second end of the sample collection pad with the first end of the sample collection pad protruding from the holder for absorbing the body fluid; a flexible cap, the cap having a first end and a second end, the second end of the flexible cap supported on the holder; wherein the sample collection pad is movable relative to the holder between a first configuration in which the sample collection pad has not been moved and a first portion of the cap is spaced from the sample collection pad, and a second configuration in which the sample collection pad has been pulled and moved relative to the holder; and wherein the sample collection pad is gripped between inner walls of the cap by a pinching action and moved into the second configuration, the sample collection pad remaining in the cap with the cap replaced on the holder for protection of the sample collection pad in the second configuration. 2. The test device of claim 1, wherein the analyte is one of a drug, a hormone, an antigen, an antibody, a hapten, a lectin, an apoprotein, or a cofactor. 3. A test device for determining a presence of an analyte in a body fluid, the device integrating a prescreen test and a confirmation test sample collection with the device, the device comprising: a sample collection pad, the sample collection pad having a first end and a second end, the sample collection pad being adapted for absorbing the body fluid; a holder having a front-end, a middle, and a rear end, the front end of the holder removably holding the second end of the sample collection pad with the first end of the sample collection pad protruding from the holder for absorbing the body fluid; a flexible cap, the cap having a first end and a second end, the second end of the flexible cap supported on the holder; wherein the sample collection pad is movable relative to the holder between a first configuration in which the sample collection pad has not been moved and a first portion of the cap is spaced from the sample collection pad, and a second configuration in which the sample collection pad has been pulled and moved relative to the holder; and wherein the cap has a hollow volume in the first portion sized to receive the sample collection pad after separation of the sample collection pad from the holder, the second configuration comprising a separation of the sample collection pad from the holder. 4. The test device of claim 3, comprising a socket in the front end of the holder removably holding the sample collection pad, wherein the socket is surro... | Branan Medical Corporation | A61B005/00 B65D081/00
US06884215 | B1 | 2005-04-26 | Frequency estimation of electro-islet-graphy | 1999-05-28 | The fundamental frequency of EIG events is determined directly, without first detecting the events themselves. Use is made of an analysis window containing more than one event (3-5). Techniques (1-8) are utilized which are somewhat similar to those utilized to detect, and estimate pitch in speech processing. The invention is based on two important biological discoveries. The first is that the EIG is generated by a functional pace maker. The second is that the EIG signal is quasi-periodic most of the time. The use of pitch detection algorithms is based on the essentially quasi-periodic nature of the EIG signal. Quasi-periodic means that the intervals between successive events are not exactly identical, but may vary slightly, that the amplitude and shape of successive events may also exhibit some variance. | 1. A method for monitoring biological micro-structure activity which produces detectable signals characterizing events, comprising the steps of estimating the fundamental frequency of the occurrence of events from the detectable signals, without detecting the occurrence of individual events, wherein the estimating step includes an autocorrelation operation. | 1. A method for monitoring biological micro-structure activity which produces detectable signals characterizing events, comprising the steps of estimating the fundamental frequency of the occurrence of events from the detectable signals, without detecting the occurrence of individual events, wherein the estimating step includes an autocorrelation operation. 2. The method of claim 1 further including one of the following steps: estimating the fundamental frequency based upon a lower autocorrelation value disposed among several adjacent peaks; treating “unvoiced” segments of the detectable signal as undecided as to pitch and estimating the pitch of those segments through subsequent processing; seeking to estimate the fundamental frequency in the range of 0.25 to 5 Hertz, utilizing an analysis window duration in the range of several seconds; performing a pre-processing operation which has the effect of increasing the effective duration of an event; and utilizing an autocorrelation process which performs segmented autocorrelation. 3. A method for monitoring biological micro-structure activity which produces detectable signals characterizing events, comprising the steps of estimating the fundamental frequency of the occurrence of events from the detectable signals, without detecting the occurrence of individual events, further comprising the step of using the number of “unvoiced” windows within a predetermined time as a measure of a blood constituent level of a patient. 4. A method for monitoring biological micro-structure activity which produces detectable signals characterizing events, comprising the steps of estimating the fundamental frequency of the occurrence of events from the detectable signals, without detecting the occurrence of individual events, further comprising the step of using a sureness grade as measure of a blood or tissue constituent level of a patient. 5. In a system for monitoring biological micro-structure activity which produces detectable signals characterizing events, a sensor capable of receiving the sensible signals and a processor including a module for estimating the fundamental frequency of the occurrence of events from the detectable signals, without first detecting the occurrence of individual events, wherein the processor further comprises a module for producing an analysis window during which events are analyzed, the analysis window spanning more than one event, wherein the module for estimating includes components to perform an autocorrelation operation. 6. The system of claim 5 wherein the module for estimating further includes one of the following submodules: a submodule which estimates the fundamental frequency based upon a lower autocorrelation value disposed among several adjacent peaks; a sub-module which identifies “unvoiced” segments of the detectable signal as undecided as to pitch, the pitch of those segments being estimated by a subsequent processing submodule; a submodule which controls the estimat... | Carmel Biosensors Ltd. | A61B005/00
US06956646 | B1 | 2005-10-18 | Direct method for the correction of pressure induced scrambling of polarized fluorescence intensities | 2000-09-11 | The present invention provides a simple and direct method for the simultaneous correction of steady-state polarized fluorescence intensities, depolarized (or scrambled) by the effects of applied hydrostatic pressure without having to first determine the scrambling factors from a separate experiment. The method involves direct recalculation of the measured polarized data obtained for the sample of interest at the time of data collection. | 1. A method for the extraction of true values of emission anisotropy (corr) from fluorescence intensities obtained for a sample under an applied hydrostatic pressure (p), comprising the steps of measuring polarized fluorescence intensities and then determining excitation and emission correction factors. | 1. A method for the extraction of true values of emission anisotropy (&lt;r>corr) from fluorescence intensities obtained for a sample under an applied hydrostatic pressure (p), comprising the steps of measuring polarized fluorescence intensities and then determining excitation and emission correction factors. 2. The method of claim 1 wherein said true values of emission anisotropy are obtained from said fluorescence intensities without performing a separate pressurized calibration experiment. 3. The method of claim 1 wherein said excitation correction factor (X) and said emission correction factor (Y) are determined for a given pressure (p) from said fluorescence intensities substantially according to the equations: X⁡(p)=G·iHV-iHHG·iHV-iHH+E·(G·iVV-iVH)and:Y⁡(p)=E·iVH-iHHE·iVH-iHH+G·(E·iVV-iHV)wherein iVV, iVH, iHH, and iHV represent the measured and pressure induced distorted polarized intensities for the sample of interest, and E and G, are both sample and pressure independent instrument factors characteristic for chosen excitation and emission wavelength conditions. 4. The method of claim 3 wherein the E-factor corrects for any inequality in intensities of vertical and horizontal polarized excitation light, the G-factor corrects for unequal sensitivity of a detection system to vertical and horizontal polarized emission light, and said E and G factors are determined at atmospheric pressure according to the equations: G=iHH0iHVo⁢⁢and⁢⁢E=iHH0iVH0where iVH0, iHH0, and iHV0 are polarized fluorescence intensities obtained at atmospheric pressure. 5. The method of claim 4 wherein said true values of emission anisotropy (&lt;r>corr) are obtained from the equations: 〈r〉corr=R-1R+2-3·(X+Y-X·Y+R·Y-R·X·Y);R=G·iVViVH. 6. The method of claim 4, further comprising determining corrected total intensities (Scorr) in accordance with the following formula: Scorr=G·1-3·(Y-X·Y)1-X-2·(Y-X·Y)·iVV+2-3·(X+Y-X·Y)1-X-2·(Y-X·Y)·iVH. 7. The method of claim 3, further comprising using said excitation and emission correction factors to detect abnormalities in an optical window. 8. A method as recited in claim 1, further comprising determining a steady state fluorescence emission anisotropy value (&lt;r>corr). 9. A method for the extraction of corrected values of total intensities (Scorr) from polarized fluorescence intensities obtained for a sample under an applied hydrostatic pressure (p), comprising the steps of measuring polarized fluorescence intensities and then determining excitation and emission correction factors. 10. The method of claim 9 wherein said corrected total intensities (Scorr) are obtained from said polarized fluorescence intensities without performing a separate pressurized calibration experiment. 11. The method of claim 8 wherein said excitation correction factor (X) and said emission correction factor (Y) are determined for a given pressure (p) from said fluorescence intensities substantially according to the equations: X⁡(p)=G·iHV-iHHG... | - | G01J003/30 G01J004/00 A61B005/00
US06889073 | B2 | 2005-05-03 | Breast biopsy and therapy system for magnetic resonance imagers | 2001-05-02 | The present invention describes a device for performing breast biopsies and/or therapy within magnetic resonance imaging (MRI) systems. The apparatus includes a RF receiver antenna for magnetic resonance imaging of the breast. The RF coil includes openings in the front and side to provide access to the breast during the procedure. Compression plates are integrated into the breast coil which compress the breast either laterally or in the head/feet direction as required for optimal access to the breast. The apparatus includes a mechanical device for positioning interventional instruments in the breast such as biopsy or therapy instruments. The mechanical positioning devices position the instrument along the desired trajectory to the target site and insert the instrument into the breast while the patient remains inside the MRI scanner. Real time MR images may be acquired during instrument alignment and insertion to verify the trajectory. The mechanical positioning devices allow manipulation of instruments in any type of MRI scanner, including high field MRI systems with cylindrical magnets. The positioning devices provide a means to overcome limited access to the patient in MRI scanners. The positioning devices may be manually operated by means of gears, drive shafts, cables or other mechanical means. Or they may be electronically controlled by means of MR compatible motorized drive systems. The devices may be remotely controlled from outside the magnet for MRI systems that have limited access to the patient in the magnet. An interface between the electronically controlled drivers and the MRI scanner computer can provide robotic control of the instrument. | 1. A MR compatible mechanical device for positioning an interventional instrument such as a biopsy needle or therapy probe inside a MRI scanner, said positioning device comprising; an instrument platform in which a biopsy or therapy instrument is secured to a mechanical positioning device, said platform being movable to align the instrument trajectory, insert the instrument to the patient and remove the instrument from the patient. a mounting block to control the vertical and horizontal motion of the instrument platform, said mounting block including or attached to a mechanical means for horizontal motion of the instrument platform and another mechanical means for vertical motion of the instrument platform. a base-plate that is attached to a mechanical means that moves said base-plate in a horizontal or vertical direction, a single post or a plurality of posts that support the aforementioned mounting block and physically link said mounting block to the aforementioned base-plate, said support post(s) attached to a mechanical means for moving said mounting block, a plurality of mechanical means for adjusting the position of the aforementioned instrument platform, mounting block and base-plate. | 1. A MR compatible mechanical device for positioning an interventional instrument such as a biopsy needle or therapy probe inside a MRI scanner, said positioning device comprising; an instrument platform in which a biopsy or therapy instrument is secured to a mechanical positioning device, said platform being movable to align the instrument trajectory, insert the instrument to the patient and remove the instrument from the patient. a mounting block to control the vertical and horizontal motion of the instrument platform, said mounting block including or attached to a mechanical means for horizontal motion of the instrument platform and another mechanical means for vertical motion of the instrument platform. a base-plate that is attached to a mechanical means that moves said base-plate in a horizontal or vertical direction, a single post or a plurality of posts that support the aforementioned mounting block and physically link said mounting block to the aforementioned base-plate, said support post(s) attached to a mechanical means for moving said mounting block, a plurality of mechanical means for adjusting the position of the aforementioned instrument platform, mounting block and base-plate. 2. The apparatus of claim 1, further comprising a means of mounting the positioning device to said imaging coil. 3. The apparatus of claim 1, whereby the positioning device is located in front of the coil, providing a means to insert instruments into the patient from the direction of the patient's head. 4. The apparatus of claim 1, whereby the positioning device is located on the side of the coil, providing a means to insert intents into the patient from the lateral direction. 5. The apparatus of claim 1, further comprising a plurality of positioning devices that can be used to position a plurality of instruments in one or both breasts. 6. The apparatus of claim 1, further comprising an acme screw for moving the base-plate along a horizontal or vertical axis. 7. The apparatus of claim 1, further comprising an inchworm gear for moving the instrument platform along a horizontal or vertical axis. 8. The apparatus of claim 1, further comprising a rack and pinion mechanism for moving the instrument platform along a horizontal or vertical axis. 9. The apparatus of claim 1, further comprising a mechanical means of rotating the support post about an axis. 10. The apparatus of claim 9, further comprising a plurality of gears for rotating the support post about an axis. 11. The apparatus of claim 1, further comprising manually operated drive shafts and gears to adjust the position of the instrument platform inside the MRI scanner. 12. The apparatus of claim 1, further comprising cables to adjust the position of the instrument platform. 13. The apparatus of claim 1, further comprising an instrument platform with side rails for guiding an instrument along a predetermined trajectory. 14. The apparatus of claim 1, further comprising an instrument platform wit... | - | A61B005/05
US06889075 | B2 | 2005-05-03 | Optical imaging of subsurface anatomical structures and biomolecules | 2001-05-03 | The present invention provides various methods/systems of optical imaging of subsurface anatomical structures and biomolecules utilizing red and infrared radiant energy. Also provided are various applications of such methods/systems in medical diagnosis and treatment. | 1. A method of optically imaging subsurface anatomic structures and biomolecules in an individual or animal comprising: illuminating a region of interest with light having a wavelength from the red to radiant infrared region of the light spectrum via a light source; detecting red and infrared light from said region of interest with a red and infrared light sensitive image detector; enhancing said detected image via pulsatile enhanced imaging, confocal enhanced imaging, Raman enhanced imaging, laser speckle enhanced imaging, multiphoton interaction enhanced imaging, optical coherence tomography enhanced imaging, time correlated single photon counting enhanced imaging, optical rotary dispersion image, circular dichroism imaging or polarization enhanced imaging; transmitting said enhanced image to a video monitor external to said image detector; and displaying said enhanced image on the video monitor thereby optically imaging subsurface anatomic structures and biomolecules. | 1. A method of optically imaging subsurface anatomic structures and biomolecules in an individual or animal comprising: illuminating a region of interest with light having a wavelength from the red to radiant infrared region of the light spectrum via a light source; detecting red and infrared light from said region of interest with a red and infrared light sensitive image detector; enhancing said detected image via pulsatile enhanced imaging, confocal enhanced imaging, Raman enhanced imaging, laser speckle enhanced imaging, multiphoton interaction enhanced imaging, optical coherence tomography enhanced imaging, time correlated single photon counting enhanced imaging, optical rotary dispersion image, circular dichroism imaging or polarization enhanced imaging; transmitting said enhanced image to a video monitor external to said image detector; and displaying said enhanced image on the video monitor thereby optically imaging subsurface anatomic structures and biomolecules. 2. The method of claim 1, wherein said wavelength is about 600 nm to about 1100 nm. 3. The method of claim 1, wherein said detected red and infrared light is transmitted light, reflected light, absorbed light, or emitted light. 4. The method of claim 1, further comprising: adding an exogenous chromophore to the region of interest. 5. The method of claim 4, wherein said chromophore is indocyanine Green (ICG) or d-aminolevulinic acid. 6. The device of claim 1, wherein said light is a light-emitting diode (LED) filtered with a bandpass filter, a diode laser or filtered broadband illumination. 7. The device of claim 1, wherein said image detector is a charge-coupled device (CCD) or a CCD video camera. | Rocky Mountain Biosystems, Inc. | A61B005/00
US06879693 | B2 | 2005-04-12 | Method and system for external assessment of hearing aids that include implanted actuators | 2002-02-26 | A noninvasive method and system are provided for assessing the performance of implanted actuators of semi or fully-implantable hearing aid systems. The invention utilizes an externally positioned test measurement device to obtain measurements of the electrical impedance of an implanted actuator when driven by a test signal of predetermined characteristics. In one embodiment, the test measurement device may comprise a signal generator for generating the test signal for the actuator, a signal processing unit to compute the electrical impedance from voltage and current measurements, and a user interface to provide an output that is usable to asses the performance of the actuator. The electrical impedance is computable from the voltage and current of the signal passing through the actuator. The electrical impedance is directly related to the mechanical impedance present at the interface between the actuator and middle ear of a patient. As such, by driving the actuator at one or more predetermined frequencies the resultant voltage and current measurements may be utilized to assess whether the implanted transducer is operative and whether a desired interface between the transducer and the middle ear of patient (e.g. the ossicular chain) is present. | 1. A method for assessing the performance of a hearing aid that includes an implanted hearing aid actuator comprising: positioning a test measurement device external to a patient having an implanted hearing aid that includes a hearing aid actuator; utilizing the test measurement device to obtain at least one impedance measure of the hearing aid actuator responsive to an electrical signal passing through the hearing aid actuator; employing the at least one impedance measure to assess an interface between the actuator and a component of an auditory system of the patient; and providing an electrical input to a positioning system responsive to said assessment of said interface, to selectively position the hearing aid actuator relative to the component of the auditory system. | 1. A method for assessing the performance of a hearing aid that includes an implanted hearing aid actuator comprising: positioning a test measurement device external to a patient having an implanted hearing aid that includes a hearing aid actuator; utilizing the test measurement device to obtain at least one impedance measure of the hearing aid actuator responsive to an electrical signal passing through the hearing aid actuator; employing the at least one impedance measure to assess an interface between the actuator and a component of an auditory system of the patient; and providing an electrical input to a positioning system responsive to said assessment of said interface, to selectively position the hearing aid actuator relative to the component of the auditory system. 2. The method of claim 1, wherein the step of providing the electrical input comprises: providing a wireless signal to the positioning system from a position external to the patient. 3. The method of claim 1, wherein the step of providing the electrical input comprises: inductively coupling the electrical input to the positioning system. | Otologics, LLC. | H04R025/00 A61N001/08 A61B005/05 A61B005/12 H02N002/02
US06866641 | B2 | 2005-03-15 | Skin prickers | 2001-11-26 | A skin pricker for blood sampling has a barrel housing a lancet which can be pushed back against a spring to a primed condition, by an elongate cap over its needle tip there to be held by a trigger. The cap (4) is removed by twist and pull action. The lancet body has integrally formed spring arms extending rearwardly and alongside, and when the lancet is fired these are momentarily flexed inwardly as their tips snap past abutments within the barrel. These act as ratchets, providing a first defence against re-priming of the pricker. If that first defence is overcome by a substantial rearward force on the lancet, the spring arms rearward pointing V's which wedge between the abutments and the lancet body. The cap may have a weakness leaving it rigid enough for the initial priming but which causes it to buckle if used to try to overcome the ratchet. | 1. A disposable pricker comprising: an elongate housing with a spring-loaded, longitudinally movable lancet having a body and a needle tip carried therein, the needle tip normally being within the housing; a trigger mechanism to retain the lancet in a fully retracted position energizing the spring loading and actuable to release or fire the lancet to cause the needle tip to have a momentary position projecting from a forward end of the housing; means for preventing repeated use including a spring finger having a tip and extending rearwardly from the lancet alongside but spaced from the body thereof; and an abutment on the inside of the housing past which the tip of the finger can snap during forward motion of the lancet, any attempt to push the lancet back with a greater than predetermined force after firing causing the finger, with the finger tip arrested by the abutment to buckle into a rearward pointing V that wedges between the abutment and the body of the lancet. | 1. A disposable pricker comprising: an elongate housing with a spring-loaded, longitudinally movable lancet having a body and a needle tip carried therein, the needle tip normally being within the housing; a trigger mechanism to retain the lancet in a fully retracted position energizing the spring loading and actuable to release or fire the lancet to cause the needle tip to have a momentary position projecting from a forward end of the housing; means for preventing repeated use including a spring finger having a tip and extending rearwardly from the lancet alongside but spaced from the body thereof; and an abutment on the inside of the housing past which the tip of the finger can snap during forward motion of the lancet, any attempt to push the lancet back with a greater than predetermined force after firing causing the finger, with the finger tip arrested by the abutment to buckle into a rearward pointing V that wedges between the abutment and the body of the lancet. 2. The disposable pricker according to claim 1, wherein the lancet is symmetrical, with two fingers on opposite sides thereof, and the housing has two opposed abutments. 3. The disposable pricker according to claim 1, wherein the abutment is shaped as a barb, pointing inwardly of and towards the forward end of the housing. 4. The disposable pricker according to claim 1, wherein the finger inclines outwardly from the lancet body and also extends rearwardly. 5. The disposable pricker according to claim 1, wherein the finger is generally parallel to and spaced from an adjacent part of the body of the lancet but with the finger tip flaring outwardly. 6. The disposable pricker according to claim 1, wherein the needle tip is initially protected by an elongate cap by which the lancet is adapted to be pushed back to the fully retracted position from an initial pre-fired position in which latter position, the finger tip is immediately to the rear of the associated abutment. 7. The disposable pricker according to claim 6, wherein the elongate cap has a weakness that is not significant when the cap is used for retracting the lancet from said initial position, but which causes the cap to deform if used to try to push the lancet back after firing. 8. The disposable pricker according to claim 7, wherein the weakness is by way of a local reduction of the cross-section of the elongate cap, the deformation being buckling. | Owen Mumford Limited | A61B005/00
US06865410 | B2 | 2005-03-08 | Apparatus and method for measuring current flow in an animal or human body | 2001-08-01 | An apparatus for measuring contact current includes data acquisition circuitry with at least two sensor contacts to measure the voltage drop across an animal or human body. A portable data processing unit is connected to the data acquisition circuitry to process the voltage data to produce current flow data. | 1. An apparatus for measuring current flow through a living body, comprising: data acquisition circuitry consisting essentially of a first contact and a second contact configured for placement on a living body to measure voltage between said first contact and said second contact thereby producing voltage data, and a first connector connected at a first end to said first contact and a second connector connected at a first end to said second contact; and a portable data processing unit connected to a second end of said first connector and to a second end of said second connector to process said voltage data and body impedance data to produce current flow data representing an amount of current flow through the living body, wherein said portable data processing unit includes acquired data conditioning circuitry to condition a data signal from said first contact and said second contact. | 1. An apparatus for measuring current flow through a living body, comprising: data acquisition circuitry consisting essentially of a first contact and a second contact configured for placement on a living body to measure voltage between said first contact and said second contact thereby producing voltage data, and a first connector connected at a first end to said first contact and a second connector connected at a first end to said second contact; and a portable data processing unit connected to a second end of said first connector and to a second end of said second connector to process said voltage data and body impedance data to produce current flow data representing an amount of current flow through the living body, wherein said portable data processing unit includes acquired data conditioning circuitry to condition a data signal from said first contact and said second contact. 2. The apparatus of claim 1 further comprising an autonomous power source. 3. The apparatus of claim 1, wherein said acquired data conditioning circuitry includes an amplifier. 4. The apparatus of claim 1 wherein said portable data processing unit includes data input interface circuitry. 5. The apparatus of claim 4 further comprising a keypad connected with said data input interface circuitry. 6. The apparatus of claim 1 wherein said portable data processing unit includes data output interface circuitry. 7. The apparatus of claim 6 further comprising a visual display connected to said data output interface circuitry. 8. The apparatus of claim 6 further comprising a compact removable flash memory card connected to said data output interface circuitry. 9. An apparatus for measuring current flow through a living body, comprising: data acquisition circuitry consisting essentially of a first contact and a second contact configured for placement on a living body to measure voltage between said first contact and said second contact thereby producing voltage data, and a first connector connected at a first end to said first contact and a second connector connected at a first end to said second contact; and a portable data processing unit connected to a second end of said first connector and to a second end of said second connector to process said voltage data and body impedance data to produce current flow data representing an amount of current flow through the living body, wherein said portable data processing unit includes a central processing unit and a memory storing a set of executable programs and said body impedance data comprised of known body impedances. 10. The apparatus of claim 9 wherein said known body impedances include estimated impedances. 11. The apparatus of claim 9 wherein said known body impedances include calculated impedance measurements. 12. The apparatus of claim 9 wherein said memory stores an output module that controls data storage to a removable flash memory. 13. The apparatus of claim 9 wherein said memory stores an output module that... | Electric Power Research Institute, Inc. | A61B005/05
US06893404 | B2 | 2005-05-17 | Breathing movement measurements and apparatus | 2001-04-04 | A method and apparatus for measuring breathing movements and determining breathing patterns, by measuring the simultaneous movement of a plurality of points of a human subject, such that a breathing pattern may be determined based on data obtained in a single acquisition. The method provides a symmetry breathing pattern parameter. | 1. A method for measuring breathing movements comprising: measuring the simultaneous movement of a plurality of points of a human body, wherein the movement is measured by the change in the distance from each point of the plurality of points to one or more reference points; providing a set of data for breathing movement measurements over a time period; calculating breathing pattern parameters comprising: type of breathing, rhythm of breathing, magnitude of breathing movements, frequency of breathing movements, and symmetry of breathing movements, where type of breathing indicates whether the breathing movements are predominantly abdominal, low-costal or high-costal; and determining a breathing pattern based on the breathing pattern parameters. | 1. A method for measuring breathing movements comprising: measuring the simultaneous movement of a plurality of points of a human body, wherein the movement is measured by the change in the distance from each point of the plurality of points to one or more reference points; providing a set of data for breathing movement measurements over a time period; calculating breathing pattern parameters comprising: type of breathing, rhythm of breathing, magnitude of breathing movements, frequency of breathing movements, and symmetry of breathing movements, where type of breathing indicates whether the breathing movements are predominantly abdominal, low-costal or high-costal; and determining a breathing pattern based on the breathing pattern parameters. 2. The method according to claim 1, wherein a breathing pattern analysis is obtained for a subject in a single acquisition. 3. The method according to claim 1, wherein the plurality of points comprises one or more points on each side of the sternum. 4. The method according to claim 1, wherein the plurality of points is symmetrically distributed with respect to the sternum. 5. The method according to claim 1, wherein a computer system receives movement measurement data and calculates the breathing pattern parameters. 6. The method according to claim 1, further comprising correlating breathing patterns with other physical and/or chemical functions of a human body by a simultaneous measurement selected from the group comprising: electroencephalographic measurements, electrocardiographic measurements, electromyographic measurements, sonographic measurements, heart rate measurements, measurements with a pulse oxymeter, blood pressure measurements, and sound measurement such as for snoring. 7. A method for evaluating the effect of a medical condition on breathing functions by measuring breathing patterns according to any one of claims 1 to 6. 8. The method according to claim 7, wherein the medical condition is selected from the group comprising diseases of the breathing organs including emphysema, respiratory insufficiency, and asthma; post-operative conditions including conditions after operations of organs in the chest and/or abdomen such as lung operations, coronary bypass operations, operations of the digestion system; post-injury conditions; and rheumatism conditions. 9. The method according to claim 8, wherein the effect evaluated is an asymmetrical effect on breathing functions. 10. An apparatus for determining breathing patterns of a human comprising: an imaging device comprising: one or more distance measurement devices that measure the change in distance from each point of a plurality of points to one or more reference points; wherein the apparatus measures the simultaneous movement of a plurality of points of a human body to calculate breathing pattern parameters comprising: type of breathing, rhythm of breathing, magnitude of breathing movements, frequency of breathing movements, and symmetr... | Remo Ehf | A61B005/08
US06847838 | B1 | 2005-01-25 | Apparatus and method for high dose rate brachytherapy radiation treatment | 2002-09-19 | A method and apparatus for the in vivo location and tracking of a radioactive seed source during and after brachytherapy treatment. The method comprises obtaining multiple views of the seed source in a living organism using: 1) a single PSPMT detector that is exposed through a multiplicity of pinholes thereby obtaining a plurality of images from a single angle; 2) a single PSPMT detector that may obtain an image through a single pinhole or a plurality of pinholes from a plurality of angles through movement of the detector; or 3) a plurality of PSPMT detectors that obtain a plurality of views from different angles simultaneously or virtually simultaneously. The plurality of images obtained from these various techniques, through angular displacement of the various acquired images, provide the information required to generate the three dimensional images needed to define the location of the radioactive seed source within the body of the living organism. | 1. A method for tracking a high intensity radioactive seed source in a living organism comprising: A) acquiring from outside of the living organism a plurality of images of the high intensity radioactive seed source in the living organism by exposing: I) a single position sensitive photomultiplier tube coupled to a low efficiency scintillator viewing the high intensity radioactive seed source in the living organism through a pinhole collimator array; II) a single position sensitive photomultiplier tube coupled to a low efficiency scintillator viewing the high intensity radioactive seed source in the living organism through a single pinhole collimator or a pinhole collimator array from a plurality of angles; or III) a plurality of position sensitive photomultiplier tubes each coupled to a low efficiency scintillator viewing the radioactive seed source in the living organism through a single pinhole collimator or a pinhole collimator array from a plurality of angles; and B) reconstructing the acquired images through computer based analysis and synthesis to produce a three dimensional image of the high intensity radioactive seed source in three dimensional space. | 1. A method for tracking a high intensity radioactive seed source in a living organism comprising: A) acquiring from outside of the living organism a plurality of images of the high intensity radioactive seed source in the living organism by exposing: I) a single position sensitive photomultiplier tube coupled to a low efficiency scintillator viewing the high intensity radioactive seed source in the living organism through a pinhole collimator array; II) a single position sensitive photomultiplier tube coupled to a low efficiency scintillator viewing the high intensity radioactive seed source in the living organism through a single pinhole collimator or a pinhole collimator array from a plurality of angles; or III) a plurality of position sensitive photomultiplier tubes each coupled to a low efficiency scintillator viewing the radioactive seed source in the living organism through a single pinhole collimator or a pinhole collimator array from a plurality of angles; and B) reconstructing the acquired images through computer based analysis and synthesis to produce a three dimensional image of the high intensity radioactive seed source in three dimensional space. 2. The method of claim 1 wherein said position sensitive photomultiplier tube is contained within a protective screen and said pinhole collimator is provided in a plug across an opening in said protective screen. 3. The method of claim 2 wherein said protective screen is cone shaped. 4. The method of claim 2 wherein said plug comprises a film or foil of tungsten. 5. The method of claim 2 wherein said scintillator comprises a plastic scintillator. 6. The method of claim 2 wherein said pinhole collimator is between about 0.5 mm and about 1.0 mm in diameter. 7. The method of claim 2 wherein said images are acquired using a plurality of position sensitive photomultiplier tubes each coupled to a plastic scintillator viewing the high intensity radioactive seed source in the living organism through a single pinhole collimator or a pinhole collimator array from a plurality of angles of less than 180°. 8. The method of claim 7 wherein said plurality of angles are below about 90°. 9. A method for tracking a high intensity radioactive seed source in a living organism comprising: A) acquiring from outside of the living organism a plurality of images of the high intensity radioactive seed source in the living organism by exposing: I) a single position sensitive photomultiplier tube coupled to a plastic scintillator viewing the high intensity radioactive seed source in the living organism through a pinhole collimator array; II) a single position sensitive photomultiplier tube coupled to a plastic _scintillator viewing the high intensity radioactive seed source in the living organism through a single pinhole collimator or a pinhole collimator array from a plurality of angles; or III) a plurality of position sensitive photomultiplier tubes each coupled to a plastic _scintillator viewing the high... | Southeastern Universities Research Assn., Inc. | A61B005/05
US06837858 | B2 | 2005-01-04 | Method and apparatus for obtaining blood for diagnostic tests | 2001-10-05 | Method and apparatus for obtaining a sample of blood from a patient for subsequent diagnostic tests, e.g., glucose monitoring. In one aspect of the invention, the method comprises the steps of: (a) placing a blood collection device over a region on the surface of the skin from which said sample is to be obtained,(b) forming a seal between said blood collection device and said surface of the skin,(c) creating a vacuum sufficient to result in said surface of the skin becoming stretched and engorged with blood,(d) triggering a lancing assembly and causing a lancet to penetrate said skin,(e) retracting said lancet,(f) withdrawing blood toward and onto a fluid collector, and(g) releasing the vacuum.;In another aspect of the invention, an apparatus for carrying out the method described previously is provided. The apparatus comprises: (a) a housing having a sealable chamber located therein and a sealable opening in fluid communication with said sealable chamber,(b) a power source,(c) a vacuum pump operably connected to said power source, said vacuum pump in communication with said sealable chamber,(d) a lancing assembly positioned within said housing, said lancing assembly capable of moving a lancet towards said sealable opening, and(e) a fluid collector positioned in said sealable chamber, said fluid collector in fluid communication with said sealable opening. | 1. An apparatus suitable for obtaining a sample of blood for analysis in a diagnostic test, comprising: (a) a housing having a sealable chamber located therein and a sealable opening in fluid communication with said sealable chamber, (b) a power source positioned in said housing, (c) a vacuum pump positioned in said housing, said vacuum pump operably connected to said power source, said vacuum pump in communication with said sealable chamber, (d) a lancing assembly positioned in said housing, said lancing assembly both capable of moving a lancet towards said sealable opening at a sufficient speed to form an opening in said skin and capable of moving said lancet away from said sealable opening, whereby said lancet forms an unobstructed opening in skin when said skin is adjacent said sealable opening, and (e) electronics in said housing for controlling power for said pump, whereby said electronics controls said pump to maintain a vacuum at a specified level. | 1. An apparatus suitable for obtaining a sample of blood for analysis in a diagnostic test, comprising: (a) a housing having a sealable chamber located therein and a sealable opening in fluid communication with said sealable chamber, (b) a power source positioned in said housing, (c) a vacuum pump positioned in said housing, said vacuum pump operably connected to said power source, said vacuum pump in communication with said sealable chamber, (d) a lancing assembly positioned in said housing, said lancing assembly both capable of moving a lancet towards said sealable opening at a sufficient speed to form an opening in said skin and capable of moving said lancet away from said sealable opening, whereby said lancet forms an unobstructed opening in skin when said skin is adjacent said sealable opening, and (e) electronics in said housing for controlling power for said pump, whereby said electronics controls said pump to maintain a vacuum at a specified level. 2. The apparatus of claim 1, wherein said housing comprises: (a) a body portion (b) a door portion, said door portion positionable over said body portion and including said sealable opening. 3. The apparatus of claim 2, wherein a fluid collector is positioned in a slot in said door portion. 4. The apparatus of claim 3, wherein said slot locates said fluid collector adjacent said sealable opening when said door portion is positioned over said body portion. 5. The apparatus of claim 2, wherein said body portion contains at least one slot for inserting a fluid collector. 6. The apparatus of claim 2, wherein said body portion includes at least one electrical connection. 7. The apparatus of claim 2, wherein said door portion includes at least one electrical connection. 8. The apparatus of claim 2, wherein said body portion comprises a movable projection containing a slot, and a fluid collector is capable of being positioned in said slot. 9. The apparatus of claim 8, wherein said movable projection includes at least one electrical connection. 10. The apparatus of claim 1, wherein said housing comprises: (a) a body portion, (b) an interior cover portion positionable over said lancing assembly, and (c) a door portion positionable over said interior cover portion and including said sealable opening. 11. The apparatus of claim 10, wherein said interior cover portion contains at least one alignment channel to position a fluid collector. 12. The apparatus of claim 10, wherein said door portion contains at least one alignment fixture to position a fluid collector. 13. The apparatus of claim 10, wherein said housing further comprises a projection. 14. The apparatus of claim 13, wherein a fluid collector is positioned in at least one slot in said projection. 15. The apparatus of claim 14, wherein said at least one slot is parallel to the upper surface of said interior cover. 16. The apparatus of claim 14, wherein said at least one slot is not parallel to the upper surface of said interior co... | Abbott Laboratories | A61B005/00
US06896873 | B2 | 2005-05-24 | MR signal-emitting coatings | 2002-03-12 | The present invention provides a coating that emits magnetic resonance signals and a method for coating medical devices therewith. The coating includes a paramagnetic metal ion-containing polymer complex that facilitates diagnostic and therapeutic techniques by readily visualizing medical devices coated with the complex. | 1. A method of making an existing medical device magnetic-resonance imageable, wherein the medical device has a surface, a portion of which is made from or coated with a solid polymer, the method comprising: treating at least a portion of the solid-polymer surface of the medical device to yield at least one surface functional group thereon, the surface functional group being an amino group or a carboxyl group; covalently linking a chelate to the surface functional group; and coordinating a paramagnetic-metal ion (Mn+) with the chelate to form a paramagnetic-metal-ion/chelate complex that is magnetic-resonance imageable, the paramagnetic-metal-ion/chelate complex being covalently linked to the medical device, wherein M is a paramagnetic ion and n is an integer that is 2 or greater. | 1. A method of making an existing medical device magnetic-resonance imageable, wherein the medical device has a surface, a portion of which is made from or coated with a solid polymer, the method comprising: treating at least a portion of the solid-polymer surface of the medical device to yield at least one surface functional group thereon, the surface functional group being an amino group or a carboxyl group; covalently linking a chelate to the surface functional group; and coordinating a paramagnetic-metal ion (Mn+) with the chelate to form a paramagnetic-metal-ion/chelate complex that is magnetic-resonance imageable, the paramagnetic-metal-ion/chelate complex being covalently linked to the medical device, wherein M is a paramagnetic ion and n is an integer that is 2 or greater. 2. The method of claim 1, wherein treating at least a portion of the solid-polymer surface comprises plasma treating at least a portion of the solid-polymer surface. 3. The method of claim 2, wherein plasma treating at least a portion of the solid-polymer surface comprises plasma treating the surface with a plasma gas which is urea, ammonia, a nitrogen-hydrogen combination or a combination thereof, and wherein the surface functional group is an amino group. 4. The method of claim 2, wherein plasma treating at least a portion of the solid-polymer surface comprises plasma treating the surface with a plasma gas which is carbon dioxide or oxygen, and wherein the surface functional group is a carboxyl group. 5. The method of claim 1, wherein the chelate is covalently linked to the surface functional group by an amide linkage. 6. The method of claim 1, wherein the polymer is selected from the group consisting of polyethylenes, polypropylenes, polyesters, polyamides, polyfluoroethylene and polyurethanes. 7. The method of claim 1, wherein M is a lanthanide or is a transition metal which is iron, manganese, chromium, cobalt or nickel. 8. A medical device capable of being magnetic-resonance imaged, the device comprising: a solid-polymer surface having functional groups linked thereto, the functional groups being amino groups or carboxyl groups; a chelate covalently linked to at least one of the functional groups on the solid-polymer surface; and a paramagnetic-metal ion coordinated with the chelate to form a paramagnetic-metal-ion/chelate complex capable of being magnetic-resonance imaged, the paramagnetic-metal-ion/chelate complex being covalently bonded to the medical device such that the complex is substantially non-absorbable by a living organism upon being inserted therein. 9. The device of claim 8, wherein the polymer is selected from the group consisting of polyethylenes, polypropylenes, polyesters, polyamides, polyfluoroethylene and polyurethanes. 10. The device of claim 8, wherein the paramagnetic metal is a lanthanide or is a transition metal which is iron, manganese, chromium, cobalt or nickel. 11. The device of claim 8, wherein the functional groups linked... | Wisconsin Alumni Research Foundation | A61B005/055
US06849403 | B1 | 2005-02-01 | Apparatus and method for drug screening | 2000-11-28 | The present invention provides kits and methods for screening drugs and drug candidates for activity by determining the presence or absence of high integrity nucleic acid in a sample. | 1. A method for screening drug activity, the method comprising the steps of: (a) determining a baseline level of high integrity nucleic acid associated with a nucleic acid locus in a first sample obtained from a patient having a disease or at least one indicium of a disease, wherein the high integrity nucleic acid is derived from a genome of the patient and the disease comprises at least one of cancer, pre-cancer, a disease associated with apoptosis, and a disease associated with programmed cell death; (b) treating the patient with a drug candidate; and (c) determining whether the high integrity nucleic acid associated with the nucleic acid locus in a second sample obtained from the patient is reduced. | 1. A method for screening drug activity, the method comprising the steps of: (a) determining a baseline level of high integrity nucleic acid associated with a nucleic acid locus in a first sample obtained from a patient having a disease or at least one indicium of a disease, wherein the high integrity nucleic acid is derived from a genome of the patient and the disease comprises at least one of cancer, pre-cancer, a disease associated with apoptosis, and a disease associated with programmed cell death; (b) treating the patient with a drug candidate; and (c) determining whether the high integrity nucleic acid associated with the nucleic acid locus in a second sample obtained from the patient is reduced. 2. The method of claim 1 wherein the determining steps comprise amplifying a target nucleic acid to determine the presence of high integrity nucleic acid in the samples. 3. The method of claim 2 wherein the target nucleic acid is amplified with one forward primer and at least two reverse primers. 4. The method of claim 2 wherein the target nucleic acid is amplified with at least two pairs of forward and reverse primers. 5. The method of claim 1 wherein the determining steps comprise capturing the high integrity nucleic acid. 6. The method of claim 5 wherein the high integrity nucleic acid is captured on a support-based complementary nucleic acid probe. 7. The method of claim 1 wherein a positive screen is determined by the presence of a lower amount of high integrity nucleic acid in the second sample relative to an amount of high integrity nucleic acid in the first sample. 8. The method of claim 1 wherein the samples are selected from the group consisting of stool, sputum, pus, blood serum, blood plasma, urine, saliva, colostrum, bile, and pancreatic juice. 9. The method of claim 7 wherein a positive screen indicates activity of the drug candidate. 10. The method of claim 7 wherein a positive screen indicates induction of programmed cell death. 11. The method of claim 7 wherein a positive screen indicates induction of apoptotic activity. 12. The method of claim 1 wherein the drug candidate comprises a nucleic acid. 13. The method of claim 1 wherein the drug candidate comprises a peptide. 14. The method of claim 1 wherein the drug candidate comprises a chemical compound. 15. The method of claim 1 wherein the drug candidate comprises a candidate for an anti-cancer drug. 16. The method of claim 1 wherein activity of a drug candidate is predictive of alleviation of disease symptoms by the drug candidate. 17. The method of claim 1 wherein at least one of the samples is contained in a buffer comprising at least 150 mM EDTA. 18. A method for screening drug activity, the method comprising the steps of: (a) determining a baseline level of high integrity nucleic acid associated with a nucleic acid locus in a sample from a subject having a disease or at least one indicium of a disease, wherein the high integrity nucleic acid is derived from ... | Exact Sciences Corporation | C12Q001/68 C12P019/34 C07H021/04 A61B005/055
US07079882 | B1 | 2006-07-18 | Method and apparatus for quantifying nerve and neural-muscular integrity related to pelvic organs or pelvic floor functions | 2000-01-22 | The present invention is a method and apparatus for quantifying nerve and neural-muscular integrity related to pelvic organs or pelvic floor functions. In one embodiment, the invention is a method and device to measure muscle or nerve activity in the pelvis or pelvic floor. In a modification of this embodiment, the invention includes means for inducing an efferent nerve signal in the spinal cord or sacral nerves and then measuring the resulting nerve activity of the nerves of the pelvic floor region. In a preferred embodiment, the invention includes placing a probe in the anal canal or vagina, the probe having proximal recording electrodes along its outer surface. In one possible embodiment, the probe has a distensible sheath having recording electrodes located on its outside surface. In another embodiment, the probe has an inflatable balloon at its distal end with recording electrodes located on the outside surface of the inflatable balloon. In a further embodiment, the proximal recording electrodes will be located on the outside of the distal end of the probe itself. In yet another embodiment, a disk having recording electrodes is placed in the rectum in contact with the tissue above the submucous space. In yet another embodiment of the invention, an essentially “C” shaped clip is provided having a core made of a spring-like material. | 1. A method of quantifying nerve and neural-muscular integrity related to pelvic organs or pelvic floor functions comprising: providing a probe comprising a body, a distensible member, and a recording electrode, wherein the distensible member has a first end connected to the body and a distal end movable relative to the body, and wherein the recording electrode is attached to the distensible member adjacent the distal end; placing the probe into the anal canal or vagina; moving the distal end of the distensible member away from the body to engage tissue of the anal canal or vagina with the recording electrode; placing a stimulation electrode, connected to a pulse generator, near the spinal cord or sacral nerves; producing electrical stimulation pulses at the stimulation electrode; and detecting the resulting electrical activity in the pelvic organs or pelvic floor by the recording electrode. | 1. A method of quantifying nerve and neural-muscular integrity related to pelvic organs or pelvic floor functions comprising: providing a probe comprising a body, a distensible member, and a recording electrode, wherein the distensible member has a first end connected to the body and a distal end movable relative to the body, and wherein the recording electrode is attached to the distensible member adjacent the distal end; placing the probe into the anal canal or vagina; moving the distal end of the distensible member away from the body to engage tissue of the anal canal or vagina with the recording electrode; placing a stimulation electrode, connected to a pulse generator, near the spinal cord or sacral nerves; producing electrical stimulation pulses at the stimulation electrode; and detecting the resulting electrical activity in the pelvic organs or pelvic floor by the recording electrode. 2. The method of claim 1 further comprising the step of timing the time between the application of a stimulation pulse by the stimulation electrode and the detection of the resulting electrical activity in the pelvic organs or pelvic floor by the recording electrode. 3. The method of claim 1 wherein moving the distal end of the distensible member away from the body to engage tissue of the anal canal or vagina comprises moving the distensible member away from the body to engage the mucosal tissue of the rectal wall, wherein the mucosal tissue of the rectal wall is pushed against the levator ani muscles. 4. The method of claim 1 wherein the distensible member comprises two or more wings each comprising a distal end, and wherein the moving step comprises moving the distal ends of the wings away from the body. 5. The method of claim 4 further comprising one or more additional recording electrodes, wherein a recording electrode is adjacent each wing, and wherein the moving step comprises moving each of the wings away from the body to engage tissue of the anal canal or vagina with the recording electrodes. 6. The method of claim 1 wherein the probe further comprises a distending member, wherein at least a portion of the distending member is positioned between the body and the distensible member, and wherein the moving step comprises distending the distending member thereby causing the distensible member to move away from the body. 7. The method of claim 6 wherein the distending member comprises a balloon, and wherein the moving step comprises expanding the balloon, wherein the balloon exerts pressure on the distensible member thereby causing the distensible member to move away from the body. 8. A method of quantifying nerve and neural-muscular integrity related to pelvic organs or pelvic floor functions comprising: providing a probe comprising a body, a distensible member, and a recording electrode, wherein the distensible member has a first end connected so he body and a distal end movable relative to the body, and wherein the recording electrode is attach... | - | A61B005/04
US07123952 | B2 | 2006-10-17 | Cardiac magnetic field diagnozer for atrial flutter and atrial fibrillation and method for identifying electric turning path of atrial flutter and atrial fibrillation | 2001-07-17 | A magnetic field distribution measurement device (1) provides a non-contact magnetic field measurement on a subject's chest at a plurality of coordinates and forms therefrom time-series magnetic field distribution data. A first arithmetic device (2) in response generates image data representing a three-dimensional, intramyocardial current density distribution. A second arithmetic device (3) receives a plurality of tomographic image data separately obtained by a tomographic diagnosis apparatus and processes the data to generate three-dimensional, anatomical image data. A display device (4) receives these data and displays on an anatomical image an image representing an intramyocardial current density. This can facilitate identifying an anatomical, positional relationship of an abnormal, electrical reentry circuit caused in heart muscle. Furthermore, the anatomical image may be replaced with an image representing a normal stimulation propagation circuit and serving as a template. | 1. A magnetocardiographic diagnosis apparatus for atrial flutter and fibrillation, comprising: a magnetic field distribution measurement device (1) performing a non-contact magnetic field measurement on a subject's chest at a plurality of coordinates to obtain a plurality of time-series magnetic field data corresponding to said plurality of coordinates, and also using said plurality of time-series magnetic field data to generate time-series magnetic field distribution data on said chest;an arithmetic device (5) using said generated time-series magnetic field distribution data to generate data representative of a three-dimensional, intramyocardial, electrical behavior of said subject; anda display device (6) using the data generated by said arithmetic device to superimpose together an image representing a stimulation propagation path of said subject extending from a sinoatrial node to a bundle of His-Purkinje fiber network and an image representing an abnormal, intramyocardial, electrical reentry circuit and display said images, thereby capable of three-dimensionally identifying an abnormal, intramyocardial, electrical reentry circuit. | 1. A magnetocardiographic diagnosis apparatus for atrial flutter and fibrillation, comprising: a magnetic field distribution measurement device (1) performing a non-contact magnetic field measurement on a subject's chest at a plurality of coordinates to obtain a plurality of time-series magnetic field data corresponding to said plurality of coordinates, and also using said plurality of time-series magnetic field data to generate time-series magnetic field distribution data on said chest;an arithmetic device (5) using said generated time-series magnetic field distribution data to generate data representative of a three-dimensional, intramyocardial, electrical behavior of said subject; anda display device (6) using the data generated by said arithmetic device to superimpose together an image representing a stimulation propagation path of said subject extending from a sinoatrial node to a bundle of His-Purkinje fiber network and an image representing an abnormal, intramyocardial, electrical reentry circuit and display said images, thereby capable of three-dimensionally identifying an abnormal, intramyocardial, electrical reentry circuit. 2. The apparatus of claim 1, wherein said data generated by said arithmetic device and representative of said three-dimensional, intramyocardial, electrical behavior is intramyocardial, time-series current density distribution data and wherein said display device three-dimensionally displays a location of a plurality of abnormal, electrical reentry circuits on said image of said stimulation propagation path, as based on said time-series current density distribution data. 3. A method of identifying an electrical reentry circuit for atrial flutter and fibrillation, comprising the steps of: performing a non-contact magnetic field measurement on a subject's chest at a plurality of coordinates to obtain a plurality of time-series magnetic field data corresponding to said plurality of coordinates and used to generate time-series magnetic field distribution data of said chest and generating first data representative of a three-dimensional, intramyocardial, electrical behavior of said subject from the generated time-series magnetic field distribution data;processing separately fed, tomographic, thoracic image data of said subject to generate second data representative of an anatomical image; anddisplaying an image of said three-dimensional, intramyocardial, electrical behavior represented by said first data, as superimposed on said anatomical image represented by said second data, to permit an abnormal, intramyocardial, electrical reentry circuit to be three-dimensionally identified. 4. The method of claim 3, wherein said three-dimensional, intramyocardial, electrical behavior represented by said first data is an intramyocardial current density distribution. 5. A method of identifying an electrical reentry circuit for atrial flutter and fibrillation, comprising the steps of: performing a non-contact magnetic field measu... | Japan Science and Technology Corporation Takenaka Corporation | A61B005/04
US06997879 | B1 | 2006-02-14 | Methods and devices for reduction of motion-induced noise in optical vascular plethysmography | 2002-07-09 | Methods and devices are provided for reducing motion artifacts when monitoring volume changes in blood vessels. Light having a first wavelength and light having a second wavelength are transmitted through a human appendage, toward the epidermis of a patient, or through tissue within the body of a patient. A portion of the light having the first wavelength and a portion of the light having the second wavelength is received. A first signal is produced based on the received portion of light having the first wavelength. A second signal is produced based on the received portion of light having the second wavelength. One of the first and second signals is subtracted from the other to produce a plethysmography signal that is representative of volume changes in blood vessels of patient tissue. | 1. A method for reducing motion artifacts when monitoring volume changes in blood vessels, comprising the following steps: (a) transmitting, from a light source implanted within a patient's body, light having a first wavelength and light having a second wavelength toward tissue within the patient's body;(b) receiving a portion of the light having the first wavelength and a portion of the light having the second wavelength reflected from the tissue;(c) producing a first signal, based on the received portion of light having the first wavelength, and a second signal, based on the received portion of light having the second wavelength;(d) subtracting one of the first and second signals from the other to produce a plethysmography signal that is representative of volume changes in blood vessels of the tissue; and(e) setting intensities such that a ratio of an intensity of the light having the first wavelength over an intensity of the light having the second wavelength is approximately equal to a ratio of an effective absorption coefficient of non-blood tissue at the second wavelength over an effective absorption coefficient of non-blood tissue at the first wavelength. | 1. A method for reducing motion artifacts when monitoring volume changes in blood vessels, comprising the following steps: (a) transmitting, from a light source implanted within a patient's body, light having a first wavelength and light having a second wavelength toward tissue within the patient's body;(b) receiving a portion of the light having the first wavelength and a portion of the light having the second wavelength reflected from the tissue;(c) producing a first signal, based on the received portion of light having the first wavelength, and a second signal, based on the received portion of light having the second wavelength;(d) subtracting one of the first and second signals from the other to produce a plethysmography signal that is representative of volume changes in blood vessels of the tissue; and(e) setting intensities such that a ratio of an intensity of the light having the first wavelength over an intensity of the light having the second wavelength is approximately equal to a ratio of an effective absorption coefficient of non-blood tissue at the second wavelength over an effective absorption coefficient of non-blood tissue at the first wavelength. 2. The method of claim 1 wherein the setting intensities step comprises setting the intensities of the light transmitted at step (a). 3. A method for reducing motion artifacts when monitoring volume changes in blood vessels, comprising the following steps of: (a) transmitting light having a first wavelength and light having a second wavelength through a human appendage;(b) setting the transmitted intensity of at least one of the light having the first wavelength and the light having the second wavelength such that a ratio of an intensity of the light having the first wavelength over an intensity of the light having the second wavelength is approximately equal to a ratio of an effective absorption coefficient of non-blood tissue at the second wavelength over an effective absorption coefficient of non-blood tissue at the first wavelength;(c) receiving a portion of the light having the first wavelength and a portion of the light having the second wavelength;(d) producing a first signal, based on the received portion of light having the first wavelength, and a second signal, based on the received portion of light having the second wavelength; and(e) subtracting one of the first and second signals from the other to produce a plethysmography signal that is representative of volume changes in blood vessels of the human appendage. 4. The method of claim 3, wherein the setting intensities step comprises setting the intensities of the light transmitted at step (a). 5. A method for reducing motion artifacts when monitoring volume changes in blood vessels, comprising the following steps: (a) applying light having a first wavelength and light having a second wavelength to an epidermis of a patient;(b) detecting, at the epidermis of the patient, a portion of the light having the first wavelength an... | Pacesetter, Inc. | A61B005/02 A61B005/00
US07179449 | B2 | 2007-02-20 | Enhanced ultrasound detection with temperature-dependent contrast agents | 2001-01-30 | Methods and devices for enhanced ultrasound detection based upon changing temperature and ultrasound reflectivity of a temperature-dependent contrast agent bound to an ultrasound target are disclosed. The methods and devices can be used for enhanced imaging alone or in conjunction with drug delivery, with therapeutic approaches such as hyperthermia or cryotherapy or with other imaging modalities. | 1. A method for comparing acoustic reflectivity of a target for ultrasound imaging at a lower and a higher temperature, the method comprising (a) measuring reflectivity prior to raising the temperature of liquid nanoparticles bound to a target;(b) raising the temperature of the liquid nanoparticles bound to said target sufficiently to produce a measurable enhancement in acoustic reflectivity of the target;(c) measuring reflectivity after raising the temperature of the bound liquid nanoparticles; and(d) determining the change in reflectivity of the bound liquid nanoparticles after raising the temperature compared to reflectivity prior to raising the temperature,wherein said nanoparticles comprise at least one fluorocarbon,said nanoparticles having been administered to said target in a non-gaseous emulsion, and wherein said nanoparticles are maintained in the liquid state during all of steps (a)–(c). | 1. A method for comparing acoustic reflectivity of a target for ultrasound imaging at a lower and a higher temperature, the method comprising (a) measuring reflectivity prior to raising the temperature of liquid nanoparticles bound to a target;(b) raising the temperature of the liquid nanoparticles bound to said target sufficiently to produce a measurable enhancement in acoustic reflectivity of the target;(c) measuring reflectivity after raising the temperature of the bound liquid nanoparticles; and(d) determining the change in reflectivity of the bound liquid nanoparticles after raising the temperature compared to reflectivity prior to raising the temperature,wherein said nanoparticles comprise at least one fluorocarbon,said nanoparticles having been administered to said target in a non-gaseous emulsion, and wherein said nanoparticles are maintained in the liquid state during all of steps (a)–(c). 2. The method according to claim 1 wherein the fluorocarbon is perfluorooctane. 3. The method according to claim 1 wherein the nanoparticles comprise at least one liquid fluorocarbon encapsulated with at least one lipid surfactant which comprises at least one ligand that binds to said target. 4. The method according to claim 1 wherein the emulsion further comprises a biologically active agent. 5. The method according to claim 1 wherein raising the temperature comprises providing the target with ultrasound or electromagnetic energy or a combination thereof, sufficient to raise the temperature of said nanoparticles, so as to enhance acoustic reflectivity. 6. The method according to claim 1 wherein changing the temperature comprises changing the temperature of the bound nanoparticles by at least 5° C. 7. The method of claim 1 wherein the target resides in a mammalian subject. 8. The method of claim 1 wherein said nanoparticles comprise at least one liquid fluorocarbon encapsulated with at least one lipid surfactant. 9. The method according to claim 3 wherein the ligand is a polypeptide, a peptidomimetic, a polysaccharide, a lipid, or a nucleic acid. 10. The method according to claim 9 wherein the polypeptide is at least a portion of an antibody. 11. The method of claim 7 wherein said subject is human. 12. A method for obtaining an image resulting from enhanced acoustic reflectivity of a target for ultrasound imaging, the method comprising (a) changing the temperature of liquid nanoparticles bound to said target sufficiently to produce a measurable enhancement of acoustic reflectivity of the target, and(b) obtaining an ultrasound image of said target, bound to said liquid nanoparticles,wherein said nanoparticles comprise at least one fluorocarbon, said nanoparticles having been administered to said target in a non-gaseous emulsion; andwherein said nanoparticles are maintained in the liquid state during both of steps (a) and (b). 13. The method according to claim 12 wherein the fluorocarbon is perfluorooctane. 14. The method according to claim ... | Barnes-Jewish Hospital | A61B005/055 A61B008/00 A61K009/127
US07008384 | B2 | 2006-03-07 | Sensor controlled analysis and therapeutic delivery system | 2002-09-25 | A method and apparatus for non-invasively withdrawing and accurately evaluating analytes quickly, painlessly and reliably from a biological subject automatically and controlling subsequent administration of therapeutic agents in response to such analyte sample analysis. Improvements are provided in electro-osmotic sample withdrawal, dosimetry and iontophoretic delivery subsystems, biosensors electrode construction and arrangement, intervenors, mediators, bolus delivery, and related subsystems. | 1. In a compact patch adapted to be worn by a biological subject, the combination comprising: first means within said patch for electro-osmotic withdrawal of fluid samples;second biosensor means within the same said patch for evaluation of said samples; andthird means within the same said patch for providing iontophoretic treatment to a biological subject in response to said biosensor means whereby said compact patch is provided including said first, second and third means in an integral package. | 1. In a compact patch adapted to be worn by a biological subject, the combination comprising: first means within said patch for electro-osmotic withdrawal of fluid samples;second biosensor means within the same said patch for evaluation of said samples; andthird means within the same said patch for providing iontophoretic treatment to a biological subject in response to said biosensor means whereby said compact patch is provided including said first, second and third means in an integral package. | - | A61B005/00 B65D081/00
US07063667 | B1 | 2006-06-20 | Isotopic gas analyzer | 1998-09-17 | An NDIR spectrometer based on the use of wavelength specific lamp sources, whose emission spectrum consists of discrete, narrow lines characteristic of the isotope present in the lamp, and which it is desired to measure with the spectrometer. This allows very high intrinsic sensitivity, enabling the use of an extremely compact absorption cell with a very short path length. In addition, the source can be self-modulated, such that problems associated with external choppers are avoided. Furthermore, there is insignificant cross sensitivity between the isotopes themselves and between the isotopes and other ambient gases in the operating environment. | 1. An isotopic gas analyzer for determining the ratio of at least a first and a second isotopic component of a gas, comprising: at least one wavelength-stable source of radiation of wavelengths characteristic of at least one of said at least first and second isotopic components;at least a first sample chamber comprising a first sample of said gas to be analyzed;at least a first reference chamber containing a reference gas comprising said first and second isotopic components;at least a second sample chamber comprising a second sample of said gas to be analyzed;at least a second reference chamber containing a reference gas comprising said first and second isotopic components; andat least one detector detecting transmission through said at least first sample chamber and transmission through said at least first reference chamber, of radiation of wavelengths characteristic of said first isotopic component of said gas, and transmission through said at least second sample chamber and transmission through said at least second reference chamber, of radiation of wavelengths characteristic of said second isotopic component of said gas. | 1. An isotopic gas analyzer for determining the ratio of at least a first and a second isotopic component of a gas, comprising: at least one wavelength-stable source of radiation of wavelengths characteristic of at least one of said at least first and second isotopic components;at least a first sample chamber comprising a first sample of said gas to be analyzed;at least a first reference chamber containing a reference gas comprising said first and second isotopic components;at least a second sample chamber comprising a second sample of said gas to be analyzed;at least a second reference chamber containing a reference gas comprising said first and second isotopic components; andat least one detector detecting transmission through said at least first sample chamber and transmission through said at least first reference chamber, of radiation of wavelengths characteristic of said first isotopic component of said gas, and transmission through said at least second sample chamber and transmission through said at least second reference chamber, of radiation of wavelengths characteristic of said second isotopic component of said gas. 2. An isotopic gas analyzer according to claim 1, and wherein said at least one detector is a single detector. 3. An isotopic gas analyzer according to claim 2, and wherein said single detector detects said transmission through said chambers simultaneously at wavelengths characteristic of said first isotopic component of said gas, and of said second isotopic component of said gas. 4. An isotopic gas analyzer according to claim 2, and wherein said at least one wavelength-stable source of radiation comprises at least two wavelength-stable sources of radiation, each of wavelengths characteristic of one isotopic component of said gas. 5. An isotopic gas analyzer according to claim 4, said at least two wavelength-stable sources of radiation operating with respective at least first and second different timing characteristics, said analyzer also comprising a detection differentiator receiving an output from said at least one detector and distinguishing outputs corresponding to said at least two wavelength-stable sources on the basis of said first and second different timing characteristics. 6. An isotopic gas analyzer according to claim 5 and also comprising a second detector, wherein said single detector views absorption signal outputs from said at least two wavelength-stable sources and said second detector views zero calibration signals from said at least two wavelength-stable sources. 7. An isotopic gas analyzer according to claim 2, and wherein said at least one wavelength-stable source of radiation is one wavelength-stable source of wavelengths characteristic of at least said first and second isotopic components. 8. An isotopic gas analyzer according to claim 2, and wherein said at least one wavelength-stable source is at least one gas discharge lamp. 9. An isotopic gas analyzer according to claim 2 and wherein at leas... | Oridion BreathID Ltd. | A61B005/00
US06993383 | B2 | 2006-01-31 | Anomaly detection based on signal variations | 2002-05-23 | A method of analyzing the type of tissue of a body portion of a subject. The method includes applying an excitation signal to the subject, sensing electrical signals from the subject responsive to the excitation signal, determining a stationary impedance map of the body portion responsive to the sensed signals, determining a change waveform of at least a part of the stationary impedance map, and providing an indication on a tissue type of at least part of the body portion, responsive to the change waveform. | 1. A method of analyzing the type of tissue of a body portion of a subject, comprising: applying an excitation signal to the subject; sensing electrical signals from the subject responsive to the excitation signal; determining a stationary impedance map of the body portion responsive to the sensed signals; determining, for at least one pixel of the map, a waveform of a change over time in a parameter of the sensed electrical signals corresponding to the at least one pixel, which change waveform corresponds to a related event in the subject that is at least partially due to application of an external stimuli to the subject; and providing an indication on a tissue type of at least part of the body portion, responsive to the change waveform. | 1. A method of analyzing the type of tissue of a body portion of a subject, comprising: applying an excitation signal to the subject; sensing electrical signals from the subject responsive to the excitation signal; determining a stationary impedance map of the body portion responsive to the sensed signals; determining, for at least one pixel of the map, a waveform of a change over time in a parameter of the sensed electrical signals corresponding to the at least one pixel, which change waveform corresponds to a related event in the subject that is at least partially due to application of an external stimuli to the subject; and providing an indication on a tissue type of at least part of the body portion, responsive to the change waveform. 2. A method according to claim 1, comprising generating a change map which is a map of a parameter of the change waveform for the at least part of the stationary map. 3. A method according to claim 2, wherein generating the change map of a parameter of the change waveform comprises generating a change map of the amplitude of the change waveform. 4. A method according to claim 2, comprising displaying the change map. 5. A method according to claim 2, comprising correlating the change map and the stationary map. 6. A method according to claim 5, comprising displaying the change map overlaid on the stationary map. 7. A method according to claim 2, comprising determining areas of increased perfusion of blood in the body portion, responsive to the change map. 8. A method according to claim 1, wherein sensing the signals comprises sensing from a plurality of points on the surface of the body portion and determining the change waveform for each of the plurality of points. 9. A method according to claim 1, wherein sensing the signals comprises sensing electrical currents. 10. A method according to claim 1, wherein sensing the signals comprises sensing electrical voltages. 11. A method according to claim 1, wherein applying signals to the subject comprises applying electrification signals from at least two electrodes which are electrified with signals of different phases. 12. A method according to claim 1, wherein the application of the external stimuli to the subject comprises injecting a stimuli to the blood stream of the subject. 13. A method according to claim 1, wherein the application of the external stimuli to the subject comprises applying pressure to the body portion of the subject. 14. A method according to claim 1, wherein providing the indication comprises filtering at least one frequency from the change waveform and providing the indication responsive to the filtering waveform. analyzing the type of tissue of a body portion of a subject, comprising: applying an excitation signal to the subject; sensing signals from the subject responsive to the excitation signal; injecting a stimulus to the blood stream of the subject; extracting from the sensed signals a time dependent change waveform responsi... | Mirabel Medical Systems Ltd. | A61B005/05
US07276027 | B2 | 2007-10-02 | System, for monitoring the concentration of analytes in body fluids | 2002-02-02 | The invention concerns a system for monitoring the concentration of analytes in body fluids, in particular in interstitial fluid and comprises a catheter having an implantable region and an outlet opening for withdrawing fluid in particular body fluid. A first and a second analytical zone are contacted sequentially with fluid from the catheter and undergo a detectable change when an analyte is present. The analytical zones can be contacted manually with fluid and also preferably in an automated fashion by means of a device. A system according to the invention additionally has an analytical device for analysing the analytical zones in order to determine the concentration of the analyte on the basis of changes caused by the analyte. A further subject matter of the present invention are catheters for use in systems according to the invention as well as magazines containing test zones. | 1. System for monitoring the concentration of analytes in body fluids, comprising a) a catheter having an implantable region and an outlet opening for the withdrawal of fluid,b) a first and a second analytical zone which, after contact with the withdrawn fluid, undergo a detectable change when an analyte is present in the fluid,c) a device configured to contact the first analytical zone with fluid from the catheter and to subsequently contact the second analytical zone with fluid from the catheter,d) an analytical device configured to analyze the changes on the analytical zones caused by the analyte-containing fluid in order to determine the concentration of an analyte to be monitored, ande) a pump configured to apply underpressure to the outlet opening in order to convey liquid,wherein the analytical device or an additional application detection device detects the presence of fluid or the presence of an adequate amount of fluid in or on the analytical zones and interrupts the application of the underpressure to the outlet opening based thereon,and wherein the first and second analytical zones are areas of a continuous test element, and the continuous test element is a tape. | 1. System for monitoring the concentration of analytes in body fluids, comprising a) a catheter having an implantable region and an outlet opening for the withdrawal of fluid,b) a first and a second analytical zone which, after contact with the withdrawn fluid, undergo a detectable change when an analyte is present in the fluid,c) a device configured to contact the first analytical zone with fluid from the catheter and to subsequently contact the second analytical zone with fluid from the catheter,d) an analytical device configured to analyze the changes on the analytical zones caused by the analyte-containing fluid in order to determine the concentration of an analyte to be monitored, ande) a pump configured to apply underpressure to the outlet opening in order to convey liquid,wherein the analytical device or an additional application detection device detects the presence of fluid or the presence of an adequate amount of fluid in or on the analytical zones and interrupts the application of the underpressure to the outlet opening based thereon,and wherein the first and second analytical zones are areas of a continuous test element, and the continuous test element is a tape. 2. System as claimed in claim 1, wherein the device is configured to contact the first analytical zone and the second analytical zone with the fluid from the catheter by bringing together the outlet opening and the first and second analytical zones. 3. System as claimed in claim 1, wherein the device is configured to contact the first analytical zone and the second analytical zone with the fluid from the catheter by moving portions of the fluid out of the catheter onto the first and second analytical zones via an ejector unit. 4. System as claimed in claim 1, in which the first and second analytical zones are separate objects that are attached to a common support. 5. System as claimed in claim 1, in which the catheter is designed such that no liquid emerges from the outlet opening until the underpressure is applied to the outlet opening. 6. System as claimed in claim 1, further comprising a control device which synchronizes the bringing together of the first and second analytical zones with the outlet opening. 7. System as claimed in claim 1, in which the amount of fluid taken up by either of the first and second analytical zones is essentially equal to or more than the active inner volume of the catheter. 8. System as claimed in claim 1, wherein the device is configured to contact the first analytical zone and the second analytical zone with the fluid from the catheter while the catheter remains implanted. 9. System as claimed in claim 1, wherein the analytical zones are configured to be used once. 10. System as claimed in claim 1, in which the system further comprises a carrying unit for carrying the system on the body and a magazine in which the analytical zones are disposed and which is configured to be inserted into the carrying unit. 11. System as claimed in c... | Roche Diagnostics Operations, Inc. | A61B005/00
US07016715 | B2 | 2006-03-21 | Selection of preset filter parameters based on signal quality | 2003-01-13 | Methods and devices for reducing noise effects in a system for measuring a physiological parameter, including receiving an input signal; obtaining an assessment of the signal quality of the input signal; selecting coefficients for a digital filter using the assessment of signal quality; and filtering the input signal using the digital filter, without comparing the filter's output signal with the input signal. | 1. In a system for measuring a physiological parameter, a method of reducing noise effects, comprising: receiving input signals of more than one wavelength;obtaining an assessment of the signal quality of the input signals by obtaining a measure of the correlation between the signals of different wavelengths;selecting coefficients for a digital filter using the assessment of signal quality; andfiltering the input signals using the digital filter, without comparing the filter's output signal with the input signals. | 1. In a system for measuring a physiological parameter, a method of reducing noise effects, comprising: receiving input signals of more than one wavelength;obtaining an assessment of the signal quality of the input signals by obtaining a measure of the correlation between the signals of different wavelengths;selecting coefficients for a digital filter using the assessment of signal quality; andfiltering the input signals using the digital filter, without comparing the filter's output signal with the input signals. 2. The method of claim 1 wherein selecting comprises selecting coefficients from a plurality of discrete preset values. 3. The method of claim 1 wherein selecting comprises selecting coefficients from a plurality of constant and preset values. 4. The method of claim 1 wherein filtering comprises operating a digital filter having a linear input-output relationship. 5. The method of claim 1 wherein filtering comprises operating a digital filter having a non-linear input-output relationship. 6. The method of claim 1 wherein selecting is performed in real time, with the coefficients of the digital filter being determined using a current input sample. 7. The method of claim 1 wherein selecting is performed using a previously stored input signal sample. 8. The method of claim 1 wherein receiving comprises receiving signals from an oximeter, wherein the signals are a function of an oxygen saturation. 9. The method of claim 1 wherein receiving comprises receiving signals from an oximeter, wherein the signals are a function of a pulse rate. 10. The method of claim 1 wherein receiving the input signals comprises receiving signals corresponding to sensed optical energies from a plurality of wavelengths. 11. In a system for a measuring a physiological parameter, an apparatus for reducing noise effects, comprising: means for receiving input signals of more than one wavelength;means for obtaining an assessment of the signal quality of the input signals by obtaining a measure of the correlation between the signals of different wavelengths;means for selecting coefficients for a digital filter using the assessment of signal quality; andmeans for filtering the input signals using the digital filter, without comparing the filter's output signal with said input signals. 12. The apparatus of claim 11 wherein the means for selecting are configured to select the coefficients from a plurality of discrete preset values. 13. The apparatus of claim 11 wherein the means for selecting are configured to select the coefficients from a plurality of constant and preset values. 14. The apparatus of claim 11 wherein the means for filtering comprise a digital filter having a linear input-output relationship. 15. The apparatus of claim 11 wherein the means for filtering comprise a digital filter having a non-linear input-output relationship. 16. The apparatus of claim 11 wherein the means for selecting are configured to select the coefficients in real time, ... | NellcorPuritan Bennett Incorporated | A61B005/00
US06986746 | B2 | 2006-01-17 | Biased vascular temperature measuring device | 2002-05-03 | The present invention provides a vascular catheter apparatus for temperature measurement of vascular tissue, comprising a flexible body, at least one resiliently biased projection depended from the body, a sensor carried by the projection and an electrical carrier for transmitting temperature data at the vascular wall from the sensor to a remote device. The electrical carrier may be coiled around the length of the projection to reduce the strain at critical points where it is necessary to maintain a seal and hence electrical isolation. | 1. A catheter apparatus for temperature measurement of vascular tissue, comprising a flexible body, at least one resiliently biased projection depended from the body, a sensor carried by the projection and an electrical carrier for transmitting data from the sensor to a remote device, wherein the electrical carrier is coiled around the body of the projection in a manner so as to reduce strain when the projection flexes. | 1. A catheter apparatus for temperature measurement of vascular tissue, comprising a flexible body, at least one resiliently biased projection depended from the body, a sensor carried by the projection and an electrical carrier for transmitting data from the sensor to a remote device, wherein the electrical carrier is coiled around the body of the projection in a manner so as to reduce strain when the projection flexes. 2. An apparatus according to claim 1, in which the pitch of the coil is arranged such that there are 5 to 10 turns per cm over the length of the projection. 3. An apparatus according to claim 1, further comprising a heat shrink wrapping applied around at least a portion of the length of the projection. 4. An apparatus according to claim 1, wherein the body has a longitudinal axis and at least part of the projection is extensible radially from the longitudinal axis of the catheter body. 5. An apparatus according to claim 1, wherein the body has a longitudinal axis and at least part of the projection is extensible radially from the longitudinal axis of the catheter body. 6. An apparatus according to claim 5, wherein the ratio of the area of the cross-sectional profiles of the apparatus in the retracted and deployed configurations is in the range 4:1-1.1:1, preferably 3:1-1.25:1, more preferably 2.5:1-2.0:1, most preferably 1.75:1-1.25:1. 7. An apparatus according to claim 1, further comprising multiple sensors and wherein the sensors are positioned circumferentially about a longitudinal axis of the catheter body. 8. An apparatus according to claim 1, wherein the sensor is a thermistor. 9. An apparatus according to claim 1, wherein the projections comprises a superelastic material. 10. An apparatus according to claim 1, further comprising a sheath adapted to substantially encompass the sensor in a retracted configuration. 11. An apparatus according to claim 10, wherein the sheath is retractable, thus enabling the projection to take up a deployed configuration. 12. An apparatus according to claim 1, wherein the projection, in a deployed configuration, adopts an arcuate shape along at least part of its length, and wherein the gradient of the arcuate portion of the projection, with respect to the longitudinal axis of the catheter, increases as a function of distance along the projection from the end attached to the catheter body. | Thermocore Medical Systems NV | A61B005/00
US07020917 | B1 | 2006-04-04 | Radiolucent surgical table with low shadow accessory interface profile | 2001-03-12 | A radiolucent surgical table (40) is provided with a low shadow accessory interface profile (70). The interface profile includes a plurality of interconnected curved surfaces formed along the entire outer edge of the patient support member for selective attachment to a range of medical accessory devices. The support interface includes a non-planar first connection area (72) having a recess (82) and a curved lip area (90) adapted to engage correspondingly formed surfaces (122, 126) on a medical accessory connector (110, 140). A second connection area (76) of the support interface (70) includes a planar locating surface (100) formed at an angle oblique with the top surface (74) of the patient support member. Engagement with the locating surface by corresponding surfaces formed on the medical accessory connectors (110, 140) generate a downward force that causes an increase in the engagement forces between the recess and lip areas of the interface and corresponding surfaces on the accessory connectors. | 1. In a radiolucent patient support table including substantially planar top and bottom surfaces held apart in an opposed relationship, a medical appliance support interface for selectively connecting an associated medical appliance to the table, the interface comprising: a non-planar first connection area defined by a first portion of said top surface of the support table, the first connection area being shaped to provide a first supporting force against an associated medical appliance in a first direction substantially parallel to the top and bottom surfaces, and a second supporting force against the associated medical appliance in a second direction substantially perpendicular to the top and bottom surfaces, wherein said first connection area includes at least one recess defined between a pair of wall surfaces that converge at a bight of the at least one recess;a second connection area defined by a second portion of said table and shaped to provide a third supporting force against the associated medical appliance in a third direction substantially parallel to the top and bottom surfaces, and a fourth supporting force against the associated medical appliance in a fourth direction substantially perpendicular to the top and bottom surfaces, wherein said second connection area includes a substantially planar surface held at an oblique angle relative to said substantially planar top and bottom surfaces; and,a rounded lip area formed between said at least one recess and said substantially planar surface, the rounded lip area defining a crest located between said bight and a plane defined by the top surface of the patient support table. | 1. In a radiolucent patient support table including substantially planar top and bottom surfaces held apart in an opposed relationship, a medical appliance support interface for selectively connecting an associated medical appliance to the table, the interface comprising: a non-planar first connection area defined by a first portion of said top surface of the support table, the first connection area being shaped to provide a first supporting force against an associated medical appliance in a first direction substantially parallel to the top and bottom surfaces, and a second supporting force against the associated medical appliance in a second direction substantially perpendicular to the top and bottom surfaces, wherein said first connection area includes at least one recess defined between a pair of wall surfaces that converge at a bight of the at least one recess;a second connection area defined by a second portion of said table and shaped to provide a third supporting force against the associated medical appliance in a third direction substantially parallel to the top and bottom surfaces, and a fourth supporting force against the associated medical appliance in a fourth direction substantially perpendicular to the top and bottom surfaces, wherein said second connection area includes a substantially planar surface held at an oblique angle relative to said substantially planar top and bottom surfaces; and,a rounded lip area formed between said at least one recess and said substantially planar surface, the rounded lip area defining a crest located between said bight and a plane defined by the top surface of the patient support table. 2. The medical appliance support interface according to claim 1 wherein said rounded lip area is spaced apart from the plane defined by the top surface by a predetermined distance. 3. A surgical table comprising: a base member;a column connected with the base member;a rectangular radiolucent patient support member carried on the column, the patient support member defining substantially planar top and bottom surfaces; and,a low radiographic shadow accessory connection interface defined by a plurality of curved surfaces of the patient support member along at least one edge of the patient support member for selectively connecting an associated accessory to the patient support member, the plurality of curved surfaces being without planar portions oriented in a substantially perpendicular relation to said planar top surface of the patient support member so that first portions of an associated x-ray signal passing through the connection interface along a path substantially perpendicular to the planar top surface are attenuated substantially the same as second portions of the x-ray signal passing through the patient support member. 4. The surgical table according to claim 3 wherein said plurality of curved surfaces of the low shadow connection interface includes a first connection area adjacent the planar top surface of t... | Steris Corporation | A61G013/10 A61B005/05
US07018346 | B2 | 2006-03-28 | Guide wire with adjustable flexibility | 2001-12-18 | A guide wire with a distal portion having adjustable flexibility. The guide wire may include a distal polymeric member and a heat source. The heat source may be activated to cause the polymeric member to increase in temperature and increase in flexibility. The increase in flexibility of the distal portion of the guide wire enhances the ability of the guide wire to navigate tortuous vasculature to a target site. After the guide wire has been navigated to the target site, the heat source may be deactivated to cause the polymeric member to decrease in temperature and increase in stiffness. The increase in stiffness of the distal portion of the guide wire enhances support provided for devices (e.g., catheters) advanced thereon. | 1. A variable stiffness guide wire, comprising: a shaft having a proximal portion and a distal portion, the distal portion having a flexibility;a polymeric member disposed on and non-releasably attached to the distal portion of the shaft, the polymeric member having a first flexibility at a first temperature and a second flexibility at a second temperature, wherein the first temperature is less than the second temperature and the first flexibility is less than the second flexibility; anda heat source disposed on the distal portion of the shaft, the heat source being in thermal communication with the polymeric member, whereby activation of the heat source causes the polymeric member to rise from the first temperature to the second temperature to thereby change the flexibility of the distal portion of the guide wire while the polymeric member remains attached to the shaft. | 1. A variable stiffness guide wire, comprising: a shaft having a proximal portion and a distal portion, the distal portion having a flexibility;a polymeric member disposed on and non-releasably attached to the distal portion of the shaft, the polymeric member having a first flexibility at a first temperature and a second flexibility at a second temperature, wherein the first temperature is less than the second temperature and the first flexibility is less than the second flexibility; anda heat source disposed on the distal portion of the shaft, the heat source being in thermal communication with the polymeric member, whereby activation of the heat source causes the polymeric member to rise from the first temperature to the second temperature to thereby change the flexibility of the distal portion of the guide wire while the polymeric member remains attached to the shaft. 2. A variable stiffness guide wire as in claim 1, wherein the polymeric member comprises a shape memory polymer. 3. A variable stiffness guide wire as in claim 2, wherein the shape memory polymer has a glass transition temperature, and wherein the first temperature is below the glass transition temperature. 4. A variable stiffness guide wire as in claim 3, wherein the second temperature is above the glass transition temperature. 5. A variable stiffness guide wire as in claim 1, wherein the heat source comprises a resistive heating element. 6. A variable stiffness guide wire as in claim 5, wherein the distal portion of the shaft includes a tip portion. 7. A variable stiffness guide wire as in claim 6, wherein the tip portion includes the polymeric member and resistive heating element. 8. A variable stiffness guide wire as in claim 7, wherein the polymeric member comprises a tube. 9. A variable stiffness guide wire as in claim 8, wherein the resistive heating element comprises a coiled wire. 10. A variable stiffness guide wire as in claim 9, further comprising one or more lead wires connected to and extending proximally from the coiled wire. 11. A variable stiffness guide wire as in claim 10, wherein the polymeric tube is disposed on the coiled wire. 12. A variable stiffness guide wire as in claim 11, wherein a core wire extends through the tip portion. 13. A variable stiffness guide wire as in claim 12, wherein the proximal portion of the shaft includes a hypotube. 14. A variable stiffness guide wire as in claim 13, wherein the distal portion of the shaft includes a slotted hypotube. 15. A variable stiffness guide wire as in claim 14, wherein the coiled wire is disposed in the slots of the slotted hypotube. 16. A variable stiffness guide wire, comprising: a shaft including a proximal portion and a distal portion having a flexibility;a polymeric member disposed on and non-releasably attached to the distal portion of the shaft;a heat source in thermal communication with the polymeric member, whereby activation of the heat source causes the polymeric member to change t... | Scimed Life Systems, Inc. | A61B005/00 A61M020/00
US06988989 | B2 | 2006-01-24 | Patient monitoring system | 2001-05-18 | A medical monitoring system includes a central station adapted to receive vital signs data concerning a plurality of patients, and one or more transmitter/receivers associated with the central station. At least one patient monitor is configured to monitor a particular one of the plurality of patients by collecting vital signs data from the particular patient, and configured to establish communications with the central station and communicate the collected vital signs data to the central station via the one or more transmitter/receivers. The at least one patient monitor is operable by a user to identify the particular patient from the plurality of patients, and to inform the central station of the identity of the particular patient. The central station is configured to associate the vital signs data received from the at least one patient monitor with the particular patient. The system, transmitter/receivers and at least one patient monitor may be wireless or wired. There is also a method for monitoring a patient. | 1. A medical monitoring system, comprising: a central station adapted to receive vital signs data concerning a plurality of patients; one or more transmitter/receivers associated with the central station; and at least one patient monitor configured to monitor a particular one of the plurality of patients by collecting vital signs data from the particular patient, and configured to establish communications with the central station and communicate the collected vital signs data to the central station via the one or more transmitter/receivers; where the at least one patient monitor is operable by a user to identify the particular one patient from the plurality of patients, and to inform the central station of the identity of the particular patient; and where the central station is configured to access a database system in order to identify the plurality of patients and to display a list of the plurality of patients in order to associate the vital signs data received from the at least one patient monitor with the particular patient from said list wherein the central station is configured to transmit the list of the plurality of patients to the at least one patient monitor, and where the at least one patient monitor includes a display device adapted to display the list to the user. | 1. A medical monitoring system, comprising: a central station adapted to receive vital signs data concerning a plurality of patients; one or more transmitter/receivers associated with the central station; and at least one patient monitor configured to monitor a particular one of the plurality of patients by collecting vital signs data from the particular patient, and configured to establish communications with the central station and communicate the collected vital signs data to the central station via the one or more transmitter/receivers; where the at least one patient monitor is operable by a user to identify the particular one patient from the plurality of patients, and to inform the central station of the identity of the particular patient; and where the central station is configured to access a database system in order to identify the plurality of patients and to display a list of the plurality of patients in order to associate the vital signs data received from the at least one patient monitor with the particular patient from said list wherein the central station is configured to transmit the list of the plurality of patients to the at least one patient monitor, and where the at least one patient monitor includes a display device adapted to display the list to the user. 2. The system of claim 1, where the central station is adapted to store the identity of the particular patient and, in the event that the communications with the central station are lost and then restored, to transmit a request to the at least one patient monitor for a confirmation that the at least one patient monitor is still configured to monitor the particular patient. 3. The system of claim 1, where the system, the one or more transmitter/receivers, and the at least one patient monitor is wireless. 4. The system of claim 3, where the central station is configured to transmit a list of the plurality of patients to the at least one patient monitor, and where the at least one patient monitor includes a display device adapted to display the list to the user. 5. The system of claim 4, where the central station is adapted to store the identity of the particular patient and, in the event that the communications with the central station are lost and then restored, to transmit a request to the at least one patient monitor for a confirmation that the at least one patient monitor is still configured to monitor the particular patient. 6. A method for monitoring a patient, comprising the steps of: establishing communications between a patient monitor and a central station via a transmitter/receiver associated with the central station, where the central station is adapted to receive vital signs data from an identified plurality of patients, said plurality of patients being identified by said central station accessing a database system containing the identities of said patients; configuring the patient monitor to collect vital signs data from a particular one of the plurality o... | Welch Allyn Protocol, Inc. | A61B005/00
US07207951 | B1 | 2007-04-24 | Apparatus for obtaining biological samples | 2000-08-11 | A device for collection of a fluid sample, particularly a biological sample from an internal cavity of a mammal, comprises a barrel having an opening at one end thereof, a plunger operable axially within the barrel, the barrel and the plunger defining a fluid chamber having a volume which varies on axial movement of the plunger within the barrel, and a hollow, elongate tube extending from the fluid chamber through the opening in the barrel, the tube being in operative engagement with the plunger for axial movement to extend and retract the tube within respect to the barrel on axial movement of the plunger, and the tube being in fluid communication with the fluid chamber to provide a fluid flow path to and from the fluid chamber through the hollow tube. A method for collection of a fluid sample from an internal cavity of a mammal is also disclosed. | 1. A device for collection of a fluid sample, comprising a barrel having an opening at one end thereof, a plunger operable axially within the barrel, said barrel and said plunger defining a fluid chamber having a volume which varies on axial movement of the plunger within the barrel, and a flexible, hollow, elongate catheter extending from the fluid chamber through said opening in the barrel, said catheter being sufficiently flexible so as to be adapted to bend and follow the conformation of an internal cavity upon encountering walls of said cavity during insertion into said cavity without damaging tissue of said cavity, said catheter further being in operative engagement with said plunger for axial movement to extend and retract the catheter within respect to the barrel on axial movement of the plunger, and said catheter being in fluid communication with the fluid chamber so as to provide fluid flow to and from the fluid chamber through the catheter during extending and retracting, respectively, of the catheter. | 1. A device for collection of a fluid sample, comprising a barrel having an opening at one end thereof, a plunger operable axially within the barrel, said barrel and said plunger defining a fluid chamber having a volume which varies on axial movement of the plunger within the barrel, and a flexible, hollow, elongate catheter extending from the fluid chamber through said opening in the barrel, said catheter being sufficiently flexible so as to be adapted to bend and follow the conformation of an internal cavity upon encountering walls of said cavity during insertion into said cavity without damaging tissue of said cavity, said catheter further being in operative engagement with said plunger for axial movement to extend and retract the catheter within respect to the barrel on axial movement of the plunger, and said catheter being in fluid communication with the fluid chamber so as to provide fluid flow to and from the fluid chamber through the catheter during extending and retracting, respectively, of the catheter. 2. A device according to claim 1, wherein said catheter extends into a chamber in the plunger which is in fluid communication with said fluid chamber. 3. A device according to claim 1, wherein one end of said catheter is attached to said plunger and said catheter is provided with perforations in the wall thereof at or near the end thereof attached to the plunger for fluid communication with said fluid chamber. 4. A device according to claim 1, wherein one end of said catheter is attached to said plunger and the end of said catheter remote from the plunger is sealed, and the catheter is provided with perforations in the wall thereof at or near the sealed end for passage of fluid in and out of the hollow catheter. 5. A device according to claim 1, further comprising a filter located in the fluid flow path to and from the fluid chamber through the hollow catheter. 6. A device according to claim 5, wherein the filter is adapted to substantially remove cells and cellular debris from a fluid in said fluid flow path. 7. A device according to claim 5, wherein the filter is located in the hollow catheter. 8. A device according to claim 5, wherein the filter is located in the plunger. 9. A device according to claim 5, wherein the filter is located in the barrel. 10. A device according to claim 1, further comprising a return device located between the barrel and the plunger of the device. 11. A device according to claim 1, further comprising means for rotating the plunger on axial movement of the plunger within the barrel of the device. 12. A device according to claim 11, wherein said means for rotating is adapted to rotate the plunger from 90° to 360° on full axial movement of the plunger within the barrel. 13. A device according to claim 1, further comprising means for collecting a sample of cells or cellular debris, said means being located on said barrel at or adjacent to said opening at one end thereof. 14. A device according to c... | Diagnotech Pty. Ltd. | A61B005/00 B65D081/00
US06842638 | B1 | 2005-01-11 | Angiography method and apparatus | 2001-11-13 | A two-dimensional slice formed of pixels (376) is extracted from the angiographic image (76) after enhancing the vessel edges by second order spatial differentiation (78). Imaged vascular structures in the slice are located (388) and flood-filled (384). The edges of the filled regions are iteratively eroded to identify vessel centers (402). The extracting, locating, flood-filling, and eroding is repeated (408) for a plurality of slices to generate a plurality of vessel centers (84) that are representative of the vascular system. A vessel center (88) is selected, and a corresponding vessel direction (92) and orthogonal plane (94) are found. The vessel boundaries (710) in the orthogonal plane (94) are identified by iteratively propagating (704) a closed geometric contour arranged about the vessel center (88). The selecting, finding, and estimating are repeated for the plurality of vessel centers (84). The estimated vessel boundaries (710) are interpolated (770) to form a vascular tree (780). | 1. An apparatus for producing an angiographic image representation of a subject, the apparatus comprising: an imaging scanner that acquires imaging data from at least a portion of a subject, the imaging data including vascular contrast; a reconstruction processor that reconstructs a three-dimensional image representation from the imaging data, the image representation formed of image elements and exhibiting vascular contrast; and a processor that converts the image representation into an edge-enhanced image representation having enhanced vascular edges and divides the edge-enhanced image representation into a plurality of two-dimensional slices formed of pixels spanning the angiographic image representation, and for each slice: flood-fills the vascular edges to form filled regions defined by pixels having a first value, identifies vessel centers through iterative removal of pixels having the first value from around the edges of the filled regions, and the vessel centers identified being representative of a three-dimensional vascular structure, the processor further segmenting, tracking, extracting, enhancing, or identifying information about the three-dimensional vascular structure using the identified vessel centers as operative inputs. | 1. An apparatus for producing an angiographic image representation of a subject, the apparatus comprising: an imaging scanner that acquires imaging data from at least a portion of a subject, the imaging data including vascular contrast; a reconstruction processor that reconstructs a three-dimensional image representation from the imaging data, the image representation formed of image elements and exhibiting vascular contrast; and a processor that converts the image representation into an edge-enhanced image representation having enhanced vascular edges and divides the edge-enhanced image representation into a plurality of two-dimensional slices formed of pixels spanning the angiographic image representation, and for each slice: flood-fills the vascular edges to form filled regions defined by pixels having a first value, identifies vessel centers through iterative removal of pixels having the first value from around the edges of the filled regions, and the vessel centers identified being representative of a three-dimensional vascular structure, the processor further segmenting, tracking, extracting, enhancing, or identifying information about the three-dimensional vascular structure using the identified vessel centers as operative inputs. 2. The apparatus as set forth in claim 1, wherein the converting of the image representation into an edge-enhanced imaged representation includes: conditional upon the vascular contrast including black blood vascular contrast, inverting the intensities of the image elements to generate an intensity-inverted image. 3. The apparatus as set forth in claim 1, further including: a magnetic resonance contrast agent administered to the subject to improve vascular contrast. 4. The apparatus as set forth in claim 1, wherein the imaging scanner includes at least one of a magnetic resonance imaging scanner and a computed tomography scanner. 5. The apparatus as set forth in claim 1, wherein the identifying of vessel centers through iterative removal of pixels includes for each iteration: a first erosion pass operating in a first direction across the slice using a moving window having a first shape; and a second erosion pass operating in a second direction across the slice using a moving window having a second shape. 6. An apparatus for producing an angiographic image representation of a subject, the apparatus comprising: an imaging scanner that acquires imaging data from at least a portion of a subject, the imaging data including vascular contrast; a reconstruction processor that reconstructs an image representation from the imaging data, the image representation formed of image elements and exhibiting vascular contrast; a processor that converts the image representation into an edge-enhanced image representation having enhanced vascular edges and divides the edge-enhanced image representation into at least one two-dimensional slice formed of pixels, and for each slice: flood-fills the vascular edges to form filled... | Koninklijke Philips Electronics N.V. | A61B005/05
US06931275 | B2 | 2005-08-16 | System for reduction of undesirable brain wave patterns using selective photic stimulation | 2002-06-07 | A method of reducing undesirable brain wave pattern using selective photic stimulation. Electrodes, attached to a subject's scalp, transmit electroencephalographic (EEG) signals from the subject. These signals are in response to the photic stimulation being displayed to the subject. The resultant EEG signals are then filtered, while control logic monitors a predetermined EEG frequency band and feeds back the filtered EEG signal to the subject via the photic stimulation. The photic stimulator is then turned on and off rapidly, in response to the amount of energy in a predetermined band of the subject's EEG, with the intent of reducing an undesirable brain rhythm. | 1. A method of reducing undesirable brain wave patterns using selective photic stimulation, comprising the steps of: securing electrodes to a subject's scalp; displaying stimuli to the subject; receiving an electroencephalographic (EEG) signal from the subject while the stimuli is being displayed; transmitting the resultant EEG signal to a filter; filtering the EEG signal; monitoring a predetermined EEG frequency band; and feeding back the filtered EEG signal to the subject via photic stimulation, by controlling the presence or absence of the photic stimulation; wherein a fixed photic stimulation frequency is employed, regardless of the frequency content of a trainee's EEG rhythms. | 1. A method of reducing undesirable brain wave patterns using selective photic stimulation, comprising the steps of: securing electrodes to a subject's scalp; displaying stimuli to the subject; receiving an electroencephalographic (EEG) signal from the subject while the stimuli is being displayed; transmitting the resultant EEG signal to a filter; filtering the EEG signal; monitoring a predetermined EEG frequency band; and feeding back the filtered EEG signal to the subject via photic stimulation, by controlling the presence or absence of the photic stimulation; wherein a fixed photic stimulation frequency is employed, regardless of the frequency content of a trainee's EEG rhythms. 2. A method of reducing undesirable brain wave patterns using selective photic stimulation, comprising the steps of: securing electrodes to a subject's scalp; displaying stimuli to the subject; receiving an electroencephalagraphic (EEG) signal from the subject while the stimuli is being displayed; transmitting the resultant EEG signal to a filter; filtering the EEG signal; monitoring a predetermined EEG frequency band; and feeding back the filtered EEG signal to the subject via photic stimulation, by controlling the presence or absence of the photic stimulation; wherein the photic stimulation is presented by a photic stimulator, which is a display lamp fitted with LED's positioned in front of the subject; wherein a photic stimulator is turned on and off rapidly, in response to the amount of energy in a predetermined frequency band of a subject's EEG; wherein the photic stimulator is activated at specific times, with the intent of reducing the amount of an undesirable brain rhythm in a subject's EEG; wherein the process is used for brief times, such as 1 to 10 minutes, preferably as a precursor to conventional EEG neurofeedback training, as an assist to the neurofeedback training to cause the brain rhythm to reduce the amount of an undesirable rhythm; wherein the process employs classical conditioning instead of operant conditioning, so that learning takes place without instructions to the subject, and independent of the subject's intention to undertake training; wherein the process employs non-volitional feedback, so that the effect on the subject's EEG does not depend on instructions to the subject, or the subject's intent to undertake training. 3. A method of reducing undesirable brain wave patterns using selective magnetic, tactile or other stimulation, comprising the steps of: securing electrodes to a subject's scalp; displaying stimuli to the subject; receiving an electroencephalographic (EEG) signal from the subject while the stimuli is being displayed; transmitting the resultant EEG signal to a filter; filtering the EEG signal; monitoring a predetermined EEG frequency band; and feeding back the filtered EEG signal to the subject via stimulation, by controlling the presence or absence of stimulation; wherein the stimulation is turned on and off rapidly, in re... | - | A61B005/04
US06856831 | B2 | 2005-02-15 | Method for the early diagnosis of subacute, potentially catastrophic illness | 2001-01-29 | In one aspect of the invention, there is provided a method for early detection of subacute, potentially catastrophic illness in an infant. The method comprises: (a) monitoring heart rate variability in the infant; and (b) identifying at least one characteristic abnormality in the heart rate variability that is associated with the illness. This method can be used to diagnose illnesses such as, but not limited to, sepsis, necrotizing enterocolitis, pneumonia and meningitis. In another aspect of the present invention, there is provided a method for early detection of subacute, potentially catastrophic illness in an infant, which comprises: (a) monitoring the patient's RR intervals; (b) generating a normalized data set of the RR intervals; (c) calculating one or more of (i) moments of the data set selected from the third and higher moments and (ii) percentile values of the data set; and (d) identifying an abnormal heart rate variability associated with the illness based on one or more of the moments and the percentile values. | 1. A method for early detection of subacute, potentially catastrophic illness in an infant comprising: (a) monitoring frequency histograms of RR intervals in the infant; (b) identifying at least one characteristic abnormal pattern or distribution; and (c) correlating the at least one abnormal pattern or distribution with said illness. | 1. A method for early detection of subacute, potentially catastrophic illness in an infant comprising: (a) monitoring frequency histograms of RR intervals in the infant; (b) identifying at least one characteristic abnormal pattern or distribution; and (c) correlating the at least one abnormal pattern or distribution with said illness. 2. The method of claim 1, wherein the illness is infectious. 3. The method of claim 2, wherein antibiotic therapy is initiated and a diagnostic work-up for the illness, comprising obtaining a blood culture from the patient, is provided when the at least one characteristic abnormal pattern or distribution is identified. 4. The method of claim 2, wherein the illness is necrotizing enterocolitis. 5. The method of claim 4, wherein a diagnostic work-up for the illness, comprising an X-ray of the infant or a pathological specimen from the infant, is provided when the at least one characteristic abnormal pattern or distribution is identified. 6. The method of claim 2 wherein the illness is selected from the group consisting of pneumonia, sepsis, and meningitis. 7. The method of claim 1, wherein a diagnostic work-up is provided when the at least one characteristic abnormal pattern or distribution is identified. 8. The method of claim 1, wherein the infant is a neonate. 9. The method of claim 1, wherein the at least one characteristic abnormal pattern or distribution is identified from a data set of RR intervals. 10. The method of claim 9, wherein the data set is normalized. 11. The method of claim 10, wherein the data set contains on the order of about 103 to 104 sequential RR intervals. 12. The method of claim 10, wherein the at least one characteristic abnormal pattern or distribution is identified based on at least one of the third and higher moments of the data set. 13. The method of claim 12, wherein the at least one moment of the data set includes the skewness of the data set. 14. The method of claim 13, wherein the illness is sepsis or necrotizing enterocolitis. 15. The method of claim 12, wherein the at least one moment of the data set includes the kurtosis of the data set. 16. The method of claim 15, wherein the illness is sepsis or necrotizing enterocolitis. 17. The method of claim 10, wherein the at least one characteristic abnormal pattern or distribution is identified based on at least one percentile value of the data set. 18. The method of claim 17, wherein the at least one percentile value is the 10th percentile value. 19. The method of claim 18, wherein the illness is sepsis or necrotizing enterocolitis. 20. The method of claim 10, wherein the at least one characteristic abnormal pattern or distribution is identified based on the variance, standard deviation or coefficient of variation of the data set. 21. The method of claim 20, wherein the illness is sepsis or necrotizing enterocolitis. 22. The method of claim 14, further comprising a diagnostic work-up. 23. The method of claim 15, f... | University of Virginia Patent Foundation | A61B005/0452
US06850794 | B2 | 2005-02-01 | Endoscopic targeting method and system | 2001-09-20 | The invention provides a system for enhancing the ability of a surgeon to access a target site within a patient. The system includes a data file containing volumetric scan data of a region of the patient that includes the target site, a display device, and a movable imagining tool for producing on the display device, an image of visible patient structure seen by the tool. A computer in the system operates to (i) determine the position and/or orientation of the tool in the frame of reference of the patient, (ii) identify the scan-data coordinates (either x,y or x,y,z coordinates) of the target site, and (iii) project on the video image on the display device, indicia that indicate the lateral position of the target site with respect to the patient structure imaged on the display device. | 1. A system for enhancing the ability of a surgeon to access a target site within a patient, comprising a data file containing volumetric scan data of a region of the patient that includes the target site, a display device, a movable imagining tool for producing on the display device, an image of visible structure seen by the tool, where the position and/or orientation of the tool is tracked with respect to the patient coordinate system, and a computer operatively connected to data file and display screen for (i) determining the position and/or orientation of the tool in a patient coordinate system, (ii) identifying the scan-data coordinates of the target site in the patient coordinate system, and (iii) projecting on the video image on the display device, indicia that indicate (a) the direction of the target site, if the target site is outside the patient structure imaged on the display device, and (b) the lateral position of the target site with respect to the patient structure imaged on the display device and the distance between the tool and the target site, if the target site is within the patient structure imaged on the display device. | 1. A system for enhancing the ability of a surgeon to access a target site within a patient, comprising a data file containing volumetric scan data of a region of the patient that includes the target site, a display device, a movable imagining tool for producing on the display device, an image of visible structure seen by the tool, where the position and/or orientation of the tool is tracked with respect to the patient coordinate system, and a computer operatively connected to data file and display screen for (i) determining the position and/or orientation of the tool in a patient coordinate system, (ii) identifying the scan-data coordinates of the target site in the patient coordinate system, and (iii) projecting on the video image on the display device, indicia that indicate (a) the direction of the target site, if the target site is outside the patient structure imaged on the display device, and (b) the lateral position of the target site with respect to the patient structure imaged on the display device and the distance between the tool and the target site, if the target site is within the patient structure imaged on the display device. 2. The system of claim 1, wherein the system includes a tracking device for tracking the position and direction of the instrument with respect to patient position. 3. The system of claim 2, wherein said indicia define an area whose center corresponds to the lateral position of the target site, and whose size is related to the distance between the instrument and the target site. 4. The system of claim 3, wherein the size of the area defined by the indicia is calculated by an error function representing the degree uncertainty in patient target site in the imaged structure. 5. The system of claim 1, wherein said computer is operatively connected to a calibration device for receiving information therefrom with respect to the imaging tool's field of view, position, direction orientation. 6. The system of claim 1, wherein (i) said imaging tool is an endoscope having an interior channel through which a surgical tool can be manipulated, (ii) the system further includes structure for tracking the position of said surgical tool relative to said imaging device, and (iii) the computer operates to project on the video image on the display device, indicia relating the position of the target site in relation to the tip of said tool. 7. A method for enhancing the ability of a surgeon to access a target site within a patient, comprising storing in a data file, volumetric scan data of a region of the patient, including a target site within the patient region, moving an imagining tool for producing on a display device, an image of visible patient structure seen by the tool, tracking the position of the tool relative to the position of the patient, and by means of a computer a computer operatively connected to data file and display screen, (i) determining the position and/or orientation of the tool in the frame of referenc... | The Trustees of the Leland Stanford Junior University | A61B005/05
US06915155 | B2 | 2005-07-05 | Medical testing system with an illuminating component and automatic shut-off | 2001-07-26 | A medical testing system has an instrument for monitoring a characteristic of a patient and an illuminating component for illuminating the instrument. The instrument includes a component for selectively activating and deactivating the illuminating component and a deactivating component for automatically deactivating the illuminating component, after a predetermined period of time has elapsed. | 1. A medical testing system for monitoring the electrical activity of a patient's heart by successively performing both electrocardiogram and echocardiogram monitoring procedures in the same location, the system comprising: a movable cart housing a central control system for controlling both the electrocardiogram and echocardiogram procedures; the central control system comprising a display for displaying monitoring output from the electrocardiogram and/or echocardiogram monitoring procedures, a printer for printing monitoring output from the electrocardiogram and/or echocardiogram monitoring procedures, a work surface upon which the printed monitoring output is disposed, a keypad comprising a plurality of keys for inputting control commands to the control system to conduct the echocardiogram procedures, and a keyboard comprising a plurality of keys for inputting commands to the control system to conduct the electrocardiogram and/or echocardiogram procedures; and an illuminating component for selectively illuminating the work surface, keypad and keyboard and thereby facilitating successive performance of the electrocardiogram and echocardiogram monitoring procedures by allowing a caregiver to view and manipulate the keys on the keypad and keyboard and view the printed monitoring output regardless of surrounding lighting conditions of the location, wherein the illuminating component is automatically de-activated if a key on the keypad is not actuated for a predetermined time period. | 1. A medical testing system for monitoring the electrical activity of a patient's heart by successively performing both electrocardiogram and echocardiogram monitoring procedures in the same location, the system comprising: a movable cart housing a central control system for controlling both the electrocardiogram and echocardiogram procedures; the central control system comprising a display for displaying monitoring output from the electrocardiogram and/or echocardiogram monitoring procedures, a printer for printing monitoring output from the electrocardiogram and/or echocardiogram monitoring procedures, a work surface upon which the printed monitoring output is disposed, a keypad comprising a plurality of keys for inputting control commands to the control system to conduct the echocardiogram procedures, and a keyboard comprising a plurality of keys for inputting commands to the control system to conduct the electrocardiogram and/or echocardiogram procedures; and an illuminating component for selectively illuminating the work surface, keypad and keyboard and thereby facilitating successive performance of the electrocardiogram and echocardiogram monitoring procedures by allowing a caregiver to view and manipulate the keys on the keypad and keyboard and view the printed monitoring output regardless of surrounding lighting conditions of the location, wherein the illuminating component is automatically de-activated if a key on the keypad is not actuated for a predetermined time period. 2. The medical testing system of claim 1, wherein the movable cart comprises a plate for supporting the display above the work surface and the illuminating component is attached to the plate. 3. The medical testing system of claim 2, wherein the illuminating component comprises a circuit board having a plurality of light emitting diodes. 4. The medical testing system of claim 1, wherein the illuminating component comprises a circuit board having a plurality of light emitting diodes. 5. The medical testing system of claim 1, wherein the illuminative component can be activated or de-activated by actuating a key on the keypad. 6. The medical testing system of claim 1, wherein the illuminating component comprises a plurality of light emitting diodes. | GE Medical Systems Information Technologies, Inc. | A61B005/0402
US06901277 | B2 | 2005-05-31 | Methods for generating a lung report | 2001-07-17 | Software, methods and user interfaces for viewing and generating a lung report. In exemplary methods, a first baseline image is analyzed to localize lung nodules. The lung nodule information can be stored and used to relocate the previously localized lung nodules in a second, follow-up imaging scan. The follow-up scan can be analyzed—either manually or automatically—to generate statistics for the lung nodules localized in the follow-up scan. Information derived from the baseline scan can be used to ease the localization of the lung nodules in the follow-up scan. In some exemplary methods, both the baseline scan and follow-up scan are displayed and the nodules are selected for comparison. Typically, the graphical user interface can display a third, composite image that displays the change of size of the nodules. | 1. A method of analyzing lung nodules in a body slice image and generating a patient report, the method comprising: providing patient information; displaying a first image along a body slice taken at t1; displaying a second image along the body slice taken at t2; localizing lung nodules in the first image and second image; registering the first image with the second image so that lung nodules in the first image are substantially registered with corresponding lung nodules in the second image; comparing the lung nodules in the first image with the lung nodules in the second image; and generating the patient report using a decision tree or a set of logic steps that combines patient information, and the comparison of selected parameters of the lung nodules and the number of lung nodules in the first image and the second image. | 1. A method of analyzing lung nodules in a body slice image and generating a patient report, the method comprising: providing patient information; displaying a first image along a body slice taken at t1; displaying a second image along the body slice taken at t2; localizing lung nodules in the first image and second image; registering the first image with the second image so that lung nodules in the first image are substantially registered with corresponding lung nodules in the second image; comparing the lung nodules in the first image with the lung nodules in the second image; and generating the patient report using a decision tree or a set of logic steps that combines patient information, and the comparison of selected parameters of the lung nodules and the number of lung nodules in the first image and the second image. 2. The method of claim 1 wherein registering comprises automatically registering the first image and second image, wherein the comparing identifies changes in the parameters of the lung nodules and the number of lung nodules. 3. The method of claim 1 wherein registering comprises manually panning at least one of the first and second images to align the first image with the second image, wherein the comparing identifies changes in the parameters of the nodules and the number of lung nodules. 4. The method of claim 1 wherein registering comprises automatically or semi-automatically panning at least one of the first and second images to register the first image with the second image, wherein the comparing identifies changes in the parameters of the nodules and the number of lung nodules. 5. The method of claim 4 wherein registering comprises automatically correcting for differences at acquisition in spatial parameters prior to panning at least one of the first and second images to align the first image with the second image. 6. The method of claim 1 wherein comparing comprises computing a third image which is a subtraction of the first image and second image. 7. The method of claim 1 wherein comparing comprises computing a third image which is an absolute difference of the first image and second image. 8. The method of claim 1 wherein comparing comprises computing a third image which is a sum of the first image and second image. 9. The method of claim 1 wherein registering comprises matching the lung nodules in the first image and second image, wherein matching comprises demanding that at least one voxel be overlapped. 10. The method of claim 1 wherein registering comprises matching the nodule in the first image and second image, wherein matching comprises demanding that the edges of the lung nodule be no more than a fixed number of pixels apart. 11. The method of claim 10 wherein nodule in the first image and second image are matched by demanding that the edges of the lung nodule be no more than one pixel apart. 12. The method of claim 10 wherein nodules are matched by demanding that the edges of the lung nodule be no... | AccuImage Diagnostics Corp. | A61B005/05
US06969359 | B2 | 2005-11-29 | Body fluid sampling device | 2002-09-19 | Body fluid is sampled from a skin incision by placing a ring against the skin and repeatedly applying an elastic pressing force to the ring, whereby a stimulator surface of the ring depresses a ring of skin and body tissue in surrounding relationship to the incision to force body fluid from the incision. The stimulator surface is inclined at an angle of 10 to 65 degrees; a width of the stimulator surface is from 5 mm to 20 mm, and an inner diameter of the stimulator surface is no less than 6.0 mm. | 1. A sampling device for sampling a bodily fluid through an incision in a skin surface, comprising: a housing;a stimulator member attached to the housing, the stimulator member being adapted to contact the skin surface around the incision and express a drop of the bodily fluid from the incision, the stimulator member defining an interior opening having a dimension of at least 6.0 mm to minimize smearing of the drop by contact between the drop and the stimulator member;a second skin contacting member configured to contact the skin surface;wherein the housing extends along a longitudinal axis;wherein the second skin contacting member is longitudinally movable relative to the housing; andwherein the stimulator member has a first position relative the second skin contacting member in which the second skin contacting member extends longitudinally outward from the stimulator member and a second position in which the stimulator member extends longitudinally outward from the second skin contacting member. | 1. A sampling device for sampling a bodily fluid through an incision in a skin surface, comprising: a housing;a stimulator member attached to the housing, the stimulator member being adapted to contact the skin surface around the incision and express a drop of the bodily fluid from the incision, the stimulator member defining an interior opening having a dimension of at least 6.0 mm to minimize smearing of the drop by contact between the drop and the stimulator member;a second skin contacting member configured to contact the skin surface;wherein the housing extends along a longitudinal axis;wherein the second skin contacting member is longitudinally movable relative to the housing; andwherein the stimulator member has a first position relative the second skin contacting member in which the second skin contacting member extends longitudinally outward from the stimulator member and a second position in which the stimulator member extends longitudinally outward from the second skin contacting member. 2. The device of claim 1, wherein both the stimulator member and the second skin contacting member are ring shaped and co-axially oriented with respect to one another. 3. The device of claim 1, further comprising: a first force transmitting member positioned between the second skin contacting member and the housing; anda second force transmitting member positioned between the stimulator member and the second skin contacting member. 4. The device of claim 3, wherein: the first force transmitting member includes an outer spring; andthe second force transmitting member includes a sleeve spring. 5. The device of claim 3, wherein the first force transmitting member and the second force transmitting member are constructed and arranged to relieve pressure applied by the stimulator member against the skin surface and to allow fresh bodily fluid to flow toward the incision while the second skin contacting member remains in contact with the skin surface upon movement of the housing away from the skin surface. 6. The device of claim 3, wherein the first force transmitting member and the second force transmitting member are constructed and arranged to extend the stimulator member to the second position and engage the skin surface upon movement of the housing toward the skin surface. 7. The device of claim 3, wherein both the stimulator member and the second skin contacting member are ring shaped and co-axially oriented with respect to one another. 8. The device of claim 3, wherein the second skin contacting member is longitudinally movable relative to the housing. 9. The device of claim 8, wherein the stimulator member is slidable relative to the housing. 10. A sampling device for sampling a bodily fluid through an incision in a skin surface, comprising: a housing;a stimulator member attached to the housing, the stimulator member being adapted to contact the skin surface around the incision and express a drop of the bodily fluid from the incision, the stim... | Roche Diagnostics Operations, Inc. | A61B005/00
US06926678 | B1 | 2005-08-09 | Sample collection and delivery device | 1996-09-27 | A specimen collecting device is disclosed. The device has an elongated foam member that is circumscribed by a hollow tubular member along a portion of its longitudinal axis. The hollow tubular member has a cross sectional area less than the uncompressed cross sectional area of the foam member so that the foam member is compressed along the circumscribed portion. | 1. A specimen collecting device useful for collecting a fluid specimen for diagnostic testing consisting essentially of an elongated absorbent foam member having a longitudinal axis and an uncompressed cross sectional area, the elongated foam member being circumscribed by an elongated hollow, one-piece tubular member having a flexible side wall along a portion of the longitudinal axis, the hollow tubular member having a cross sectional area less than the uncompressed cross sectional area of the foam member so that the foam member is compressed by the hollow tubular member along the circumscribed portion, wherein the foam member consists of a material suitable for collecting a specimen of a patient fluid and the foam member remains circumscribed by the tubular member during collection of the specimen, andwherein pressure can be exerted on the side wall of the tubular member to compress the foam and deliver a sample of the fluid specimen. | 1. A specimen collecting device useful for collecting a fluid specimen for diagnostic testing consisting essentially of an elongated absorbent foam member having a longitudinal axis and an uncompressed cross sectional area, the elongated foam member being circumscribed by an elongated hollow, one-piece tubular member having a flexible side wall along a portion of the longitudinal axis, the hollow tubular member having a cross sectional area less than the uncompressed cross sectional area of the foam member so that the foam member is compressed by the hollow tubular member along the circumscribed portion, wherein the foam member consists of a material suitable for collecting a specimen of a patient fluid and the foam member remains circumscribed by the tubular member during collection of the specimen, andwherein pressure can be exerted on the side wall of the tubular member to compress the foam and deliver a sample of the fluid specimen. 2. The specimen collecting device of claim 1, wherein the hollow tubular member is made of paper or plastic. 3. The specimen collecting device of claim 1, wherein the foam member comprises a material selected from the group consisting of polyurethane foam, polyethylene foam, polyvinylchloride foam, ethylvinylacetate foam, polyethylene/ethylvinylacetate foam, polyester foam and polyether foam. 4. The specimen collecting device of claim 1, wherein the foam member comprises a polyurethane foam. 5. The specimen collecting device of claim 1, wherein the foam member has a mean diameter in an uncompressed state, and wherein the foam member protrudes from an end of the tubular member a distance of about 25% to about 125% of said mean diameter of the uncompressed cross sectional area of the foam member. 6. The specimen collecting device of claim 1, wherein the hollow tubular member has a length of from about 1.0 inch to about 5.0 inches. 7. The specimen collecting device of claim 1, wherein the hollow tubular member has a mean diameter of from about 0.12 to about 0.5 inch. 8. The specimen collecting device of claim 1, wherein the hollow tubular member has a length of from about 1.5 inches to about 3.5 inches. 9. The specimen collecting device of claim 1, wherein the hollow tubular member has a mean diameter of from about 0.18 to about 0.38 inch. 10. A specimen collecting device useful for collecting a fluid specimen for diagnostic testing consisting essentially of an elongated absorbent foam member having a longitudinal axis and an uncompressed cross sectional area, the elongated foam member being circumscribed by an elongated hollow tubular member having a flexible side wall along a portion of the longitudinal axis, the hollow tubular member having a cross sectional area less than the uncompressed cross sectional area of the foam member so that the foam member is compressed by the hollow tubular member along the circumscribed portion, wherein the hollow tubular member has a mean diameter of from about 0.12 to abo... | Avitar, Inc. | A61B005/00 A61B010/00
US06912412 | B2 | 2005-06-28 | System and methods of fluorescence, reflectance and light scattering spectroscopy for measuring tissue characteristics | 2002-01-18 | The present invention utilizes a plurality of spectroscopic systems and methods to measure characteristics of tissue useful in the diagnosis of disease. In a preferred embodiment, a combination of fluorescence, reflectance and light scattered spectra can be measured and processed to provide biochemical, architectural and morphological state of tissue. The methods and systems can be used particularly in the early detection of carcinoma within tissue in vivo and in vitro. | 1. A method of analyzing spectral data to measure a structure in a layer of tissue comprising: providing a light collection system that collects fluorescent and reflected light from the tissue at a plurality of wavelengths and detects the collected light; forming a fluorescence representation and a scattered light representation as a function of wavelength from the detected light; and determining a characteristic of the tissue layer with the fluorescence representation and the scattered light representation. | 1. A method of analyzing spectral data to measure a structure in a layer of tissue comprising: providing a light collection system that collects fluorescent and reflected light from the tissue at a plurality of wavelengths and detects the collected light; forming a fluorescence representation and a scattered light representation as a function of wavelength from the detected light; and determining a characteristic of the tissue layer with the fluorescence representation and the scattered light representation. 2. The method of claim 1 further comprising using the fluorescence representation to determine one or more components of the tissue. 3. The method of claim 1 further comprising using the scattered light representation to determine a size of a structure within the tissue layer. 4. The method of claim 1 further comprising measuring a periodic component of the detected light to provide the scattered light representation. 5. The method of claim 1 further comprising analyzing a diffuse reflectance spectrum and a light scattering spectrum. 6. The method of claim 1 further comprising calibrating with a reflectance standard and a fluorescence standard. 7. The method of claim 1 further comprising generating a lookup table with different sizes of scatterers. 8. The method of claim 1 further comprising correcting the fluorescence spectrum to obtain an intrinsic fluorescence spectrum. 9. The method of claim 5 further comprising transforming the light scattering spectrum and removing spectral components to determine size distribution of a tissue component. 10. The method of claim 1 further comprising performing discriminate analysis to diagnose the tissue. 11. The method of claim 1 further comprising separating a plurality of scattering coefficient values into a plurality of different regions. 12. An apparatus for optically measuring tissue comprising: a broadband light source that generates light in the range of 330-700 nm and that illuminates a region of interest in tissue to be measured with incident radiation; an optical system that collects scattered and fluorescent light from the tissue at a plurality of wavelengths; a detector system that senses the collected light and provides fluorescence data and scattered light data as a function of wavelength; and a data processor that determines a characteristic of the region of interest with the fluorescence data and the scattered light data. 13. The apparatus of claim 12 further comprising a second laser light source that generates light in the range of 330-700 nm. 14. The apparatus of claim 12 further comprising a fiber optic probe that couples the source to the tissue. 15. The apparatus of claim 14 wherein the probe is insertable in an endoscope. 16. The apparatus of claim 12 further comprising a fast excitation-emission matrix instrument. 17. A method for analyzing tissue spectra comprising: acquiring fluorescence spectra; acquiring a reflectance spectra; processing the fluorescence spe... | Massachusetts Institute of Technology | A61B005/00
US06899676 | B1 | 2005-05-31 | Method for determining the stress capacity of a person | 2000-03-08 | The invention relates to a method for determining the stress capacity of a person taking into consideration the individual anaerobic threshold by measuring lactate concentrations in relation to physical effort. To make it possible to deduce from the changes in lactate concentration other performance data characteristic of the person tested, the invention provides for the following process steps: measurement of time-dependent lactate concentration changes above the individual anaerobic threshold; adjustment of the measurement cure, in which lactate concentration is recorded in relation to time, to the measurement values obtained in this way; determination of a first rise in the measurement curve at a time tIAT which corresponds to the individual anaerobic threshold: determination of at least one other rise in the measurement curve at a time tx where tx>tIAT and subtraction of the second rise from the first rise to determine a difference ΔA. | 1. Process for determining the stress capacity of a person while taking into consideration the individual anaerobic threshold by measuring lactate concentrations in dependence upon work performed, comprising the steps of measuring the time-dependent lactate concentration changes above the individual anaerobic threshold, plotting a curve of the measured values of the lactate concentration with respect to time, determining a first slope of the curve at time point tIAT, at the individual anaerobic threshold, determining at least one more slope of the measured curve at a time point tx, with tx>tIAT, and subtracting the second slope from the first slope for determining a difference ΔA. | 1. Process for determining the stress capacity of a person while taking into consideration the individual anaerobic threshold by measuring lactate concentrations in dependence upon work performed, comprising the steps of measuring the time-dependent lactate concentration changes above the individual anaerobic threshold, plotting a curve of the measured values of the lactate concentration with respect to time, determining a first slope of the curve at time point tIAT, at the individual anaerobic threshold, determining at least one more slope of the measured curve at a time point tx, with tx>tIAT, and subtracting the second slope from the first slope for determining a difference ΔA. 2. Process according to claim 1, wherein differences ΔAx with x=1, 2 . . . are recorded with respect to the effort expended and the measured curves formed in this way are compared with standard curves, measured curves of different persons, or measured curves of the same person at different stress times. 3. Process according to claim 1, wherein the curves are measured during a staged effort change of the person. 4. Process according to claim 1, wherein the curves are measured during a continuous effort change of the person. | - | A61B005/00
US06904305 | B2 | 2005-06-07 | System for MRI-guided interventional mammary procedures | 2001-12-20 | The combination of contrast enhanced magnetic resonance imaging (MRI) and MR-guided subcutaneous core biopsy can be used as a robust approach for the diagnosis and treatment of breast cancer. MRI provides the means to accurately position and monitor interventional procedures such as biopsy, removal of tissue or other transcanular procedures. MRI may also be used in this invention to position and monitor the progress of breast conserving therapies (BCT), such as laser photo-ablation, cryoablation and localized hyperthermia. The general practice of this invention is to provide a remotely controlled apparatus for MR-guided interventional procedures in the breast. The apparatus allows the practice of a method that provides flexibility in conditioning the breast, i.e. orientation and degree of compression, and in setting the trajectory of the intervention. To that end, a robust conditioning/positioning device, fitted with the appropriate degrees of freedom, enhances the efficacy and efficiency of breast interventions by providing the flexibility in planning and executing an appropriate procedure strategy that better suits interventional procedures, either those in current use or yet to be developed. The novelty and potential commercial success of the device originates from its high maneuverability to set and perform the procedure strategy and its adaptability to accommodate an array of interventional probes. Remote control of this device can allow planning the operation and performing the relevant tasks in a short period, for example, within the contrast window provided by a single injection of a contrast agent, and this feature can be operator-independent. | 1. A system for guiding a probe in a magnetic resonance imaging apparatus comprising: a first compression surface; a second compression surface aligned substantially parallel with the first compression surface wherein the first compression surface and the second compression surface are each positioned along a first axis; a compression adjuster adapted to adjust a distance between the first compression surface and the second compression surface along the first axis; a stage coupled to the first compression surface and the second compression surface and adapted to rotate on a second axis substantially orthogonal to the first axis, the second axis having a user selectable position; and a probe guide having an intervention axis and adapted to guide the probe for insertion into a region between the first compression surface and the second compression surface along the intervention axis. | 1. A system for guiding a probe in a magnetic resonance imaging apparatus comprising: a first compression surface; a second compression surface aligned substantially parallel with the first compression surface wherein the first compression surface and the second compression surface are each positioned along a first axis; a compression adjuster adapted to adjust a distance between the first compression surface and the second compression surface along the first axis; a stage coupled to the first compression surface and the second compression surface and adapted to rotate on a second axis substantially orthogonal to the first axis, the second axis having a user selectable position; and a probe guide having an intervention axis and adapted to guide the probe for insertion into a region between the first compression surface and the second compression surface along the intervention axis. 2. The system of claim 1 further comprising a height adjuster coupled to the stage and adapted to adjust a height between the stage and a supporting surface. 3. The system of claim 2 further including a stage height remote control coupled to the height adjuster. 4. The system of claim 1 further comprising a lateral adjuster coupled to the stage and adapted to adjust a lateral distance between the stage and a supporting surface. 5. The system of claim 4 further including a stage lateral position remote control coupled to the lateral adjuster. 6. The system of claim 1 further comprising a probe guide angle adjuster coupled to the stage and the probe guide and adapted to adjust an angle between the probe guide and the stage. 7. The system of claim 6 further including a probe guide angle remote control coupled to the probe guide angle adjuster. 8. The system of claim 1 further comprising a probe guide height adjuster coupled to the stage and the probe guide and adapted to adjust a height of the probe guide relative to the stage. 9. The system of claim 8 further including a probe guide height remote control coupled to the probe guide height adjuster. 10. The system of claim 1 further comprising a probe guide depth stop adjuster coupled to the probe guide and adapted to adjust an insertion distance of a probe coupled to the probe guide. 11. The system of claim 10 further including a depth stop remote control coupled to the probe guide depth stop adjuster. 12. The system of claim 1 further comprising a patient supporting surface adapted to rotate on a supporting surface axis aligned orthogonal to the second axis and substantially parallel with a spine of a patient when a patient is positioned on the supporting surface. 13. The system of claim 12 further including a patient support remote control coupled to the patient supporting surface. 14. The system of claim 1 further including one or more linear actuators coupled to the compression adjuster. 15. The system of claim 14 further including one or more remote controls coupled to the one or more linear actuators.... | Regents of the University of Minnesota | A61B005/05
US07056294 | B2 | 2006-06-06 | Method and apparatus for accessing the left atrial appendage | 2002-03-15 | Disclosed is an apparatus for facilitating access to the left atrium, and specifically the left atrial appendage. The apparatus may comprise a sheath with first and second curved sections that facilitate location of the fossa ovalis and left atrial appendage. The apparatus may further comprise tissue piercing and dilating structures. Methods are also disclosed. | 1. A transseptal access system, comprising a sheath, a dilator and a needle; wherein the sheath comprises:a first curved section shaped to abut the interior wall of the right atrium substantially opposite the septum and directed toward the septum, and a second curved section shaped to direct a distal end of the sheath toward the left atrial appendage;the first curved section having an outer edge that definesa first proximal edge point, anda first distal edge pointwherein the intersection between a tangent drawn to the first curved section at the first proximal edge point and a tangent drawn to the first curved section at the first distal edge point, define a first angle and a first plane;the second curved section having an outer edge that definesa second proximal edge point, anda second distal edge pointwherein the intersection between a tangent drawn to the second curved section at the second proximal edge point and a tangent drawn to the second curved section at the second distal edge point, define a second angle and a second plane;wherein the first angle is within the range of about 75 degrees to about 105 degrees;wherein the second angle is within the range of about 75 degrees to about 105 degrees;wherein the angle between the first and second plane is within the range of about 75 degrees to about 105 degrees;wherein the needle is coated with PTFE. | 1. A transseptal access system, comprising a sheath, a dilator and a needle; wherein the sheath comprises:a first curved section shaped to abut the interior wall of the right atrium substantially opposite the septum and directed toward the septum, and a second curved section shaped to direct a distal end of the sheath toward the left atrial appendage;the first curved section having an outer edge that definesa first proximal edge point, anda first distal edge pointwherein the intersection between a tangent drawn to the first curved section at the first proximal edge point and a tangent drawn to the first curved section at the first distal edge point, define a first angle and a first plane;the second curved section having an outer edge that definesa second proximal edge point, anda second distal edge pointwherein the intersection between a tangent drawn to the second curved section at the second proximal edge point and a tangent drawn to the second curved section at the second distal edge point, define a second angle and a second plane;wherein the first angle is within the range of about 75 degrees to about 105 degrees;wherein the second angle is within the range of about 75 degrees to about 105 degrees;wherein the angle between the first and second plane is within the range of about 75 degrees to about 105 degrees;wherein the needle is coated with PTFE. 2. The transseptal access system of claim 1, wherein the first angle is within the range of about 80 degrees to about 100 degrees. 3. The transseptal access system of claim 1, wherein the second angle is within the range of about 80 degrees to about 100 degrees. 4. The transseptal access system of claim 1, wherein the dilator comprises an inflatable balloon. 5. The transseptal access system of claim 1, wherein the dilator comprises a spiral helical structure. 6. The transseptal access system of claim 1, wherein the dilator is coated with PTFE. 7. The transseptal access system of claim 1, wherein the needle is advanced by a biasing means. 8. A transseptal access system, comprising a sheath, a dilator and a needle; wherein the sheath comprises:a first curved section shaped to abut the interior wall of the right atrium substantially opposite the septum and directed toward the septum, and a second curved section shaped to direct a distal end of the sheath toward the left atrial appendage;the first curved section having an outer edge that definesa first proximal edge point, anda first distal edge pointwherein the intersection between a tangent drawn to the first curved section at the first proximal edge point and a tangent drawn to the first curved section at the first distal edge point, define a first angle and a first plane;the second curved section having an outer edge that definesa second proximal edge point, anda second distal edge pointwherein the intersection between a tangent drawn to the second curved section at the second proximal edge point and a tangent drawn to the second curved sect... | ev3 Sunnyvale, Inc | A61B005/00 A61M025/00
US07254438 | B2 | 2007-08-07 | Multimodality medical imaging system and method with intervening patient access area | 2001-10-19 | The invention comprises a system and method for creating medical images of a subject patient using a plurality of imaging devices, such as tomographic imaging scanners. The system comprises a plurality of imaging devices, each having a bore through which a patient is translated during scanning. The bores of each device are substantially aligned axially with respect to each other. An open area is formed between the imaging devices along the path of the patient, through which a caregiver can attain line-of-sight visual contact with or other access to the patient. During a scanning operation, the scanned portion of a patient is transported through the bore of the first device, past the opening between the devices and into the bore of the second device. | 1. A medical imaging apparatus, comprising: a CT imaging device for obtaining one or more tomographic images of a subject patient, the CT imaging device having a first housing which defines a bore through which the patient axially translates during formation of one or more images by the device;a nuclear camera imaging device for obtaining one or more tomographic images of the subject patient, nuclear camera imaging device having a second housing which defines a bore through which a patient axially translates during formation of the images by the device, which first and second housings are selectably securable in a fixed abutting position to one another during the formation of one or more images of the subject patient, in which fixed position the bore of each device is substantially aligned axially with the bore of the other; andwhen the first and second housings are secured in a fixed abutting position to one another, the bores of the CT and nuclear camera imaging devices are spaced apart by a distance sufficient to allow direct access by a caregiver to a portion of the subject patient which is positioned between the first and second bores, the spacing between the bores is free of obstructions in a region above the subject patient. | 1. A medical imaging apparatus, comprising: a CT imaging device for obtaining one or more tomographic images of a subject patient, the CT imaging device having a first housing which defines a bore through which the patient axially translates during formation of one or more images by the device;a nuclear camera imaging device for obtaining one or more tomographic images of the subject patient, nuclear camera imaging device having a second housing which defines a bore through which a patient axially translates during formation of the images by the device, which first and second housings are selectably securable in a fixed abutting position to one another during the formation of one or more images of the subject patient, in which fixed position the bore of each device is substantially aligned axially with the bore of the other; andwhen the first and second housings are secured in a fixed abutting position to one another, the bores of the CT and nuclear camera imaging devices are spaced apart by a distance sufficient to allow direct access by a caregiver to a portion of the subject patient which is positioned between the first and second bores, the spacing between the bores is free of obstructions in a region above the subject patient. 2. An imaging apparatus having an imaging axis, comprising: separatable first and second imaging devices for obtaining one or more images of a patient, wherein the patient is substantially aligned with the imaging axis;a first housing which houses the first imaging device and defines a first bore; anda second housing which houses the second imaging device and defines a second bore;projecting engageable securement structures which extend from at least one of the first housing and the second housing to fixedly attach the first and second imaging devices in positions abutting each other and fixed relative to the imaging axis; andan opening formed between the first and second bores when the first and second housings are abutting through which opening a caregiver can have line-of-sight visual contact with the patient that is aligned with the imaging axis and extends between the first and second imaging devices. 3. A medical imaging apparatus, comprising: a first tomographic medical imaging device having an opening for receipt of a subject patient;a second tomographic medical imaging device having an opening for receipt of the subject patient;a patient support structure extending through the openings of the first and second imaging devices during the formation of one or more images by at least one of the imaging devices;an imaging device support structure securing the first and second imaging devices in a fixed abutting relationship with the openings of the first and second imaging devices in alignment with an imaging axis during the formation of one or more tomographic images, by at least one of the imaging devices, of the subject patient;which imaging device support structure forms a patient access area between the first... | Koninklijke Philips Electronics N.V. | A61B005/05
US07171255 | B2 | 2007-01-30 | Virtual reality 3D visualization for surgical procedures | 2001-07-02 | A method and apparatus for three-dimensional imaging and treatment of a patient's body. The method and apparatus utilize a system for developing a therapy plan for treatment of an organ of the patient, a device for generating ultrasound image data from a treatment region and a device for providing a translucent volume image of a portion of a patient's body and a separate translucent image of the patient organ and a three dimensional viewing device to superimpose a translucent article image to enable viewing of the article image simultaneously with the patient organ and a portion of the patient's body. | 1. A computer-readable medium for facilitating performance of medical therapy treatment on a patient, the computer-readable medium comprising: one or more instructions executable by a processor for processing real-time image data to generate in substantially real time an image of a treatment region of a patient's body along with an image of a therapeutic device inserted inside the patient's body relative to the treatment region, the therapeutic device including a radiation source;one or more instructions executable by a processor for retrieving a planned dose spatial distribution for the patient's body, the planned dose spatial distribution corresponding to a planned position for the radiation source; andone or more instructions executable by a processor for displaying, in substantially real time and in spatial registration with each other, the real time treatment region image, the therapeutic device, the radiation source and the planned dose spatial distribution, wherein at least one of the real time treatment region image and the planned dose spatial distribution is displayed as a translucent image. | 1. A computer-readable medium for facilitating performance of medical therapy treatment on a patient, the computer-readable medium comprising: one or more instructions executable by a processor for processing real-time image data to generate in substantially real time an image of a treatment region of a patient's body along with an image of a therapeutic device inserted inside the patient's body relative to the treatment region, the therapeutic device including a radiation source;one or more instructions executable by a processor for retrieving a planned dose spatial distribution for the patient's body, the planned dose spatial distribution corresponding to a planned position for the radiation source; andone or more instructions executable by a processor for displaying, in substantially real time and in spatial registration with each other, the real time treatment region image, the therapeutic device, the radiation source and the planned dose spatial distribution, wherein at least one of the real time treatment region image and the planned dose spatial distribution is displayed as a translucent image. 2. The computer-readable medium of claim 1, further comprising: one or more instructions executable by a processor for retrieving a plan image of the treatment region; andone or more instructions executable by a processor for displaying the plan image concurrently and in spatial registration with the real time treatment region image, the therapeutic device, the radiation source and the planned dose spatial distribution. 3. The computer-readable medium of claim 2, further comprising: one or more instructions executable by a processor for determining an actual position of the radiation source that has been inserted within the treatment region;one or more instructions executable by a processor for calculating an actual dose spatial distribution for the patient's body, the actual dose spatial distribution corresponding to the determined actual position of the radiation source within the treatment region; andone or more instructions executable by a processor for displaying the actual dose spatial distribution concurrently and in spatial registration with the plan image, the real time treatment region image, the therapeutic device, the radiation source and the planned dose spatial distribution. 4. The computer-readable medium of claim 3, further comprising: one or more instructions executable by a processor for retrieving a planned position for the therapeutic device; andone or more instructions executable by a processor for displaying the planned position for the therapeutic device concurrently and in spatial registration with the plan image, the real time treatment region image, the therapeutic device, the radiation source, the planned dose spatial distribution and the actual dose spatial distribution, wherein the planned position for the therapeutic device is displayed as a translucent image. 5. The computer-readable medium of claim 1, further comp... | Computerized Medical Systems, Inc. | A61B005/05
US07309313 | B2 | 2007-12-18 | Vascular disease examining system and bypass vascular diagnosing device | 2002-06-21 | Disclosed are a vascular disease examination system capable of examining and diagnosing vascular diseases such as arterial obliteration by determining the feature quantities of blood flow velocity waveforms or blood pressure waveforms or blood flow rates, and a bypass vascular diagnosing system. A blood flow velocity or blood pressure at a target location of the body is measured A waveform analysis unit 4 determines the feature quantity of a blood flow waveform (or blood pressure waveform), for displaying on an output unit 5. The feature quantity includes a time constant, Fourier transform value, differentiated value, integrated value, rise/fall time, and sharpness or gradient of rise. Feature quantity of a blood flow velocity waveform (or blood pressure waveform) indicated in numeric value enables a correct examination/diagnosis. | 1. A vascular disease examination system, comprising: an equivalent constant calculation means for calculating blood vessel equivalent constants defining a relationship between blood flow velocity and blood pressure based on a blood flow velocity waveform signal and a blood pressure waveform signal output from a measurement system, which measures blood flow velocity values and blood pressure values, anda dynamic blood vessel condition calculation means for calculating blood flow velocity values under a virtual blood pressure condition, in which a load is exerted on a blood vessel using the equivalent constants to find a dynamic blood vessel condition. | 1. A vascular disease examination system, comprising: an equivalent constant calculation means for calculating blood vessel equivalent constants defining a relationship between blood flow velocity and blood pressure based on a blood flow velocity waveform signal and a blood pressure waveform signal output from a measurement system, which measures blood flow velocity values and blood pressure values, anda dynamic blood vessel condition calculation means for calculating blood flow velocity values under a virtual blood pressure condition, in which a load is exerted on a blood vessel using the equivalent constants to find a dynamic blood vessel condition. 2. A bypass vascular diagnosing system, comprising: an equivalent constant calculation means for calculating bypass blood vessel equivalent constants defining a relationship between blood flow velocity and blood pressure based on a blood flow velocity waveform signal and a blood pressure waveform signal output from a measurement system, which measures blood flow velocity values and blood pressure values, anda dynamic bypass vascular condition calculation means for calculating the blood flow velocity values under a virtual blood pressure condition, in which a load is exerted on a bypass blood vessel using the equivalent constants to find a dynamic bypass vascular condition. | Nihon University Tsukuba Technology Seed Co., Ltd. | A61B005/02 A61B008/00 A61B005/00
US07054680 | B1 | 2006-05-30 | Device for detecting electrical potentials in the forehead-area of a patient | 2000-05-04 | The invention relates to a device for detecting electrical potentials on a patient. The device comprises an electrode device that is applied in the area of the forehead, this electrode device being located on a forehead support element (6) which is connected to a breathing mask (1). The electrode device preferably comprises at least two electrode elements (9, 10, 11). According to a preferred embodiment of the invention, these electrode elements are connected to a signal processing device which is located in the direct proximity of the electrode elements. The output of said signal processing device is preferably potential-free and communicates directly with a data transmission device, so that the processed signals can be transmitted to a data processing device in a cordless manner, especially by radio transmission. The invention can be used in a patient monitoring system of a sleep laboratory or in a CPAP device. | 1. An apparatus for detecting brain electrical potentials on a patient, comprising: a breathing mask;a stiff tubular nonround member extending upwardly from said breathing mask and for supplying gas to said breathing mask;an electrode device which can be applied in the forehead region of the patient on said tubular member, the electrode device being arranged on a forehead support element of an elastomeric material which co-operates with the breathing mask in such a way that an application position of the electrode device is established in conjunction with the application position of the breathing mask, said forehead support element being transverse to said tubular member and having free ends on opposite sides of a central portion, said electrode device including at least one electrode element on each of said ends adapted to bear on the forehead region of the patient when said breathing mask is on the face of the patient and another electrode element on said central portion bearing on the forehead region, said breathing mask having a seal engaging the face of the patient around a nose and mouth region, a cavity surrounded by said seal and a drawn-in region receiving the nose of the patient and attached to the seal, the electrode elements being coupled to a signal processing device for processing brain signals. | 1. An apparatus for detecting brain electrical potentials on a patient, comprising: a breathing mask;a stiff tubular nonround member extending upwardly from said breathing mask and for supplying gas to said breathing mask;an electrode device which can be applied in the forehead region of the patient on said tubular member, the electrode device being arranged on a forehead support element of an elastomeric material which co-operates with the breathing mask in such a way that an application position of the electrode device is established in conjunction with the application position of the breathing mask, said forehead support element being transverse to said tubular member and having free ends on opposite sides of a central portion, said electrode device including at least one electrode element on each of said ends adapted to bear on the forehead region of the patient when said breathing mask is on the face of the patient and another electrode element on said central portion bearing on the forehead region, said breathing mask having a seal engaging the face of the patient around a nose and mouth region, a cavity surrounded by said seal and a drawn-in region receiving the nose of the patient and attached to the seal, the electrode elements being coupled to a signal processing device for processing brain signals. 2. An apparatus as set forth in claim 1 wherein the forehead support element is formed in one piece with a mask base member of the breathing mask device. 3. An apparatus as set forth in claim 1 wherein said electrode elements are mounted on an application surface to yield in a direction substantially perpendicular to said application surface. 4. An apparatus as set forth in claim 1 wherein the signal processing device is integrated into the forehead support element. 5. An apparatus as set forth in claim 1 wherein the signal processing device is provided with a data transmission device for the cord-less transmission of the processed signals to a data processing device. 6. A breathing mask arrangement for feeding a respiration gas to a patient under an increased pressure, comprising: a mask member which engages over the nose region of the patient;a stiff tubular nonround member extending upwardly from said mask member;a sealing device for sealing off an inner region of the mask with respect to the ambient atmosphere, anda forehead support element of an elastomeric material for supporting the mask member in the forehead region of the patient, said forehead support element having at least two electrodes said forehead support element being transverse to said tubular member and having free ends on opposite sides of a central portion, said forehead support element having at an electrode on each of said ends and another electrode in said central portion adapted to press against said forehead region for detecting brain-electrical potentials. 7. The arrangement defined in claim 6 wherein the mask member is formed from an elastomer material. 8. ... | Map Medizin-Technologie GmbH | A61B005/04 A62B018/02
US07107086 | B2 | 2006-09-12 | Device for in-vivo measurement of the concentration of a substance contained in a body fluid | 2002-09-05 | A device for the in-vivo measuring of a constituent of a body fluid, the device including a light emitter which emits light in a wavelength range to which the constituent is sensitive and which is implantable in living tissue, and a light detector which, when the device is in a measuring position, receives the light emitted by the emitter and outputs a signal depending on the amount of light received, from which signal the constituent may be measured, wherein the light detector is situated outside the tissue when the device is in the measuring position. | 1. A device for measuring in-vivo the concentration of a constituent of a body fluid, the device comprising: a) an energy source adapted to be situated outside living tissue;b) a light source powered by the energy source, emitting light in a wavelength range to which said constituent is sensitive, and adapted to be implanted in living tissue; andc) a light detector which in a measuring position of said device receives said light and outputs a signal depending on said light received, from which signal the concentration of said constituent may be determined; whereind) said light detector is adapted to be situated outside living tissue in said measuring position, and in said measuring position, said light source and light detector are connected to each other in fixed alignment relative to each other. | 1. A device for measuring in-vivo the concentration of a constituent of a body fluid, the device comprising: a) an energy source adapted to be situated outside living tissue;b) a light source powered by the energy source, emitting light in a wavelength range to which said constituent is sensitive, and adapted to be implanted in living tissue; andc) a light detector which in a measuring position of said device receives said light and outputs a signal depending on said light received, from which signal the concentration of said constituent may be determined; whereind) said light detector is adapted to be situated outside living tissue in said measuring position, and in said measuring position, said light source and light detector are connected to each other in fixed alignment relative to each other. 2. The device as set forth in claim 1, wherein said light source is operably coupled to a skin passage device, wherein said skin passage device, when implanted, forms a body port for supplying or discharging a body fluid. 3. The device as set forth in claim 1, wherein said light detector is operably coupled to a skin passage device, wherein said skin passage device, when implanted, forms a body port for supplying or discharging a body fluid. 4. The device as set forth in claim 1, wherein said constituent is glucose. 5. The device as set forth in claim 1, wherein the device is adapted for use as a measuring device in a control loop of a pump in an infusion device for administering an agent, wherein said agent is administered in doses depending on the concentration of said constituent. 6. The device of claim 1, wherein the energy source is an electric battery. 7. The device of claim 1, wherein the energy source inductively powers the light source. 8. The device of claim 1, further comprising an electrical line interconnecting the energy source to the light source. 9. The device of claim 1, wherein the light source is an infrared laser diode. 10. The device of claim 1, wherein the light source is an array of a number of infrared laser diodes. 11. A device for measuring in-vivo the concentration of a constituent of a body fluid, the device comprising: a light source adapted to emit light;a reflector having a reflective area and adapted to be implanted in living tissue, wherein the degree of reflection for a first portion of the reflective area changes depending on the concentration of the constituent and the degree of reflection for a second portion of the reflective area does not change relative to the concentration of the constituent, and wherein the reflective area reflects the light received from the light source; anda light detector for placement in a receiving position relative to the reflector and adapted to receive light reflected from the reflector and output a signal depending on the light received, the signal pertaining to the concentration of the constituent. 12. The device as set forth in claim 11, wherein the light detector is adapt... | Disetronic Licensing AG | A61B005/00
US07024000 | B1 | 2006-04-04 | Adjustment of a hearing aid using a phone | 2000-06-07 | A system and method for using a telephone to reconfigure or readjust the performance characteristics of a hearing aid or to check whether a user has a hearing problem. The telephone is used to generate one or more frequency tests covering the audible spectrum using a DSP contained in the phone, an external computer and/or a hearing aid. The keypad of the phone or keyboard of an attached computer is used as a feedback mechanism. The generated frequencies can be used to test the hearing of a user and the quality (or fit) of a hearing aid while being worn by the user. A local memory may be used to store the results of the tests for future reference or for transmission over the network for analysis at a later time. Once the hearing response of a user wearing the hearing aid has been measured, an updated compensation configuration (audiogram) can be downloaded into the hearing aid via an infra-red link, via a physical connection or a direct audio transmission from the telephone to the DSP in the hearing aid. | 1. A method for remotely adjusting a hearing aid of a user, comprising the steps of: generating a command via a first computer at a first location;transmitting the command to a second computer at a second location over a remote data link;sending the command from the second computer to a digital signal processor in the hearing aid as a DTMF tone;outputting a test tone from the digital signal processor based on the command to a user of a telephone wearing the hearing aid;receiving a user response to the test tone over the remote data link; andadjusting the hearing aid based on the user response to the test tone, wherein: said adjusting step comprises the steps of: transmitting the user response to the first computer over the remote data link;retrieving a stored audiogram from memory based on an accuracy of the response; anduploading the audiogram into the hearing aid of the user over the remote data link; andsaid audiogram is a compensation curve for adjusting performance characteristics of the hearing aid based on the user response. | 1. A method for remotely adjusting a hearing aid of a user, comprising the steps of: generating a command via a first computer at a first location;transmitting the command to a second computer at a second location over a remote data link;sending the command from the second computer to a digital signal processor in the hearing aid as a DTMF tone;outputting a test tone from the digital signal processor based on the command to a user of a telephone wearing the hearing aid;receiving a user response to the test tone over the remote data link; andadjusting the hearing aid based on the user response to the test tone, wherein: said adjusting step comprises the steps of: transmitting the user response to the first computer over the remote data link;retrieving a stored audiogram from memory based on an accuracy of the response; anduploading the audiogram into the hearing aid of the user over the remote data link; andsaid audiogram is a compensation curve for adjusting performance characteristics of the hearing aid based on the user response. 2. The method of claim 1, wherein said receiving step comprises inputting a response to the command into the second computer via a keyboard attached to the computer. 3. The method of claim 1, wherein said receiving step comprises inputting a response to the command via a key pad on the telephone. 4. A method for adjusting a hearing aid of a user, comprising the steps of: generating a command via a computer;sending the command to a digital signal processor in the hearing aid as a DTMF tone;outputting a test tone from the digital signal processor based on the command to the user of a telephone wearing the hearing aid;receiving a response to the test tone by the user;storing the response to the test tone by the user in the computer;retrieving a stored audiogram from memory based on an accuracy of the stored response; anduploading the audiogram into the hearing aid of the user. 5. The method of claim 4, wherein said receiving step comprises inputting a response to the command into the computer via a keyboard attached to the computer. 6. The method of claim 4, wherein said receiving step comprises inputting a response to the command via a keypad on the telephone. 7. The method of claim 4, wherein said audiogram is a compensation curve for adjusting performance characteristics of the hearing aid based on the user response. 8. The method of claim 4, wherein the command is generated by a first computer at a first location and is received by a second computer at a second location, and said second computer sends the command to the digital processor. 9. The method of claim 4, wherein the step of sending the command to the digital signal processor is by a wireless link. 10. A method for remotely adjusting a hearing aid of a user, comprising the steps of: generating a command via a first computer at a first location;transmitting the command to a second computer at a second location over a remote data link;sending the comma... | Agere Systems Inc. | H04B029/00 A61B005/00 H04M011/00
US06918873 | B1 | 2005-07-19 | Inverted sensor module | 2002-09-19 | An inverted sensor module for sensing physical and chemical fluid characteristics in a vessel in a living body. The inverted sensor module includes an inverted sensor bonded to a hollow carrier so that sensing occurs in the lumen of the inverted sensor module. An inverted sensor module comprising one or more inverted sensors is implantable and allows fluid to flow through, thereby allowing for measurements of acute and chronic conditions. The inverted sensor module may be used in series with other inverted sensor modules to assess trends such as pressure gradients along a bloodstream. | 1. An inverted sensor module for sensing one or more characteristics of fluid in one or more vessels in a living body, the inverted sensor module comprising; a sensor carrier having one or more openings, wherein said sensor carrier is sized for insertion within said one or more vessels, and wherein an inner surface of said sensor carrier is coated or lined with a substance for reducing turbulence of the fluid flowing through a lumen of said sensor carrier;one or more sensors mounted to said sensor carrier, such that a sensing side of said one or more sensors is exposed to the inside of said sensor carrier through said one or more openings;a surface feature formed on each end of said inverted sensor module, such that the ends of a cut vessel may be attached to both ends of said inverted sensor module; andcommunication media operatively connected to said one or more sensors, said communication media operable for communicating sensor information to a monitoring device. | 1. An inverted sensor module for sensing one or more characteristics of fluid in one or more vessels in a living body, the inverted sensor module comprising; a sensor carrier having one or more openings, wherein said sensor carrier is sized for insertion within said one or more vessels, and wherein an inner surface of said sensor carrier is coated or lined with a substance for reducing turbulence of the fluid flowing through a lumen of said sensor carrier;one or more sensors mounted to said sensor carrier, such that a sensing side of said one or more sensors is exposed to the inside of said sensor carrier through said one or more openings;a surface feature formed on each end of said inverted sensor module, such that the ends of a cut vessel may be attached to both ends of said inverted sensor module; andcommunication media operatively connected to said one or more sensors, said communication media operable for communicating sensor information to a monitoring device. 2. The inverted sensor module of claim 1, wherein said one or more sensors sense one or more physical characteristics of said fluid. 3. The inverted sensor module of claim 2, wherein said one or more characteristics comprises flow velocity. 4. The inverted sensor module of claim 2, wherein said one or more characteristics comprises pressure. 5. The inverted sensor module of claim 2, wherein said one or more characteristics comprises temperature. 6. The inverted sensor module of claim 2, wherein said one or more characteristics comprises conductance. 7. The inverted sensor module of claim 1, wherein said one or more sensors sense one or more chemical characteristics of said fluid. 8. The inverted sensor module of claim 7, wherein said one or more characteristics comprises pH. 9. The inverted sensor module of claim 7, wherein said one or more characteristics comprises O2 concentration. 10. The inverted sensor module of claim 7, wherein said one or more characteristics comprises CO2 concentration. 11. The inverted sensor module of claim 7, wherein said one or more characteristics comprises glucose concentration. 12. The inverted sensor module of claim 1, wherein the communication media are coupled to the one or more sensors with wires, fiber optics, or telemetry. 13. The inverted sensor module of claim 1, further comprising a portion of tubing bonded to said inverted sensor module such that fluids may flow through said inverted sensor module, wherein said inverted sensor module is operable to detect the one or more characteristics of the fluid flowing through said inverted sensor module. 14. The inverted sensor module of claim 13, wherein said surface feature on each end of said inverted sensor module comprises a groove. 15. The inverted sensor module of claim 13, wherein said surface feature on each end of said inverted sensor module comprises a lip. 16. The inverted sensor module of claim 13, wherein said surface feature on each end of said inverted sensor module compris... | Millar Instruments, Inc. | A61B005/00 A61B005/02
US06849052 | B2 | 2005-02-01 | Body fluid measuring apparatus with lancet and lancet holder used for the measuring apparatus | 2000-12-13 | A body fluid measuring apparatus, comprising a main body (20) and a lancet holder (30) installed on the main body (20), the lancet holder (30) further comprising a lancet (31) capable of being moved for piercing to allow body fluid to be bled from skin (S) and a biosensor (36) introducing the bled body fluid for a specified measurement, wherein the biosensor (36) is moved in a specified direction by a sensor moving mechanism (40, 42). | 1. A body fluid measuring apparatus comprising: a main body; and a lancet holder for attachment to the main body; wherein the lancet holder includes a lancet movable for stabbing a skin to draw body fluid, and a biosensor for introducing the drawn body fluid to perform predetermined measurement; and wherein a sensor moving mechanism including a motor is further provided for moving the biosensor in a predetermined direction. | 1. A body fluid measuring apparatus comprising: a main body; and a lancet holder for attachment to the main body; wherein the lancet holder includes a lancet movable for stabbing a skin to draw body fluid, and a biosensor for introducing the drawn body fluid to perform predetermined measurement; and wherein a sensor moving mechanism including a motor is further provided for moving the biosensor in a predetermined direction. 2. The body fluid measuring apparatus according to claim 1, wherein the biosensor is shaped like a plate and internally formed with a body fluid path having an inlet end, the biosensor being formed with an operative electrode and a counterpart electrode facing the body fluid path at a position spaced from the inlet end. 3. The body fluid measuring apparatus according to claim 2, further comprising body fluid detecting means for detecting the drawn body fluid, the sensor moving mechanism moving the biosensor in the predetermined direction in response to a detection signal generated by the body fluid detecting means. 4. The body fluid measuring apparatus according to claim 3, wherein the body fluid detecting means comprises a pair of conductance measuring electrodes facing the body fluid path at the inlet end of the biosensor, the body fluid being detected by sensing conduction established between the conductance measuring electrodes due to the body fluid. 5. The body fluid measuring apparatus according to claim 3, wherein the body fluid detecting means is non-contact detection means for detecting existence of the body fluid without contacting the drawn body fluid. 6. The body fluid measuring apparatus according to claim 5, wherein the non-contact detection means comprises a light source for emitting light toward the skin, and a light receiving element for receiving reflected light from the drawn body fluid. 7. The body fluid measuring apparatus according to claim 2, wherein the biosensor comprises a base plate having an upper surface on which the operative electrode and the counterpart electrode are formed, a pair of spacers spaced from each other and disposed on the base plate, and a cover plate laminated on the spacers to bridge the spacers, the body fluid path being defined between the spacers. 8. The body fluid measuring apparatus according to claim 2, wherein the body fluid path is formed with a reactive reagent layer. 9. The body fluid measuring apparatus according to claim 2, wherein the biosensor includes a first electrode facing the body fluid path at the inlet end, and a second electrode facing the body fluid path at an end opposite to the inlet end, filling of the body fluid path with the body fluid being detected by measuring conductance between the first electrode and the second electrode. 10. The body fluid measuring apparatus according to claim 9, wherein the second electrode is one of the operative electrode and the counterpart electrode. 11. The body fluid measuring apparatus according to claim 1, wh... | Arkray, Inc. | A61B005/00
US06882166 | B2 | 2005-04-19 | System and method for measuring the validity of a bioelectric impedance measurement in the presence of interference | 2002-07-03 | A method for measuring bioelectric impedance in real time, in the presence of interference and noise is disclosed. A small electric current is injected into a biopotential electrode system, and then the measurement is tested for contamination by electrical interference or other noise sources. | 1. A method for measuring bioelectric impedance in a biopotential measurement system including at least two electrodes, said method comprising the steps of: measuring bioelectric impedance by measuring a voltage at the frequency of an impedance signal estimating noise voltage by measuring a voltage at a discrete frequency near the frequency of said impedance signal; computing an impedance signal to noise ratio from same impedance voltage measured at the frequency of the impedance signal and said estimated noise voltage; and determining the validity of the bioelectric impedance measurement by comparing said impedance signal to noise ratio against an impedance signal to noise ratio threshold value. | 1. A method for measuring bioelectric impedance in a biopotential measurement system including at least two electrodes, said method comprising the steps of: measuring bioelectric impedance by measuring a voltage at the frequency of an impedance signal estimating noise voltage by measuring a voltage at a discrete frequency near the frequency of said impedance signal; computing an impedance signal to noise ratio from same impedance voltage measured at the frequency of the impedance signal and said estimated noise voltage; and determining the validity of the bioelectric impedance measurement by comparing said impedance signal to noise ratio against an impedance signal to noise ratio threshold value. 2. The method for measuring bioelectric impedance of claim 1 wherein said noise voltage of the bioelectric impedance measurement is estimated by measuring voltages at a number of discrete frequencies in a band near the frequency of the impedance signal, aggregating said voltages at each of said discrete frequencies and using an aggregated voltage value as an estimate of said noise voltage. 3. The method for measuring bioelectric impedance of claim 1 further comprising the step of enabling said system for measuring biopotential impedance if a prior measured bioelectric impedance does not exceed said impedance signal to noise ratio threshold value. 4. The method for measuring bioelectric impedance of claim 1 further comprising the step of enabling said system for measuring biopotential impedance if a prior measured bioelectric impedance exceeds said signal to noise impedance threshold and the measured impedance signal to noise ratio does not exceed the signal to noise ratio threshold. 5. The method for measuring bioelectric impedance of claim 1 further comprising the step of disabling said system for measuring biopotential impedance if a prior measured bioelectric impedance exceeds the impedance threshold and the measured impedance signal to noise ratio exceeds the signal to noise ratio threshold. | Aspect Medical Systems, Inc. | G01R027/08 A61B005/05
US06853190 | B2 | 2005-02-08 | Method and apparatus for magnetic resonance imaging with simultaneous measurement of two neighboring slices | 2002-10-25 | In a magnetic resonance tomography apparatus and a method for the operation thereof, two neighboring slices of an examination subject are simultaneously excited and the measurement time thus can be shortened without SNR loss. The radiofrequency excitation pulse encodes the first slice with a phase offset of +90° or −90° compared to the second slice, so that a response signal contains magnetization information about both slices, [this information] being separated on the basis of the phase offset. | 1. A method of generating a magnetic resonance image of a subject, comprising the steps of: placing an examination subject in a static basic magnetic field; simultaneously exciting nuclear spins in first and second adjacent slices of said subject with a radio frequency excitation pulse that encodes said first slice with a phase offset of +90 degrees or −90 degrees relative to said second slice, and thereby producing a magnetic resonance response signal containing magnetization information for both of said first and second slices; generating gradient pulses that spatially encode said magnetic resonance response signals, thereby producing spatially encoded magnetic resonance response signals; using multiple reception coils having respective coil sensitivities, which receive, said spatially encoded magnetic resonance response signals, as received signals having a real part and an imaginary part; processing said received signals, including separating the respective magnetization information for said first and second slices dependent on said phase offset by calculating respective magnetization densities of said first and second slices from said real part and said imaginary part dependent on said coil sensitivities and making respective entries which represent said received signals, into a first plurality of rows in k space, and comprising the additional step of identifying the respective coil sensitivities by making said entries wherein said first slice is encoded with a complementary phase offset of −90 degrees or +90 degrees relative to said second slice, in a second plurality of rows of k space; selectively combining said first plurality of rows and said second plurality of rows in identifying the respective coil sensitivities; and generating respective magnetic resonance images of said first slice and said second slice. | 1. A method of generating a magnetic resonance image of a subject, comprising the steps of: placing an examination subject in a static basic magnetic field; simultaneously exciting nuclear spins in first and second adjacent slices of said subject with a radio frequency excitation pulse that encodes said first slice with a phase offset of +90 degrees or −90 degrees relative to said second slice, and thereby producing a magnetic resonance response signal containing magnetization information for both of said first and second slices; generating gradient pulses that spatially encode said magnetic resonance response signals, thereby producing spatially encoded magnetic resonance response signals; using multiple reception coils having respective coil sensitivities, which receive, said spatially encoded magnetic resonance response signals, as received signals having a real part and an imaginary part; processing said received signals, including separating the respective magnetization information for said first and second slices dependent on said phase offset by calculating respective magnetization densities of said first and second slices from said real part and said imaginary part dependent on said coil sensitivities and making respective entries which represent said received signals, into a first plurality of rows in k space, and comprising the additional step of identifying the respective coil sensitivities by making said entries wherein said first slice is encoded with a complementary phase offset of −90 degrees or +90 degrees relative to said second slice, in a second plurality of rows of k space; selectively combining said first plurality of rows and said second plurality of rows in identifying the respective coil sensitivities; and generating respective magnetic resonance images of said first slice and said second slice. 2. A method as claimed in claim 1 wherein said entries in said second plurality of rows of k space include magnitude information, and comprising selectively combining said first plurality of rows and said second plurality of rows independently of said magnitude information. 3. A method as claimed in claim 1 comprising selectively combining said first plurality of rows and said second plurality of rows based on a criteria that a phase of coil profiles between said first slice and said second slice is equal to exactly +90° or −90° in regions of said first and second slices having an intensity of said received signal below a predetermined threshold. 4. A method as claimed in claim 1 comprising employing a spin-echo sequence for generating said radiofrequency excitation pulses and said gradient pulses. 5. An apparatus generating a magnetic resonance image of a subject, comprising: a basic field magnet which generates a static basic magnetic field and configured to receive an examination subject therein; a radio-frequency system that simultaneously excites nuclear spins in first and second adjacent slices of said subject in said ba... | Siemens Aktiengesellschaft | G01V003/00 A61B005/05
US06933720 | B2 | 2005-08-23 | Sequence preconditioning for ultra-fast magnetic resonance imaging | 2001-12-11 | An improved magnetic resonance imaging (MRI) methodology uses an abbreviated initial MRI sequence to generate sequence diagnostic parameters. The sequence diagnostic parameters have a fixed relationship to certain sequence-conditioning parameters, and are used for calculating characteristic values of the sequence-conditioning parameters. The read out gradient pulse sequence is modified in accordance with the calculated characteristic values of the sequence-conditioning parameters. The modified read out gradient pulse sequence is then incorporated into a subsequent MRI pulse sequence used for obtaining a diagnostic image. The methodology has particular application in so called ultra fast MRI process which include echo-planar imaging (EPI) and echo-volume imaging (EVI). | 1. A method of improving the quality of magnetic resonance images by preconditioning an EPI pulse sequence for use in an ultra-fast MRI system so as to negate abnormalities in gradient fields, said method comprising: (i) performing a modified EPI pulse sequence with the ultra-fast MRI system with phase encoding gradients switched off thereby obtaining collective sequence diagnostic parameters, said collective sequence diagnostic parameters comprising at least one of: a collective deviation of the echo centers, from the center of the echo acquisition window by Δ, a collective shift of the even-echo centers from the odd-echo centers by δ; and a collective linear drift of the echo centers by an amount of σ per echo; (ii) computing sequence-preconditioning parameters for the EPI sequence from said obtained collective sequence diagnostic parameters; and (iii) modifying a subsequently performed EPI pulse sequence in accordance with said computed preconditioning parameters as the subsequent EPI pulse sequence is performed and thereby obtaining an improved magnetic resonance image when the subsequent EPI pulse sequence is performed. | 1. A method of improving the quality of magnetic resonance images by preconditioning an EPI pulse sequence for use in an ultra-fast MRI system so as to negate abnormalities in gradient fields, said method comprising: (i) performing a modified EPI pulse sequence with the ultra-fast MRI system with phase encoding gradients switched off thereby obtaining collective sequence diagnostic parameters, said collective sequence diagnostic parameters comprising at least one of: a collective deviation of the echo centers, from the center of the echo acquisition window by Δ, a collective shift of the even-echo centers from the odd-echo centers by δ; and a collective linear drift of the echo centers by an amount of σ per echo; (ii) computing sequence-preconditioning parameters for the EPI sequence from said obtained collective sequence diagnostic parameters; and (iii) modifying a subsequently performed EPI pulse sequence in accordance with said computed preconditioning parameters as the subsequent EPI pulse sequence is performed and thereby obtaining an improved magnetic resonance image when the subsequent EPI pulse sequence is performed. 2. The method of claim 1, wherein said collective sequence diagnostic parameters are being obtained by determining the centers of each of the echo signals, fitting the echo centers to two separate lines—one for all odd echoes and another for all even echoes Ciodd=aodd×i+bodd i=0,2,4, . . . Cieven=aeven×i+beven i=1,3,5, . . . and computing at least one of said collective diagnostic parameters in accordance with the following Δ= ⁢bodd+beven2δ=beven-boddσ=aodd+aeven2. 3. The method of claim 2 wherein said computing process includes at least one of: (a) computing a mistuning of a first echo by an amount of ΔM in gradient moment which has units of gauss·ms·cm−1; (b) computing a delay of a gradient waveform by an amount d; and (c) computing an offset of a readout gradient by an amount of g0, wherein Δ⁢ ⁢M=∫0Δ⁢γ⁢ ⁢Gro⁢(t)⁢ⅆtd=-δ2g0=-∫0σ⁢Gro⁢(t)⁢ⅆt⁢ Δ⁢ ⁢TEwhere ΔTE is a time interval between two consecutive echoes. 4. The method of claim 1 wherein said modifying process involves modifying a gradient read out pulse sequence of the EPI pulse sequence. 5. The method of claim 1, said EPI pulse sequence being used in an echo-planar imaging process. 6. The method of claim 1, said EPI pulse sequence being used in an echo-volume imaging process. 7. An apparatus improving the quality of magnetic resonance images by preconditioning an EPI pulse sequence in an ultra-fast MRI system so as to negate abnormalities in gradient fields, said apparatus comprising: sequence diagnostic means for obtaining a modified EPI pulse sequence with the ultra-fast MRI system with phase encoding gradients switched off thereby obtaining collective sequence diagnostic parameters, said collective sequence diagnostic parameters comprising at least one of: a collective deviation of the echo centers from the center of the echo acquisition window by Δ, a collec... | Toshiba America MRI, Inc. | G01V003/00 A61B005/055
US06898459 | B2 | 2005-05-24 | System and method for diagnosing pathologic heart conditions | 2001-02-23 | A method of diagnosing pathologic heart conditions in which a time series of heart sounds is filtered and parsed into a sequence of individual heart cycles. A systolic interval as well as systolic sub-intervals are identified for each heart cycle. An energy value is computed for the systolic sub-interval of one or more heart cycles. The energy value computed is proportional to the energy level associated with the filtered series of heart sounds. A composite energy value is then computed for the systolic sub-intervals of one or more heart cycles and compared to a threshold level in order to distinguish between a normal heart and a pathologic heart. The system corresponding to the method is comprised of a portable computing device that manages data collection and stores data collected from new patients, and analyzes data. | 1. A method of diagnosing pathologic heart conditions comprising: identifying a systolic sub-interval of a systolic interval for a plurality of heart cycles in a sequence of heart cycles; computing an energy value for each systolic sub-interval; computing a composite energy value using the computed energy values for each systolic sub-interval; and comparing the composite energy value to a threshold level in order to distinguish between a normal heart and a pathologic heart. | 1. A method of diagnosing pathologic heart conditions comprising: identifying a systolic sub-interval of a systolic interval for a plurality of heart cycles in a sequence of heart cycles; computing an energy value for each systolic sub-interval; computing a composite energy value using the computed energy values for each systolic sub-interval; and comparing the composite energy value to a threshold level in order to distinguish between a normal heart and a pathologic heart. 2. A method of diagnosing pathologic heart conditions comprising: filtering a time series of heart sounds; parsing the time series of heart sounds into a sequence of individual heart cycles; identifying a systolic interval for each heart cycle; identifying a systolic sub-interval of the systolic interval for each heart cycle; computing an energy value for the systolic sub-interval of one or more heart cycles, said energy value being proportional to the energy level associated with the filtered series of heart sounds; computing a composite energy value for the systolic sub-intervals of one or more heart cycles; and comparing the composite energy value to a threshold level in order to distinguish between a normal heart and a pathologic heart. 3. The method of claim 2 wherein said parsing step uses electro-cardiogram (ECG) data in order to transform a time series of heart sounds into a sequence of individual heart cycles. 4. The method of claim 2 wherein said parsing step uses acoustic heart sounds obtained directly from a patient in order to transform a time series of heart sounds into a sequence of individual heart cycles. 5. The method of claim 2 wherein identifying a systolic interval for each heart cycle is achieved by identifying pulses on an electro-cardiogram (ECG). 6. The method of claim 2 wherein identifying a systolic interval for each heart cycle is achieved by acoustically locating a first and a second heart sound using a bandpass filter, said bandpass filter applied to the time series of heart sounds. 7. The method of claim 2 wherein filtering the time series of heart sounds is achieved using a bandpass filter. 8. The method of claim 2 wherein filtering the time series of heart sounds is achieved using time-frequency transforms. 9. The method of claim 8 wherein the time-frequency transform is a wavelet transform. 10. The method of claim 8 wherein the time-frequency transform is a Fourier transform. 11. The method of claim 2 wherein the systolic sub-interval is centered in systole. 12. The method of claim 2 wherein the systolic sub-interval is centered in systole and is half of the systolic interval. 13. The method of claim 2 wherein the composite energy value is computed as the median of the computed energy values for more than one of the systolic sub-intervals of the heart cycles. 14. The method of claim 2 wherein the composite energy value is computed as the weighted average of more than one of the computed energy values for the systolic sub-interva... | The Johns Hopkins University | A61B005/04
US06887210 | B2 | 2005-05-03 | Devices and methods for obtaining mammary fluid samples for evaluating breast diseases, including cancer | 2001-11-13 | Biological samples of mammary fluid or components thereof are obtained using a breast pump device coupled with a solid phase sample collection medium, alternatively facilitated by administering oxytocin to the subject. The breast pump device stimulates expression of mammary fluid and provides for collection of diagnostic samples to evaluate breast disease, including cancer. The biological sample may include whole cells or cellular components, purified or bulk proteins, glycoproteins, peptides, nucleotides or other desired constituents comprising a breast disease marker. Methods, kits and adapter devices relating to the breast pump device are also provided. Yet additional methods, devices, accessories, and materials are provided for laboratory handling and processing of breast fluid samples and for related diagnostic methods. | 1. A sample collection device for collecting a biological sample from a mammary organ of a patient, comprising: a breast engaging member constructed of a non-porous material sized and dimensioned to receive at least a nipple portion of a breast of said patient and form a suction seal therewith; a solid phase sample collection medium in fluid connection with said breast engaging member for receiving a sample of expressed breast fluid; vacuum pump means in gaseous connection with said breast engaging member for generating negative pressure through the breast engaging member to facilitate breast fluid expression; and a hand-held breast pump incorporating said breast engaging member and vacuum pump means in a compact, structurally integrated breast fluid collection apparatus that can be manipulated and operated with one hand, wherein the hand-held breast pump comprises a fluid-retaining well comprising an integral, defined compartment or enclosure within a sample collection housing for receipt of breast fluid and/or constituent samples thereof within a removable fluid reservoir member of the sample collection housing. | 1. A sample collection device for collecting a biological sample from a mammary organ of a patient, comprising: a breast engaging member constructed of a non-porous material sized and dimensioned to receive at least a nipple portion of a breast of said patient and form a suction seal therewith; a solid phase sample collection medium in fluid connection with said breast engaging member for receiving a sample of expressed breast fluid; vacuum pump means in gaseous connection with said breast engaging member for generating negative pressure through the breast engaging member to facilitate breast fluid expression; and a hand-held breast pump incorporating said breast engaging member and vacuum pump means in a compact, structurally integrated breast fluid collection apparatus that can be manipulated and operated with one hand, wherein the hand-held breast pump comprises a fluid-retaining well comprising an integral, defined compartment or enclosure within a sample collection housing for receipt of breast fluid and/or constituent samples thereof within a removable fluid reservoir member of the sample collection housing. 2. The sample collection device of claim 1, wherein the removable reservoir member is a rigid sample collection tube or vial removably connected with an outer casing member of the housing that partially or completely encloses the tube or vial. 3. The sample collection device of claim 1, wherein the removable reservoir member is a rigid sample collection tube removably, sealably connected with an outer casing member of the housing to form an airtight coupling therewith. 4. The sample collection device of claim 1, wherein the removable reservoir member is a cytology vial sealably connected with an outer casing member of the housing to form an airtight coupling therewith. 5. The sample collection device of claim 4, wherein the removable reservoir member and outer casing member of the housing are coupled to form an assembled sample collection housing, wherein the reservoir member is removably nested within the casing member to form a substantially airtight contact between an inner wall of the casing member and an outer wall of the reservoir member. 6. The sample collection device of claim 5, wherein an outer wall of the removable reservoir member features a circumferential O-ring that engages and makes a circumferential airtight seal against the inner wall of the casing member when the vial is nested within the casing member. 7. The sample collection device of claim 1, wherein the removable reservoir member is gaseously and fluidly connected with the breast engaging member to facilitate sample collection. 8. The sample collection device of claim 1, wherein vacuum pressure from the vacuum pump means is routed to the breast engaging member through the removable reservoir member of the housing. 9. The sample collection device of claim 8, wherein the removable reservoir member is modified to include one or more air ports that form a ga... | Atossa Healthcare, Inc. | A61B005/00 B65D081/00
US06934579 | B2 | 2005-08-23 | Anaesthesia control system | 2001-10-09 | An anaesthesia control system and a method of calculating an index representative of the depth of anaesthesia is disclosed. The method comprises subjecting a patient to a repetitive audio stimulus and monitoring and auditory evoked potentials (AEP) produced by the patient and then recording these auditory evoked potentials using EEG recording means and providing a signal corresponding to the coarseness of the monitored AEP signal and using this signal as an index indicative of anaesthetic depth. The raw AEP signal is divided into a series of sweeps and each sweep is synchronized with the repetitive audio stimulus and sweeps are recorded in sequence to produce a time averaged sweep from which the anaesthetic index is calculated. The anaesthetic index is constantly updated by repeatedly conducting a successive series of sweeps. The system and index signal can be used as part of an anaesthesia control system for regulating the supply of anaesthetic to the patient to maintain the anaesthetic index at a predetermined level. | 1. A method of calculating an index indicative of anaesthetic depth, the method comprising: subjecting a patient to a repetitive audio stimulus; monitoring auditory evoked potentials (AEP) produced by the patient; providing a signal corresponding to the coarseness of the monitored AEP signal, the coarseness of the signal being a single measure increasing and decreasing with both amplitude and frequency of variations in the signal AEP signal; and using said signal as said index indicative of anaesthetic depth. | 1. A method of calculating an index indicative of anaesthetic depth, the method comprising: subjecting a patient to a repetitive audio stimulus; monitoring auditory evoked potentials (AEP) produced by the patient; providing a signal corresponding to the coarseness of the monitored AEP signal, the coarseness of the signal being a single measure increasing and decreasing with both amplitude and frequency of variations in the signal AEP signal; and using said signal as said index indicative of anaesthetic depth. 2. A method as claimed in claim 1 wherein the monitored or raw AEP signal is divided into a series of sweeps or frames of a given duration, each sweep being synchronised with the repetitive audio stimulus. 3. A method as claimed in claim 2 wherein a number of sweeps n are recorded in sequence and are averaged to produce a time averaged sweep and for the time averaged sweep the anaesthesia index is calculated. 4. A method as claimed in claim 3 wherein each time a new series of sweeps is recorded, a new time averaged sweep is determined from the most recent n sweeps and the anaesthesia index for that time averaged sweep calculated. 5. A method as claimed in claim 4, wherein for a moving time averaged sweep this measure is a function of the sum of the square roots of the difference between every two adjacent sample points in the time averaged sweep. 6. A method as claimed in claim 2 wherein the raw AEP signal is sampled at regular intervals to produce a digitized AEP signal. 7. A method as claimed in claim 2 wherein an indication of coarseness is obtained by measuring the differences between neighboring sample points. 8. A method as claimed in claim 3 wherein the raw AEP signal is sampled at regular intervals to produce a digitized AEP signal. 9. A method as claimed in claim 3 wherein an indication of coarseness is obtained by measuring the differences between neighboring sample points. 10. A method as claimed in claim 3, wherein said index is calculated and used without reference to a measured amplitude of the time averaged sweep. 11. A method as claimed in claim 3, wherein said index is calculated and used without reference to a measured EEG power. 12. A method as claimed in claim 4 wherein the raw AEP signal is sampled at regular intervals to produce a digitized AEP signal. 13. A method as claimed in claim 4 wherein an indication of coarseness is obtained by measuring the differences between neighboring sample points. 14. A method as claimed in claim 1 wherein the raw AEP signal is sampled at regular intervals to produce a digitized AEP signal. 15. A method as claimed in claim 1 wherein an indication of coarseness is obtained by measuring the differences between neighboring sample points. 16. A method as claimed in claim 1, wherein said index is calculated and used without reference to a measured latency of any peak in the AEP signal. 17. A method as claimed in claim 1, wherein said index is calculated and used without relyi... | The University Court of the University of Glasgow | A61B005/04
US06896656 | B2 | 2005-05-24 | Neurological testing apparatus | 2002-11-14 | An apparatus for performing neuropsychological tests on a patient includes a microprocessing unit adapted for carrying out test procedures which is connected to an interface to a sensory signal generating device comprising at least one of a visual information display, an indication light, and a sound generator, an interface to a reaction detecting device comprising at least one of a dynamometer, a knob, and a touch screen, and an interface to a memory for reading test procedures stored in the memory. At least one test procedure comprising instructions for the generation of sensory signals and the detection of a patient reaction is stored in the memory and is performed by generating the sensory signals on the sensory generating device and by detecting a patient's reaction through the reaction detecting device. | 1. Apparatus for performing neuropsychological tests on a patient, comprising: means for generating a plurality of sensory signals comprising visual signals, auditory signals, or both visual and auditory signals; means for determining perceptual-motor function in a neuropsychological test by detecting physical movement of a patient in response to said generated plurality of sensory signals, said detected physical movement comprising at least one of (i) the trajectory or speed of a rotating motion of a rotary knob turned by the patient and (ii) the accuracy of a path traced by the patient on a touch screen; and means for reporting information on the detected physical movement of the patient through an interface to a computer or network of computers, or both. | 1. Apparatus for performing neuropsychological tests on a patient, comprising: means for generating a plurality of sensory signals comprising visual signals, auditory signals, or both visual and auditory signals; means for determining perceptual-motor function in a neuropsychological test by detecting physical movement of a patient in response to said generated plurality of sensory signals, said detected physical movement comprising at least one of (i) the trajectory or speed of a rotating motion of a rotary knob turned by the patient and (ii) the accuracy of a path traced by the patient on a touch screen; and means for reporting information on the detected physical movement of the patient through an interface to a computer or network of computers, or both. 2. Apparatus according to claim 1, wherein the plurality of sensory signals is determined prior to performing a test on a patient. 3. Apparatus according to claim 1, wherein visual signals and auditory signals are simultaneously generated. 4. Apparatus according to claim 1, further comprising means for storing, processing, or both storing and processing information obtained from the means for detecting physical movement. 5. Apparatus according to claim 1, wherein said detected physical movement comprises the trajectory or speed of a rotating motion of the rotary knob turned by the patient. 6. Apparatus according to claim 1, wherein said detected physical movement comprises the accuracy of a path traced by the patient on the touch screen. 7. The apparatus of claim 1, wherein the information is reported through a network of computers. 8. The apparatus of claim 7, wherein the information is reported through a modem. 9. The apparatus of claim 1, wherein the test procedure is selected or modified from a remote location. 10. Apparatus for performing neuropsychological tests on a patient, comprising a microprocessing unit adapted for carrying out test procedures, said microprocessing unit being connected to an interface to a sensory signal generating device comprising at least one of a visual information display, an indication light, and a sound generator; an interface to a reaction detecting device for determining perceptual-motor function in a neuropsychological test, said reaction detecting device comprising at least one of (i) a rotary knob that measures the trajectory or speed of a rotating motion of the rotary knob turned by the patient and (ii) a touch screen that measures the accuracy of a path traced by the touch screen; and an interface to a memory for reading test procedures stored in the memory, wherein at least one test procedure comprising instructions for the generation of sensory signals and the detection of a patient reaction is stored in the memory and is performed by generating the sensory signals on the sensory generating device and by detecting a patient's reaction through the reaction detecting device, and and interface to a computer or network of computers, or both, w... | - | A61B005/00
US06852083 | B2 | 2005-02-08 | System and method of determining whether to recalibrate a blood pressure monitor | 2002-01-17 | A monitor for determining a patient's physiological parameter includes a calibration device that provides a calibration signal indicative of an accurate measurement of the patient's physiological parameter. The monitor also includes a processor, which receives a noninvasive signal from a noninvasive sensor positioned over a blood vessel. The processor uses the calibration signal to calibrate a relationship between the noninvasive signal and a property of the physiological parameter. The processor also determines when to recalibrate the relationship. | 1. A method of determining whether to recalibrate blood pressure monitor which calculates blood pressure using a calibration signal from a calibration device indicative of an accurate measurement of the blood pressure of a patient and using a continuous signal from a sensor indicative of a continuous noninvasive measurement of one or more attributes of one or more waves propagating along an artery of the patient, the method characterized by: approximating a function relating a change in a trigger parameter to a change in a pressure-velocity relationship, wherein the pressure-velocity relationship is used to noninvasively determine the blood pressure of a patient; calibrating the pressure-velocity relationship using a calibration signal indicative of a first accurate measurement of the blood pressure, and using a continuous signal indicative of a continuous noninvasive measurement of one or more attributes of one or more waves propagating along an artery of a patient; calculating a first value of the trigger parameter from the continuous signal; calculating the blood pressure using the pressure-velocity relationship and the continuous signal; calculating a second value of the trigger parameter from the continuous signal; calculating a change in the pressure-velocity relationship by applying the pressure-velocity relationship and the first and second values of the trigger parameter to the function; and when the change in the pressure-velocity relationship exceeds a threshold value, recalibrating the pressure-velocity relationship using the continuous signal and using a calibration signal indicative of a second accurate measurement of the blood pressure. | 1. A method of determining whether to recalibrate blood pressure monitor which calculates blood pressure using a calibration signal from a calibration device indicative of an accurate measurement of the blood pressure of a patient and using a continuous signal from a sensor indicative of a continuous noninvasive measurement of one or more attributes of one or more waves propagating along an artery of the patient, the method characterized by: approximating a function relating a change in a trigger parameter to a change in a pressure-velocity relationship, wherein the pressure-velocity relationship is used to noninvasively determine the blood pressure of a patient; calibrating the pressure-velocity relationship using a calibration signal indicative of a first accurate measurement of the blood pressure, and using a continuous signal indicative of a continuous noninvasive measurement of one or more attributes of one or more waves propagating along an artery of a patient; calculating a first value of the trigger parameter from the continuous signal; calculating the blood pressure using the pressure-velocity relationship and the continuous signal; calculating a second value of the trigger parameter from the continuous signal; calculating a change in the pressure-velocity relationship by applying the pressure-velocity relationship and the first and second values of the trigger parameter to the function; and when the change in the pressure-velocity relationship exceeds a threshold value, recalibrating the pressure-velocity relationship using the continuous signal and using a calibration signal indicative of a second accurate measurement of the blood pressure. 2. The method of claim 1, further comprising updating the function for each patient with the calibration signal when the function produces unacceptable results. 3. The method of claim 1, wherein the trigger parameter comprises a measurement of dispersion. 4. The method of claim 3, wherein calculating the trigger parameter comprises: calculating propagation velocities along the artery for induced perturbations of a plurality of frequencies over a pulse of blood pressure; selecting a representative value of the propagation velocities; calculating a linear regression of the representative value versus data from the induced perturbations; and calculating the slope of the linear regression. 5. The method of claim 4, wherein the representative value comprises the mean value of the propagation velocities over the pulse of blood pressure. 6. The method of claim 1, wherein the trigger parameter comprises a measurement of attenuation. 7. The method of claim 6, wherein calculating the trigger parameter comprises: calculating propagation phase delays along the artery and propagation amplitudes for induced perturbations of a plurality of frequencies over a pulse of blood pressure; calculating a representative value including a measure of the propagation amplitudes; calculating a linear regression of the ... | Masimo Corporation | A61B005/00
US06890303 | B2 | 2005-05-10 | Implantable device for monitoring aneurysm sac parameters | 2002-04-04 | A device for measuring physiological parameters within an aneurysm sac that has been excluded from blood flow by an endoprosthesis. The device is comprised of at least two sensors, one placed in the aneurysm sac and another in a systemic artery. A differential between the readings of the two sensors can then be calculated, making the device easily calibrated in vivo and insensitive to changes in atmospheric pressure. | 1. A method of calibrating at least two sensors disposed within a human body, comprising: a. measuring a physiological parameter in said human body to obtain a reference value for said physiological parameter, b. providing a first sensing means disposed within an aneurysm sac, c. providing a second sensing means disposed within an artery outside of said aneurysm sac, d. providing a means for calculating a differential value between said first sensing means and said second sensing means, e. providing a means for synchronizing said second sensing means with said reference value whereby said second sensing means is calibrated to said reference value, f. providing a means for calculating a value of said first sensing means by subtracting said differential value from said reference value. | 1. A method of calibrating at least two sensors disposed within a human body, comprising: a. measuring a physiological parameter in said human body to obtain a reference value for said physiological parameter, b. providing a first sensing means disposed within an aneurysm sac, c. providing a second sensing means disposed within an artery outside of said aneurysm sac, d. providing a means for calculating a differential value between said first sensing means and said second sensing means, e. providing a means for synchronizing said second sensing means with said reference value whereby said second sensing means is calibrated to said reference value, f. providing a means for calculating a value of said first sensing means by subtracting said differential value from said reference value. 2. The method of claim 1, wherein said physiological parameter is a pressure. 3. The method of claim 1, wherein said physiological parameter is a flow velocity. 4. The method of claim 1, wherein said physiological parameter is a pressure waveform. 5. The method of claim 1, wherein said first sensing means is disposed within said aneurysm sac by means of fluoroscopic guidance. 6. The method of claim 1, wherein said first sensing means is selected from the group consisting of piezoelectric semiconductor, catheter, acoustic, and ultrasonic sensors. 7. The method of claim 1, wherein said second sensing means is selected from the group consisting of piezoelectric, semiconductor, catheter, acoustic, and ultrasonic sensors. 8. A method for measuring a physiological parameter in an aneurysm sac, comprising the steps of: a. chronically implanting a first sensing means within an aneurysm sac, b. chronically implanting a second sensing means within an artery outside of said aneurysm sac, c. calibrating said first sensing means using at least a first signal from said second sensing means, d. measuring said physiological parameter using a signal from the first sensing means and a second signal from the second sensing means. 9. The method of claim 8, wherein said physiological parameter is a pressure. 10. The method of claim 8, wherein said physiological parameter is a flow velocity. 11. The method of claim 8, wherein said physiological parameter is a pressure waveform. 12. The method of claim 8, wherein said first sensing means is selected from the group consisting of piezoelectric, semiconductor, catheter, acoustic, and ultrasonic sensors. | - | A61B005/00
US06839582 | B2 | 2005-01-04 | Pulse oximetry method and system with improved motion correction | 2002-08-12 | A pulse oximetry method and system for improved motion correction is disclosed. The method/system provides for the use of a detector output signal to obtain a different plurality of differential absorption data sets in corresponding relation to each of a succession of measurement, wherein each of the data sets includes differential absorption values for light of a first wavelength and light of a second wavelength. The data sets are processed to obtain a relative motion estimate value for each measurement. When the relative motion estimate value for a given measurement falls within a predetermined range (i.e., corresponding with clinical motion), a corresponding blood analyte indicator value is adjusted in a predetermined manner, wherein the corresponding adjusted blood analyte indicator is employable to obtain at least one blood analyte concentration value. In one embodiment, blood analyte indicator values may be readily multiplied by a predetermined adjustment factor (i.e., when clinical motion is identified). The relative motion estimate value for a given measurement may be obtained by conducting a principal component analysis of the corresponding plurality of data sets relative to a corresponding best fit function therefor to obtain corresponding variance values V1, V2. The variance value V1, and/or V2 for a given current measurement may be employed to obtain a current motion estimate value. The current motion estimate value and the relative motion estimate value obtained for a prior low motion measurement (i.e., for which no adjustment was necessary) may be used to compute the relative motion estimate value for the current measurement. The variance values V1 and/or V2 are also employable to compute an ongoing, updated motion probability factor, wherein such factor may be used to adjust relative motion estimates values in instances of rapid tissue perfusion changes. | 1. A method for use in a pulse oximetry system which provides a detector output indicative of light absorption by a tissue-under-test at each of a plurality of different light wavelengths and which employs the detector output to compute blood analyte indicator values for each of a plurality of measurements, said method comprising the steps of: utilizing said detector output to obtain a corresponding relative motion estimate value for each of said plurality of measurements; and, determining: (i) whether the corresponding relative motion estimate value for each of said plurality of measurements is within a first predetermined range, wherein for each measurement having a corresponding relative motion estimate value within the first predetermined range the corresponding blood analyte indicator value is adjusted in a first adjustment manner, and wherein said adjusted blood analyte indicator value is employable to obtain at least one blood analyte concentration value; (ii) whether the relative motion estimate value corresponding with each of said plurality of measurements is within a second predetermined range, non-overlapping with said first predetermined range, wherein for each measurement having a corresponding relative motion estimate value within said second predetermined range the corresponding blood analyte indicator value is employable to obtain at least one blood analyte concentration value without adjustment to the corresponding blood analyte indicator value; and (iii) whether the relative motion estimate value corresponding with each of said plurality of measurements is within a third predetermined range, non-overlapping with said first and second predetermined ranges, wherein for each measurement having a corresponding relative motion estimate value within the third predetermined range a previously determined blood analyte indicator value corresponding to a measurement within said first or second predetermined ranges is adjusted in a second adjustment manner different from said first adjustment manner, and wherein said adjusted previously determined blood analyte indicator value is employable to obtain at least one blood analyte concentration value. | 1. A method for use in a pulse oximetry system which provides a detector output indicative of light absorption by a tissue-under-test at each of a plurality of different light wavelengths and which employs the detector output to compute blood analyte indicator values for each of a plurality of measurements, said method comprising the steps of: utilizing said detector output to obtain a corresponding relative motion estimate value for each of said plurality of measurements; and, determining: (i) whether the corresponding relative motion estimate value for each of said plurality of measurements is within a first predetermined range, wherein for each measurement having a corresponding relative motion estimate value within the first predetermined range the corresponding blood analyte indicator value is adjusted in a first adjustment manner, and wherein said adjusted blood analyte indicator value is employable to obtain at least one blood analyte concentration value; (ii) whether the relative motion estimate value corresponding with each of said plurality of measurements is within a second predetermined range, non-overlapping with said first predetermined range, wherein for each measurement having a corresponding relative motion estimate value within said second predetermined range the corresponding blood analyte indicator value is employable to obtain at least one blood analyte concentration value without adjustment to the corresponding blood analyte indicator value; and (iii) whether the relative motion estimate value corresponding with each of said plurality of measurements is within a third predetermined range, non-overlapping with said first and second predetermined ranges, wherein for each measurement having a corresponding relative motion estimate value within the third predetermined range a previously determined blood analyte indicator value corresponding to a measurement within said first or second predetermined ranges is adjusted in a second adjustment manner different from said first adjustment manner, and wherein said adjusted previously determined blood analyte indicator value is employable to obtain at least one blood analyte concentration value. 2. The method of claim 1 wherein, for each measurement having a corresponding relative motion estimate value within the third predetermined range, said second adjustment manner includes: adjusting the previously determined blood analyte indicator value utilizing an adjustment factor. 3. The method of claim 2 further comprising: utilizing said detector output signal to obtain moving average values of total tissue absorption of AC and DC signal components corresponding with each of said plurality of different light wavelengths, wherein said moving average values are determined for each of said plurality of measurements; computing a first ratio of the moving average values obtained for the measurement corresponding to the previously determined blood analyte indicator value; computing a second ... | Datex-Ohmeda, Inc. | A61B005/00
US06865414 | B1 | 2005-03-08 | Apparatus and method for automatically sensing threshold histogram with differentiation of sinus from ectopic beats | 2001-09-20 | An implantable programmable cardiac stimulation device and associated method for differentiating between normal sinus events and ectopic beats. The stimulation device monitors the sensing thresholds of sinus and non-sinus cardiac events, and stores a history of these sensing thresholds along with temporal data for accurate event detection. The stimulation device further provides accurate and appropriate detection of sensed events including P-waves, non-conducted PACs, and conducted PACs and thus verifies correct detection of PVCs and R-waves. Furthermore, the present invention provides a history record of ectopic events, distinguished by sensing thresholds and timing intervals, giving a valuable diagnostic tool to the physician in optimizing rhythm management therapy. In addition, the stimulation device allows the sensitivity threshold to be set based on a single cardiac cycle and past history. | 1. A method for use in a cardiac stimulation device to differentiate between sinus events and ectopic events, comprising: sensing sinus events and ectopic events; automatically determining sensing thresholds of the sinus events; automatically determining sensing thresholds of the ectopic events; storing the sensing thresholds of the sinus events, the sensing thresholds of the ectopic events, and timing relationships that define sinus intervals and ectopic intervals; classifying a sensed cardiac event as a sinus event or an ectopic event based on the proximity of an amplitude of the cardiac event to any of the sensing thresholds of the sinus events or the sensing thresholds of the ectopic events; and classifying a sensed cardiac event as a sinus event or an ectopic event based on the proximity of the sensed cardiac event to a previous average cycle length in a corresponding cardiac chamber. | 1. A method for use in a cardiac stimulation device to differentiate between sinus events and ectopic events, comprising: sensing sinus events and ectopic events; automatically determining sensing thresholds of the sinus events; automatically determining sensing thresholds of the ectopic events; storing the sensing thresholds of the sinus events, the sensing thresholds of the ectopic events, and timing relationships that define sinus intervals and ectopic intervals; classifying a sensed cardiac event as a sinus event or an ectopic event based on the proximity of an amplitude of the cardiac event to any of the sensing thresholds of the sinus events or the sensing thresholds of the ectopic events; and classifying a sensed cardiac event as a sinus event or an ectopic event based on the proximity of the sensed cardiac event to a previous average cycle length in a corresponding cardiac chamber. 2. The method according to claim 1, wherein the step of automatically determining sensing thresholds of the sinus events comprises: adjusting a sensitivity setting until an event sensing threshold is determined; and determining the sinus event sensing thresholds on a rhythmic consistency of occurrence of the sensed events. 3. The method according to claim 2, wherein the step of automatically determining sensing thresholds of the ectopic events comprises determining the ectopic event sensing thresholds on any one of: a lack of rhythmic consistency of occurrence of sensed events, and a rhythmic consistency that is shorter in time than sinus event sensing. 4. The method according to claim 3, wherein the step of adjusting a sensitivity setting comprises maintaining a sensitivity setting of a first sense amplifier, while adjusting a sensitivity setting of a second sense amplifier. 5. The method according to claim 3, wherein the step of sensing sinus events comprises sensing any of: atrial events; or ventricular events. 6. The method according to claim 5, wherein the step of sensing ectopic events comprises sensing any of: premature atrial contractions (PACs), or premature ventricular contractions (PVCs). 7. The method according to claim 6, further including the step of determining an event interval by sampling at least two consecutive cardiac cycles to determine a time interval between two sensed events. 8. The method according to claim 7, further including the step of comparing the event interval to an average sinus interval measured between previously sensed sinus events, and, if the event interval is approximately equal to the average sinus interval, then confirming sinus event detection and sinus sensing threshold. 9. The method according to claim 8, further including the step of adjusting a sensitivity setting based on a sensing threshold determined for the detected sinus event. 10. The method according to claim 8, further including the step of comparing the event interval to the average sinus interval, and, if the event interval is a predetermine... | Pacesetter, Inc. | A61B005/0452
US07302288 | B1 | 2007-11-27 | Tool position indicator | 1996-11-25 | A tool for use with a surgical navigation system includes a plurality of LEDs mounted to the distal end of the tool. The LED's provide an indication of the direction in which the tool must be moved to reach a desired position. Four LED's mounted along perpendicular axes are provided. Modes for indicating the desired location, orientation, rotation, and depth are provided. A mode indicator may also be provided. The tool is designed for use with a surgical navigation system which includes an infrared localizer and computer system. The indicators assist the surgeon in positioning the surgical tool in relation to the patient. Indicators other than LEDs may be used; the indicators may also be mounted in alternate locations in the operating room environment. | 1. An apparatus for guiding the movement of a surgical toot in relation to the anatomy of a patient, the apparatus comprising: a surgical tool;a display remote from the tool for displaying a diagnostic image of the anatomy;a means for superimposing at least one of a proposed target and a proposed trajectory indicative of a desired position of the tool on the display;means for determining an actual position of the tool;means for determining a difference between the actual position of the tool and the desired position of the tool; and,indicating means different from the remote display and mounted on the tool for indicating to a human the difference between the actual and desired positions of the tool, the indicating means having an indicator reference frame different from the remote display frame of reference, the difference being indicated with respect to the indicator reference frame. | 1. An apparatus for guiding the movement of a surgical toot in relation to the anatomy of a patient, the apparatus comprising: a surgical tool;a display remote from the tool for displaying a diagnostic image of the anatomy;a means for superimposing at least one of a proposed target and a proposed trajectory indicative of a desired position of the tool on the display;means for determining an actual position of the tool;means for determining a difference between the actual position of the tool and the desired position of the tool; and,indicating means different from the remote display and mounted on the tool for indicating to a human the difference between the actual and desired positions of the tool, the indicating means having an indicator reference frame different from the remote display frame of reference, the difference being indicated with respect to the indicator reference frame. 2. The apparatus of claim 1 wherein the indicating means comprises at least one indicator. 3. The apparatus of claim 1 wherein the means for deter the actual position of the tool comprises one of an infrared localizer and an articulated arm. 4. The apparatus of claim 1 wherein the actual and desired positions include at least one of a location, trajectory, depth, and rotation of the tool. 5. The apparatus of claim 1 wherein the indicating means comprises a position indicator and a mode indicator. 6. An apparatus for guiding the movement of a surgical tool in relation to the anatomy of a patient, the apparatus comprising: a surgical tool, wherein the surgical tool includes a pointing axis;a position determining means for determining an actual position and orientation of the pointing axis the surgical tool;a difference determining means for determining a difference between the actual position and orientation of the pointing axis of the surgical tool and a desired position and orientation of the pointing axis of the surgical tool;at least two electronic indicators mounted to the surgical tool in a plane substantially orthogonal to the pointing axis, the indicators being connected with the difference determining means to indicate the determined difference between the actual and desired position and orientation of the pointing axis, in a frame of reference of the surgical tool; anda display remote from the electronic indicators which displays a diagnostic image with the actual position and orientation of the pointing axis of the surgical tool and the desired position and orientation of the pointing axis of the surgical tool superimposed thereon. 7. The apparatus of claim 6, further including at least two pairs of electronic indicators. 8. An apparatus for use with an image guided surgery system, the apparatus comprising: a surgical tool;first and second display members associated with the tool and arranged along a first line and third and fourth display members associated with the tool and arranged along a second line, the first and second lines being perpendicula... | Z-Kat, Inc. Biomet Manufacturing Corp. | A61B005/00
US07163511 | B2 | 2007-01-16 | Devices and methods for frequent measurement of an analyte present in a biological system | 2003-01-29 | Devices and methods are provided for frequently measuring the concentration of an analyte present in a biological system. A monitoring system having at least two components is employed in order to allow separation of data collection from data processing and display. Such separation allows greater flexibility and convenience for the user. | 1. A monitoring system for frequently measuring an analyte present in a subject and providing to said subject self-monitoring of said analyte, said monitoring system comprising: (A) a first component to be placed in contact with a skin or mucosal surface of said subject comprising(i) a sampling mechanism for transdermally extracting the analyte from the subject, wherein said sampling mechanism is adapted for extracting the analyte across a skin or mucosal surface of said subject;(ii) a sensing mechanism in operative contact with the analyte extracted by the sampling mechanism, wherein said sensing mechanism obtains a signal from the extracted analyte and said signal is specifically related to an amount or concentration of analyte; and(iii) a first mechanism for providing operative communication with a second component of the monitoring system, wherein (a) said operative communication comprises wireless communication technology employing electromagnetic waves, and (b) said first and second components are separately housed; and(B) the second component comprising(i) a user interface;(ii) a second mechanism for providing operative communication with the first component, wherein (a) said operative communication comprises wireless communication technology employing electromagnetic waves, (b) the second component receives the signal from the first component, and (c) is capable of communicating with a third component remote from the subject being monitored without communicating with the first component;(iii) a computing mechanism that converts the signal from the extracted analyte to an output indicative of the amount or concentration of analyte extracted by the sampling mechanism;(iv) a visual display said output to the subject;(v) an alarm that produces tactile and/or auditory signals to warn the subject when the amount or concentration of analyte is above a pre-set threshold value or below a pre-set threshold value; and(vi) the second component configured to be worn by the subject on a wrist of the subject or a belt of the subject, or configured to be carried by the subject as a credit-card-type unit or a pager-type unit. | 1. A monitoring system for frequently measuring an analyte present in a subject and providing to said subject self-monitoring of said analyte, said monitoring system comprising: (A) a first component to be placed in contact with a skin or mucosal surface of said subject comprising(i) a sampling mechanism for transdermally extracting the analyte from the subject, wherein said sampling mechanism is adapted for extracting the analyte across a skin or mucosal surface of said subject;(ii) a sensing mechanism in operative contact with the analyte extracted by the sampling mechanism, wherein said sensing mechanism obtains a signal from the extracted analyte and said signal is specifically related to an amount or concentration of analyte; and(iii) a first mechanism for providing operative communication with a second component of the monitoring system, wherein (a) said operative communication comprises wireless communication technology employing electromagnetic waves, and (b) said first and second components are separately housed; and(B) the second component comprising(i) a user interface;(ii) a second mechanism for providing operative communication with the first component, wherein (a) said operative communication comprises wireless communication technology employing electromagnetic waves, (b) the second component receives the signal from the first component, and (c) is capable of communicating with a third component remote from the subject being monitored without communicating with the first component;(iii) a computing mechanism that converts the signal from the extracted analyte to an output indicative of the amount or concentration of analyte extracted by the sampling mechanism;(iv) a visual display said output to the subject;(v) an alarm that produces tactile and/or auditory signals to warn the subject when the amount or concentration of analyte is above a pre-set threshold value or below a pre-set threshold value; and(vi) the second component configured to be worn by the subject on a wrist of the subject or a belt of the subject, or configured to be carried by the subject as a credit-card-type unit or a pager-type unit. 2. The monitoring system of claim 1, wherein the wireless communication technology employs low frequency electromagnetic waves in a frequency range of about 1 Hz. to about 1 Mega Hz. 3. The monitoring system of claim 1, wherein the wireless communication technology employs medium frequency electromagnetic waves in a frequency range of about 1 Mega Hz. to about 500 Mega Hz. 4. The monitoring system of claim 1, wherein the wireless communication technology employs high frequency electromagnetic waves in a frequency range of about 500 Mega Hz. to about 5 Giga Hz. 5. The monitoring system of claim 1, wherein the wireless communication technology employs capacitance coupling between the subject and subject's environment. 6. The monitoring system of claim 1, wherein the wireless communication technology employs inductive coupling. 7.... | Animas Technologies, LLC | A61B005/00
US07197350 | B2 | 2007-03-27 | Device for determining acoustically evoked brainstem potentials | 2002-03-07 | The invention relates to a device for determining acoustically evoked brain potentials in objective audiometry from electrodes applied to the head of a subject, where the device comprises a plurality of electrodes, i.e. at least one pickup electrode and a reference electrode, to be applied at different points of the head, where the electrodes are mounted on a structure in fixed positions in relation to each other, where the structure comprises at least one flexible cushion abutting the head of the subject around an ear, where the pickup electrode and the reference electrode are integrated in the flexible cushion. | 1. A device for determining acoustically evoked brain potentials in brainstem audiometry, the device comprising: a plurality of electrodes comprising at least one pickup electrode and a reference electrode, to be applied at different points on a subject's head, a structure mounting the electrodes in fixed positions in relation to each other, said structure comprising a flexible cushion for abutting the head of the subject around an ear, and wherein the pickup electrode and the reference electrode are integrated in the flexible cushion. | 1. A device for determining acoustically evoked brain potentials in brainstem audiometry, the device comprising: a plurality of electrodes comprising at least one pickup electrode and a reference electrode, to be applied at different points on a subject's head, a structure mounting the electrodes in fixed positions in relation to each other, said structure comprising a flexible cushion for abutting the head of the subject around an ear, and wherein the pickup electrode and the reference electrode are integrated in the flexible cushion. 2. A device according to claim 1, where the cushion forms an ear bud on a headphone comprising a output transducer for delivering and output signal to the ear. 3. A device according to claim 1, where two cushions are provided, where the structure comprises two interconnected parts each comprising a cushion and where the two interconnected parts are connected by means of a headband. 4. A device according to claim 1, where the structure comprises a headband, which at an area at its end opposite the structure is adapted for abutting the subject's head opposite the ear presently being tested hereby supporting the structure in a stable position. 5. A device according to claim 4, where a further electrode is mounted in connection with the headband. 6. A device according to claim 1, wherein each cushion is releasably connected to the remaining structure by means of a conductive coupling. 7. A device for determining acoustically evoked brain potentials in brainstem audiometry, the device comprising a plurality of electrodes comprising at least one pickup electrode and a reference electrode for application against different points of a subject's head, and a structure for mounting the electrodes in fixed positions in relation to each other, the structure comprising a flexible cushion for abutting the head of the subject around an ear, the electrodes being flat and mounted on surface parts of the flexible cushion so as to abut the head of a subject. | Maico Diagnostic GmbH | A61B005/0478
US07087372 | B2 | 2006-08-08 | Compositions and methods for detection of | 2001-01-18 | The invention provides methods and compositions for the detection of Ehrlichia canis and Ehrlichia chaffeensis antibodies and antibody fragments. | 1. A device containing one or more polypeptides consisting of SEQ ID NO:1, SEQ ID NO:2, SEQ ID NO:3, SEQ ID NO:4, SEQ ID NO:5, SEQ ID NO:6, SEQ ID NO:7, and amino acid substitution variants thereof that specifically bind to an anti-Ehrlichia antibody. | 1. A device containing one or more polypeptides consisting of SEQ ID NO:1, SEQ ID NO:2, SEQ ID NO:3, SEQ ID NO:4, SEQ ID NO:5, SEQ ID NO:6, SEQ ID NO:7, and amino acid substitution variants thereof that specifically bind to an anti-Ehrlichia antibody. 2. The device of claim 1, further comprising instructions for use of the one or more polypeptides for the identification of an Ehrlichia infection in a mammal. 3. The device of claim 2, wherein the instructions for use indicate that the identification of an Ehrlichia infection is done using a method of detecting presence of antibodies to Ehrlichia comprising: (a) contacting one or more polypeptides selected from the group consisting of the polypeptides shown in SEQ ID NO:1, SEQ ID NO:2, SEQ ID NO:3, SEQ ID NO:4, SEQ ID NO:5, SEQ ID NO:6, SEQ ID NO:7, and amino acid substitution variants thereof that specifically bind to an anti-Ehrlichia antibody, with a test sample suspected of comprising antibodies to Ehrlichia, under conditions that allow polypeptide/antibody complexes to form;(b) detecting polypeptide/antibody complexes;wherein the detection of polypeptide/antibody complexes is an indication that an Ehrlichia infection is present. 4. The device of claim 2, wherein the Ehrlichia infection is caused by Ehrlichia canis or Ehrlichia chaffeensis. 5. A device containing one or more polypeptides selected from the group consisting of SEQ ID NO:3, SEQ ID NO:4, SEQ ID NO:5, SEQ ID NO:6, SEQ ID NO:7, that specifically bind to an anti-Ehrlichia antibody. 6. The device of claim 5, further comprising instructions for use of the one or more polypeptides for the identification of an Ehrlichia infection in a mammal. 7. The device of claim 5, wherein the instructions for use indicate that the identification of an Ehrlichia infection is done using a method of detecting presence of antibodies to Ehrlichia comprising: (a) contacting one or more polypeptides selected from the group consisting of SEQ ID NO:3, SEQ ID NO:4, SEQ ID NO:5, SEQ ID NO:6, SEQ ID NO:7, that specifically bind to an anti-Ehrlichia antibody, with a test sample suspected of comprising antibodies to Ehrlichia, under conditions that allow polypeptide/antibody complexes to form;(b) detecting polypeptide/antibody complexes;wherein the detection of polypeptide/antibody complexes is an indication that an Ehrlichia infection is present. 8. The device of claim 5, wherein the Ehrlichia infection is caused by Ehrlichia canis or Ehrlichia chaffeensis. | IDEXX Laboratories, Inc. | C12Q001/00 C12N011/08 A61B005/055 A61K039/00 A61K039/02 A61K038/00
US07182734 | B2 | 2007-02-27 | Retracting needle safety device | 2002-08-08 | A retractable safety needle assembly is provided. The needle assembly includes a tubular outer body, a retraction assembly contained with the outer body and a biasing element for applying a retraction force therebetween. The retraction assembly includes a cannula having an intravenous puncture tip and a non-patient puncture tip at opposing ends. In a first position, the intravenous puncture tip extends from a first end of the tubular outer body. The retraction assembly is adapted for movement within the tubular outer body from the first position to a second position in which the intravenous puncture tip and the non-patient puncture tip are contained entirely within the tubular outer body and are offset from a general axis defining the outer body. | 1. A retractable safety needle assembly comprising: a) a generally tubular outer body including a first end, a second end and a tubular wall extending along an axis from said first end to said second end to define an interior opening within said outer body, said outer body including a guide channel at least a portion of which is offset from said axis;b) a retraction assembly positioned within said interior opening of said outer body, said retraction assembly including a cannula having an intravenous puncture tip and a non-patient puncture tip and including a fin for corresponding engagement with said guide channel of said outer body, said retraction assembly adapted for movement within said interior opening of said outer body from a first position in which said intravenous puncture tip extends from said first end of said outer body to a second position in which said intravenous puncture tip and said non-patient puncture tip are contained entirely within said outer body and said cannula is offset from said axis of said outer body;c) a retaining element for preventing movement of said retraction assembly between said first position and said second position;d) a biasing element for applying a retraction force between said retraction assembly and said outer body to move said retraction assembly from said first position to said second position upon release of said retaining element; ande) a lock mechanism for locking said retraction assembly in said second position, wherein said guide channel includes a notch,wherein said notch forms a notch pocket located at an end of the guide channel, said guide channel end being located at said portion that is offset from said axis, and wherein said notch acts as the lock mechanism for preventing said fin from moving within said guide channel and thereby preventing said retraction assembly from moving from said second position to said first position. | 1. A retractable safety needle assembly comprising: a) a generally tubular outer body including a first end, a second end and a tubular wall extending along an axis from said first end to said second end to define an interior opening within said outer body, said outer body including a guide channel at least a portion of which is offset from said axis;b) a retraction assembly positioned within said interior opening of said outer body, said retraction assembly including a cannula having an intravenous puncture tip and a non-patient puncture tip and including a fin for corresponding engagement with said guide channel of said outer body, said retraction assembly adapted for movement within said interior opening of said outer body from a first position in which said intravenous puncture tip extends from said first end of said outer body to a second position in which said intravenous puncture tip and said non-patient puncture tip are contained entirely within said outer body and said cannula is offset from said axis of said outer body;c) a retaining element for preventing movement of said retraction assembly between said first position and said second position;d) a biasing element for applying a retraction force between said retraction assembly and said outer body to move said retraction assembly from said first position to said second position upon release of said retaining element; ande) a lock mechanism for locking said retraction assembly in said second position, wherein said guide channel includes a notch,wherein said notch forms a notch pocket located at an end of the guide channel, said guide channel end being located at said portion that is offset from said axis, and wherein said notch acts as the lock mechanism for preventing said fin from moving within said guide channel and thereby preventing said retraction assembly from moving from said second position to said first position. 2. The needle assembly as in claim 1, wherein said biasing element comprises a compression spring. 3. The needle assembly as in claim 1, wherein said tubular wall of said outer body includes a recess extending within the interior opening of said outer body adapted for accommodating said non-patient puncture tip when said retraction assembly is in said second position. 4. The needle assembly as in claim 1, wherein said outer body includes an opening extending through said tubular wall and said retraction assembly includes a tab extending through said opening, said retaining element comprising an interference engagement between said tab and said tubular wall. 5. The needle assembly as in claim 1, wherein said retraction assembly is irreversibly movable from said first position to said second position. 6. The needle assembly as in claim 1, wherein attempted movement of said retraction assembly from said second position to said first position causes said intravenous puncture tip to contact an interior portion of said tubular outer body at said first end, thereby prev... | Becton, Dickinson and Company | A61B005/00
US07034113 | B2 | 2006-04-25 | Bacteriocin-metal complexes in the detection of pathogens and other biological analytes | 2002-02-22 | Complexes of bacteriocins and metals are provided that are useful in detecting bacteria, fungi and other biological analytes, and are particularly useful in detecting gram positive bacteria. The complexes are preferably chelated complexes wherein the bacteriocin is a lantibiotic, non-lanthionine containing peptide, large heat labile protein and complex bacteriocin, fusion protein thereof, mixture thereof, and fragment, homolog and variant thereof, and (b) a detectable label comprising a transition or lanthanide metal. The complex preferentially binds to viable gram positive or mycobacterial cells. The complex can also bind to gram negative bacteria and fungi. Methods of using the complexes in assays, diagnosis and imaging are also provided. | 1. An isolated chelated complex comprising: nisin, fusion proteins thereof, and mixtures thereof; anda detectable label comprising cobalt chelated with the nisin. | 1. An isolated chelated complex comprising: nisin, fusion proteins thereof, and mixtures thereof; anda detectable label comprising cobalt chelated with the nisin. 2. The complex of claim 1, wherein the complex binds to microbial cells selected from the group consisting of gram positive bacteria or mycobacteria. 3. The complex of claim 1, wherein the complex binds to gram negative bacteria or fungi. 4. A method for conducting a chemiluminescent assay of pathogens comprising: (a) contacting a sample with the complex of claim 1, under conditions and for a time sufficient to allow the complex to bind with pathogens present in the sample;(b) removing any unbound complex; and(c) detecting the pachogens by contacting the sample with a peroxide source and an oxidizable substrae. 5. The method of claim 4, wherein pathogens are isolated from the sample prior to contacting the sample with the complex. 6. The method of claim 5, wherein pathogens are isolated from the sample using antibody-attached microbeads or phage-attached microbeads. 7. The method of claim 6, wherein the microbeads comprise a magnetic matenal. 8. The method of claim 4, wherein the peroxide source is selected from the group consisting of hydrogen peroxide, benzoyl peroxide and cumyl peroxide. 9. The method of claim 4, wherein the oxidizable substrate is selected from the group consisting of luininol and its derivatives, lucigerain, penicillin, luciferin, and polyaromatic phthalylhydrazides. 10. The method of claim 4, wherein the pathogens are grain positive bacteria or mycobacteria. 11. The method of claim 4, wherein the patbogeas are gram negative bacteria or fungi. 12. The method of claim 4, wherein the bacteria are selected from the group consisting of lactococci, leuconostocs, micrococci, pediococci, actisomyces, mycoabacteria, pneuniococci, streptococci, staphylococci, aerobic bacilli, anaerobic clostridia, listeria and nocardia. 13. The method of claim 12, wherein the mycobacteria are selected from the group consisting of mycobacteriurn tuberculosis, inycobacterium avium, mnycobacterium paratuberculosis, mycobacterium bovis and mycobacterium leprae. 14. The method of claim 12, wherein the bacteria are selected from the group consisting of Bacillus ntliracis, Clostridinin botulinum and Clostridium perfringes. | Paradigm Diagnostics, LLC | C07K014/00 A61B005/00 C12Q001/00
US07149756 | B1 | 2006-12-12 | System and method for determining the probable existence of disease | 2000-10-16 | A computer implemented and interactive system for obtaining a patient's symptoms and other indications of disease, which are utilized to determine the probable existence of disease. The program generates a series of inquiries which are displayed to a patient who is asked to identify the body part which the patient believes is causing a physical or mental problem. The system's program uses a database of diseases and related symptoms to generate questions. Based upon the patient's answer to each question, the program determines the probability that the patient has one of the diseases contained in the database. When the patient has completed answering all of the questions, the system's program generates a report of the probable existence diseases, including the patient's most probable disease. | 1. A method for determining the probable existence of disease comprising: a) providing a plurality of patient Body Parts;b) associating with each Body Part a list of diseases, and as to each disease in the list of diseases, associating a list of Symptoms which indicate said disease;c) associating with each disease in each list of diseases a disease numerical value equal to the number of persons who have been confirmed to exhibit said each disease;d) associating with each Symptom in each list of Symptoms, which indicate a specific disease, a symptom numerical value equal to the disease numerical value of said specific disease;e) calculating an initial probable existence of each disease in each list of diseases by determining the quotient of the disease numerical value of each disease divided by the sum of all disease numerical values for all diseases;f) calculating a total symptom numerical value for all identical Symptoms in all lists of Symptoms;g) identifying a set of Symptoms, and each Symptom's associated total symptom numerical value, which comprise all Symptoms associated to each Body Part;h) prompting a patient to identify a target Body Part which the patient associates to his or her physical or mental condition;i) identifying a target Symptom from the set of Symptoms associated to the target Body Part which has the highest total symptom numerical value;j) generating a question presented to the patient which asks the patient if the target Symptom is indicated or is not indicated;k) prompting the patient to respond to the question;l) updating the probable existence of each disease in each list of diseases based upon the patient's answer in accordance with the following formula:Pr⁢{D⁢S}=Pr⁢{D}⁢Pr⁢{S⁢D}Pr⁢{S}where D is a disease and S is a Symptom;m) identifying a qualifying disease in each list of diseases with a probability greater than a predetermined threshold value, and adding each Symptom associated to said qualifying disease to the set of Symptoms associated to the target Body Part if the Symptom to be added is not already in the set;n) deleting the target Symptom from the set of Symptoms associated to the target Body Part;o) repeating steps i through n above until each Symptom contained in the set of Symptoms associated to the Body Part is deleted; andp) generating a list of diseases and the probable existence of each disease in the list, including the identity of a most probable disease having the highest probability of existence. | 1. A method for determining the probable existence of disease comprising: a) providing a plurality of patient Body Parts;b) associating with each Body Part a list of diseases, and as to each disease in the list of diseases, associating a list of Symptoms which indicate said disease;c) associating with each disease in each list of diseases a disease numerical value equal to the number of persons who have been confirmed to exhibit said each disease;d) associating with each Symptom in each list of Symptoms, which indicate a specific disease, a symptom numerical value equal to the disease numerical value of said specific disease;e) calculating an initial probable existence of each disease in each list of diseases by determining the quotient of the disease numerical value of each disease divided by the sum of all disease numerical values for all diseases;f) calculating a total symptom numerical value for all identical Symptoms in all lists of Symptoms;g) identifying a set of Symptoms, and each Symptom's associated total symptom numerical value, which comprise all Symptoms associated to each Body Part;h) prompting a patient to identify a target Body Part which the patient associates to his or her physical or mental condition;i) identifying a target Symptom from the set of Symptoms associated to the target Body Part which has the highest total symptom numerical value;j) generating a question presented to the patient which asks the patient if the target Symptom is indicated or is not indicated;k) prompting the patient to respond to the question;l) updating the probable existence of each disease in each list of diseases based upon the patient's answer in accordance with the following formula:Pr⁢{D⁢S}=Pr⁢{D}⁢Pr⁢{S⁢D}Pr⁢{S}where D is a disease and S is a Symptom;m) identifying a qualifying disease in each list of diseases with a probability greater than a predetermined threshold value, and adding each Symptom associated to said qualifying disease to the set of Symptoms associated to the target Body Part if the Symptom to be added is not already in the set;n) deleting the target Symptom from the set of Symptoms associated to the target Body Part;o) repeating steps i through n above until each Symptom contained in the set of Symptoms associated to the Body Part is deleted; andp) generating a list of diseases and the probable existence of each disease in the list, including the identity of a most probable disease having the highest probability of existence. 2. The method as in claim 1 in which the step of generating a question further includes assembling the question from a list of possible question phrases and a template code associated to each question phrase, whereby the syntax of the sentence is determined by the Body Part identified by the patient and the Symptom with the highest total symptom numerical value. 3. A method for determining the probable existence of disease comprising: a) providing a plurality of patient Body Parts;b) associating with ea... | Medoctor, Inc. | G06Q010/00 A61B005/00
US07072710 | B2 | 2006-07-04 | Method and a device for determining the dry weight of a patient with kidney failure | 2001-11-06 | A system and method for determining the dry weight wgtdry(t) of a patient at a time t by determining the extracellular water volume ECV(t) and the weight Wgt(t) of the patient at time t and by deriving the dry weight wgtdry(t) of the patient from an intersection of a function derived from the determined ECV(t) and Wgt(t) values with a previously established extracellular water volume (ECV) against dry weight (wgtdry(t)) reference relation representing healthy subjects. To obtain more accurate results it is also proposed to take into account a compartmental mass correction Δm(t). The invention also relates to a device for deriving the dry weight wgtdry(t). | 1. A method for determining the dry weight Wgtdry(t) of a patient at a time t comprising the steps of: determining the extracellular water volume ECV(t) of the patient at the time t,determining the weight Wgt(t) of the patient at the time t,deriving the dry weight Wgtdry(t) of the patient from an intersection of a function derived from the determined ECV(t) and Wgt(t) values with a previously established extracellular water volume (ECV) against dry weight (Wgtdry) reference relation representing healthy subjects. | 1. A method for determining the dry weight Wgtdry(t) of a patient at a time t comprising the steps of: determining the extracellular water volume ECV(t) of the patient at the time t,determining the weight Wgt(t) of the patient at the time t,deriving the dry weight Wgtdry(t) of the patient from an intersection of a function derived from the determined ECV(t) and Wgt(t) values with a previously established extracellular water volume (ECV) against dry weight (Wgtdry) reference relation representing healthy subjects. 2. The method according to claim 1 characterized in that ECV(t) is derived from a bioimpedance measurement. 3. The method according to claim 2 characterized in that the bioimpedance measurement is a whole body measurement. 4. The method according to claim 2 characterized in that the bioimpedance measurement is a segmental measurement. 5. The method according to claim 1 characterized in that Wgtdry(t) is determined using the following expression:WgtDry⁡(t)=ECV⁡(t)-βe·Wgt⁡(t)Ae-βe,where αe and βe are empirically determined coefficients. 6. The method according to claim 1 characterized in that the ECV(ti) and Wqt(ti) values of a patient at times tj, i=1 . . . j are stored and that the dry body weight Wgtdry(ti) is derived by a linear regression analysis. 7. The method according claim 1 characterized in that a compartmental mass correction Δm(t) is determined in order to derive the dry body weight Wgtdry(t) from the determined weight Wgt(t). 8. The method according to claim 7 characterized in that the dry weight Wgtdry(t) is derived by the following expression:WgtDry⁡(t)=ECV⁡(t)-βe·Wgt⁡(t)αe-βe-αe·Δ⁢⁢m⁡(t)βe-αe,where αe and βe are empirically determined coefficients. 9. The method according to claim 7 characterized in that the compartmental mass correction Δm(t) encompasses a fat mass correction Δf(t) and/or a muscle mass correction Δmmuscle(t). 10. The method according to claim 7 characterized in that the intracellular water volume ICV(t) is determined for the patient at the time t and that the determined ICV(t) is used to derive the compartmental mass correction Δm(t). 11. The method according to claim 1 characterized in that the method for determining the dry weight Wgtdry(t) of a patient at a time t further comprises the steps of determining the intracellular water volume ICV(t) of the patient at the time t and of deriving the dry weight Wgtdry(t) of the patient also from an intersection of a function derived from the determined ICV(t) and Wgt(t) values with a previously established intracellular water volume (ICV) against Wgtdry reference relation representing healthy subjects. 12. A device (10) for carrying out the method according to claim 1 comprising a microprocesaor unit (1) which in turn comprises a microprocessor program storage unit (1a),a input unit (2) to enable entering the values of EVC(t) and Wgt(t),a computer storage unit (3) for storing the ECV(t) and Wgt(t) value. 13. The device according to claim 12 characte... | Fresenius Medical Care Deutschland GmbH | A61B005/05
US07058439 | B2 | 2006-06-06 | Methods of forming prostheses | 2002-05-03 | A breast prosthesis may be formed by forming an outer layer on a mold, and filling the mold with a polymerizable foaming composition. The mold may be formed by forming a computer model of the prosthesis based on scanning a patient. The computer model may be used to form a solid model. A prosthesis may be coupled to the patient by coupling at least one metallic insert to the prosthesis and at least one magnet to the patient. Alternately, the prosthesis may be formed with a retaining harness integral to the prosthesis. In still another method, the prosthesis may be coupled to a retaining device surgically implanted in the patient. | 1. A method of forming a model of a breast prosthesis for a patient who has had one breast at least partially removed, comprising: providing a scanning system comprising one or more imaging devices and one or more alignment markers, wherein at least one of the alignment markers comprises a tangible object, wherein the tangible object is noticeable in a scanned image and may be used to manipulate a scanned image as a reference point;arranging at least one of the alignment markers in the middle of the body of the patient and within the field of view of at least one of the imaging devices;determining a first set of data elements using the scanning system is based on a shape of the patient's intact breast;determining a first computer model of the intact breast based on the first set of data elements;using at least one of the alignment markers in the middle of the body of the patient as a fixed reference point relative to the patient to determine alignment of the first computer model of the intact breast;applying a reflection transformation to the first computer model of the intact breast to form a second computer model, the second computer model comprising a second set of data points that represents a mirror image of the first computer model;determining a third set of data elements using the scanning system based on a surgical site at which the breast was at least partially removed;determining a third computer model based on the third set of data elements; and combining the third computer model and the second computer model to form a breast prosthesis model, wherein the second computer model defines an anterior portion of the breast prosthesis model that mirrors the intact breast and the third computer model defines a posterior portion of the breast prosthesis model that is custom fit to the surgical site. | 1. A method of forming a model of a breast prosthesis for a patient who has had one breast at least partially removed, comprising: providing a scanning system comprising one or more imaging devices and one or more alignment markers, wherein at least one of the alignment markers comprises a tangible object, wherein the tangible object is noticeable in a scanned image and may be used to manipulate a scanned image as a reference point;arranging at least one of the alignment markers in the middle of the body of the patient and within the field of view of at least one of the imaging devices;determining a first set of data elements using the scanning system is based on a shape of the patient's intact breast;determining a first computer model of the intact breast based on the first set of data elements;using at least one of the alignment markers in the middle of the body of the patient as a fixed reference point relative to the patient to determine alignment of the first computer model of the intact breast;applying a reflection transformation to the first computer model of the intact breast to form a second computer model, the second computer model comprising a second set of data points that represents a mirror image of the first computer model;determining a third set of data elements using the scanning system based on a surgical site at which the breast was at least partially removed;determining a third computer model based on the third set of data elements; and combining the third computer model and the second computer model to form a breast prosthesis model, wherein the second computer model defines an anterior portion of the breast prosthesis model that mirrors the intact breast and the third computer model defines a posterior portion of the breast prosthesis model that is custom fit to the surgical site. 2. A method of forming a model of a breast prosthesis for a patient who has had one breast at least partially removed, comprising: providing a scanning system comprising one or more imaging devices and one or more alignment markers, wherein at least one of the alignment markers comprises a tangible object, wherein the tangible object is noticeable in a scanned image and may be used to manipulate a scanned image as a reference point;arranging at least one of the alignment markers in the middle of the body of the patient and within the field of view of at least one of the imaging devices;determining a first set of data elements using the scanning system based on a shape of a garment at least partially covering the patient's intact breast;determining a computer model of a breast prosthesis based on the first set of data elements; andusing at least one of the alignment markers in the middle of the body of the patient as a fixed reference point relative to the patient to determine alignment of the first computer model;applying a reflection transformation to the first computer model of the garment covering the intact breast to form a second computer mod... | ContourMed, Inc. | A61B005/00 A61F002/52
US07123955 | B1 | 2006-10-17 | Control method and system and the sense organs test method and system based on electrical steady state induced response | 2000-10-30 | A control method and a sense organ testing method based on steady state brain evoked response include the steps of: generating one or multiple stimulus signal(s) varying in various frequencies that can be sensed by sense organs of human subject; detecting the brain electrical signal of stimulated human subject; analyzing the detected brain electrical signal to extract the characteristic physical quantity corresponding to the frequency of a stimulus signal; and, according to the characteristic physical quantity, controlling a specific device corresponding to the stimulus signal to perform a preprogrammed operation or objectively evaluating the condition of the sense organ of the subject. | 1. A control method based on steady state brain evoked response comprising the steps of: a. generating one or multiple stimulus signals varying in various frequencies that can be sensed by sense organs of human subject;b. detecting the brain electrical signal of stimulated human subject;c. analyzing the detected brain electrical signal to extract a characteristic quantity corresponding to the frequency of one of the one or multiple stimulus signals as a control signal; andd. on the basis of said control signal, controlling a device corresponding to the stimulus signal to perform a preprogrammed operation. | 1. A control method based on steady state brain evoked response comprising the steps of: a. generating one or multiple stimulus signals varying in various frequencies that can be sensed by sense organs of human subject;b. detecting the brain electrical signal of stimulated human subject;c. analyzing the detected brain electrical signal to extract a characteristic quantity corresponding to the frequency of one of the one or multiple stimulus signals as a control signal; andd. on the basis of said control signal, controlling a device corresponding to the stimulus signal to perform a preprogrammed operation. 2. A control method according to claim 1, wherein said multiple signals varying in various frequencies correspond to multiple different devices to be controlled respectively. 3. A control method according to claim 1 or 2, wherein said one or multiple stimulus signals are visual stimulus signals, and said sense organ of human subject is a visual organ. 4. A control method according to claim 3, wherein at least one of said one or multiple visual stimulus signals is a light signal, the brightness, shape, color, size, position, duration or their combination of which varies in a frequency. 5. A control method according to claim 3, wherein each of said one or multiple visual stimulus signals has a frequency of 1–50 hertz. 6. A control method according to claim 5, wherein each of said one or multiple visual stimulus signals has a frequency of 4–25 hertz. 7. A control method according to claim 3, wherein said multiple visual stimulus signals are a group of spots scintillating in various frequencies on a screen, each spot corresponding to a device to be controlled respectively. 8. A control method according to claim 3, wherein at least one of said one or multiple visual stimulus signals is a combination of multiple light signals which overlap in space and scintillate in various frequencies, and said characteristic quantity is a composition of frequency components corresponding to the composition of the various frequencies of said combination of multi light signals in the detected brain electrical signal. 9. A control method according to claim 1 or 2, wherein said one or multiple stimulus signals are auditory stimulus signals, and said sense organ of human subject is an auditory organ. 10. A control method according to claim 9, wherein at least one of said one or multiple auditory stimulus signals is of a simple tone having a single varying frequency. 11. A control method according to claim 9, wherein at least one of said one or multiple auditory stimulus signal signals is of a composite tone consisting of a specific composition of certain frequencies. 12. A control method according to claim 1 or 2, said one or multiple stimulus signals are tactual stimulus signals, and said sense organ of human subject is a tactile organ. 13. A control method according to claim 1 or 2, wherein said analysis includes spectrum analyzing of the brain electrical ... | Tsinghua University | A61B005/0476 G09B021/00
US07047061 | B2 | 2006-05-16 | Method of localizing the myocardium of the heart and method of determining perfusion parameters thereof | 2001-12-05 | The invention relates to a method of deriving the myocardial region from one or more MR images. Left and right ventricle contours (1, 2) are determined. The thickness or width (D) of the myocardial region is computed from the distance between these contours in a predetermined part (A) of the image (images). The location of the myocardial region outside the preset region is derived on the basis of this distance. | 1. A method of determining a position of the myocardium of the heart from a magnetic resonance (MR) image or a series of MR images of the heart of a patient, comprising the steps of: first determining from the image or images a left ventricle contour and a right ventricle contour of the heart;second determining a thickness measure of the myocardium based on a detected distance between said contours in a predetermined region in which both contours of the left and right ventricles extend such that outside said region the position of the myocardium is bounded by said left ventricle contour and an outer contour that extends outside said left ventricle contour at a distance that is equal to the thickness measure of the myocardium. | 1. A method of determining a position of the myocardium of the heart from a magnetic resonance (MR) image or a series of MR images of the heart of a patient, comprising the steps of: first determining from the image or images a left ventricle contour and a right ventricle contour of the heart;second determining a thickness measure of the myocardium based on a detected distance between said contours in a predetermined region in which both contours of the left and right ventricles extend such that outside said region the position of the myocardium is bounded by said left ventricle contour and an outer contour that extends outside said left ventricle contour at a distance that is equal to the thickness measure of the myocardium. 2. A method as set forth in claim 1, wherein the step of second determining includes that the thickness measure of the myocardium is determined by selection of the largest distance, or the mean distance, between said left ventricle contour and said right ventricle contour in the predetermined region. 3. A method as set forth in claim 1, wherein the step of second determining includes that the predetermined region in which said left ventricle and right ventricle contours extend is bounded by two tangent lines disposed at distinct locations which partly coincide with, or are tangent to, said right ventricle contour and which intersect one another at a center of gravity of the left ventricle. 4. A method as set forth in claim 3, wherein the step of second determining includes that at least two lines of intersection that segment the myocardium are formed by the tangent lines that bound the region in which the left ventricle contour and the right ventricle contour extend at some distance from one another. 5. A magnetic resonance (MR) method for determining perfusion parameters of the heart of a patient while administering a contrast liquid, comprising the steps of: first calculating a position of the myocardium of the heart from a magnetic resonance (MR) image or a series of MR images of the heart of a patient, which step of calculating further comprises the steps of: first determining from the image or images a left ventricle contour and a right ventricle contour of the heart,second determining a thickness measure of the myocardium based on a detected distance between said contours in a predetermined region in which both contours of the left and right ventricles extend such that outside the region the position of the myocardium is bound by the left ventricle contour and an outer contour that extends outside the left ventricle contour at a distance that is equal to the thickness measure of the myocardium;segmenting the myocardium;second calculating a time intensity curve that represents the variation in time of the concentration of the contrast liquid in the myocardium for at least one segment; andderiving the perfusion parameters from the time intensity curve. 6. A method as set forth in claim 5, wherein the step of second ... | Koninklijke Philips Electronics N.V. | A61B005/055
US07266405 | B1 | 2007-09-04 | Compact electrode assembly for a portable ECG signaling device | 2000-08-23 | An electrode assembly (10) for a portable ECG signaling device, comprises a thin, flexible electrode support (11) supporting thereon a plurality of electrodes in spaced relationship. The electrodes (V1, V2, V3, V4, V5 and V6, LA, RA, LL) are configured for producing a twelve-lead electrocardiogram, and the electrode assembly is foldable into a compact assembly when not in use. | 1. An electrode assembly for a portable 12-lead ECG signaling device, said electrode assembly comprising a thin, flexible electrode support supporting a plurality of electrodes (V1, V2, V3, V4, V5 and V6, LA, RA, LL) at least some of which are constructed on the electrode support in proper spaced relationship for producing electrical contact with respective areas of a patient's chest for producing an electrocardiogram when the electrode assembly is placed directly against the patient's chest; wherein: the flexible support comprises a plurality of foldable sections that fixedly support the electrodes thereon and open out to form a substantially flat base that is placeable against the patient's chest so that those of said electrodes that are in proper spaced relationship for producing electrical contact with respective areas of a patient's chest simultaneously contact the respective areas of the patient's chest without requiring adjustment or calibration, and whereby the electrode assembly can be folded into a compact unit prior to or after use, andone of the foldable sections is provided with a flap for tucking into a slot in another one of said sections, whereby the electrode assembly can be folded into a self-contained compact unit prior to use. | 1. An electrode assembly for a portable 12-lead ECG signaling device, said electrode assembly comprising a thin, flexible electrode support supporting a plurality of electrodes (V1, V2, V3, V4, V5 and V6, LA, RA, LL) at least some of which are constructed on the electrode support in proper spaced relationship for producing electrical contact with respective areas of a patient's chest for producing an electrocardiogram when the electrode assembly is placed directly against the patient's chest; wherein: the flexible support comprises a plurality of foldable sections that fixedly support the electrodes thereon and open out to form a substantially flat base that is placeable against the patient's chest so that those of said electrodes that are in proper spaced relationship for producing electrical contact with respective areas of a patient's chest simultaneously contact the respective areas of the patient's chest without requiring adjustment or calibration, and whereby the electrode assembly can be folded into a compact unit prior to or after use, andone of the foldable sections is provided with a flap for tucking into a slot in another one of said sections, whereby the electrode assembly can be folded into a self-contained compact unit prior to use. 2. An electrode assembly for a portable 12-lead ECG signaling device, said electrode assembly comprising a thin, flexible electrode support supporting a plurality of electrodes (V1, V2, V3, V4, V5 and V6, LA, RA, LL) at least some of which are constructed on the electrode support in proper spaced relationship for producing electrical contact with respective areas of a patient's chest for producing an electrocardiogram when the electrode assembly is placed directly against the patient's chest; wherein: the flexible support comprises a plurality of foldable sections that fixedly support the electrodes thereon and open out to form a substantially flat base that is placeable against the patient's chest so that those of said electrodes that are in proper spaced relationship for producing electrical contact with respective areas of a patient's chest simultaneously contact the respective areas of the patient's chest without requiring adjustment or calibration, and whereby the electrode assembly can be folded into a compact unit prior to or after use, andthere is joined to at least one of the foldable sections a serpentine strip supporting thereon one of said electrodes (RA). 3. An electrode assembly for a portable 12-lead ECG signaling device, comprising: a thin, flexible electrode support that is foldable into a compact unit prior to or after use, said electrode support supporting at least six electrodes (V1, V2, V3, V4, V5 and V6) wholly constructed on the electrode support in proper mutual spaced relationship for producing electrical contact each with a correct respective area of a patient's chest when two of said electrodes (V1, V2) are substantially symmetrically disposed about his or her vertebrae for p... | SHL Telemedicine International Ltd. | A61B005/0408
US07035679 | B2 | 2006-04-25 | Wavelet-based analysis of pulse oximetry signals | 2002-06-21 | A pulse oximetry signal, suitably a photoplethysmogram (PPG), is decomposed by wavelet transform techniques, and the decomposed signal analysed to provide selected physiological data. The signal may be processed to remove noise, artefacts, or transient features. Information on respiration may also be recovered. | 1. A method of measuring physiological parameters, comprising the steps of: using a pulse oximeter to obtain a pulse oximetry signal;decomposing the pulse oximetry signal by wavelet transform analysis;deriving a continuous wavelet transform surface:plotting the surface against a location parameter and a scale parameter; andfollowing a feature on the plotted surface to derive one or more physiological parameters. | 1. A method of measuring physiological parameters, comprising the steps of: using a pulse oximeter to obtain a pulse oximetry signal;decomposing the pulse oximetry signal by wavelet transform analysis;deriving a continuous wavelet transform surface:plotting the surface against a location parameter and a scale parameter; andfollowing a feature on the plotted surface to derive one or more physiological parameters. 2. A method according to claim 1, in which the pulse oximetry signal is a photoplethysmogram (PPG). 3. A method as claimed in claim 1, in which the step of deriving a continuous wavelet transform surface comprises deriving the wavelet energy surfaces of the pulse oximeter signal. 4. A method as claimed in claim 1, in which the step of deriving a continuous wavelet transform surface includes the step of deriving the wavelet transform modulus of the pulse oximeter signal. 5. A method as claimed in claim 1, in which the scale parameter is a characteristic frequency of the wavelet used in the decomposition. 6. A method as claimed in claim 1, in which the scale parameter is the wavelet dilation. 7. A method as claimed in claim 1, including visually displaying information derived from the pulse oximetry signal by the wavelet transform analysis. 8. A method as claimed in claim 7, in which said information is displayed in real time. 9. A method as claimed in claim 7, in which said information includes one or more of: the distribution of energies within the pulse oximetry signal, coherent structures in the processed signal, a contour plot of the decomposed waveform, a surface plot of the decomposed waveform, and a 2D or 3D energy scalogram. 10. A method according to claim 7, in which the unprocessed pulse oximetry signal is also displayed. 11. A method according to claim 1, in which the step of following a feature on the plotted surface comprises identifying a ridge caused by respiration, and following said ridge to derive information relating to respiration. 12. A method according to claim 11, in which respiration information is derived from high amplitude ridges using ridge-following methods. 13. A method according to claim 11, in which the respiration information is derived by analysis of amplitude or frequency modulation. 14. A method according to claim 11, in which the information relating to respiration comprises a breathing rate. 15. A method according to claim 14, in which respiration information is derived from high amplitude ridges using ridge-following methods. 16. A method according to claim 14, in which the respiration information is derived by analysis of amplitude or frequency modulation. 17. A method according to claim 11, wherein said ridge following comprises the step of extracting phase information along the ridge to identify and/or analyze individual breaths. 18. A method according to claim 17, in which the respiration information is derived by cross scale correlation of phase values. 19. A method according to... | Cardiodigital Limited | A61B005/00
US07196518 | B2 | 2007-03-27 | Magnetic resonance method which automatically forms water/fat separated images with different echo times and determines that proper phase unwrapping has been utilized | 2001-08-13 | A magnetic resonance imaging method for fully automatically forming a water/fat separated image by calculation after acquiring data on images of different echo times, wherein the unwrapping of a phase map showing the distribution of the phase rotation due to the inhomogeneous static magnetic field is repeated so as to determine the distribution of the inhomogeneous static magnetic field by using an index used for judging whether or not the unwrapping is properly being performed, and wherein during the formation of a water/fat separated image with correction of the static magnetic field, the unwrapping is automatically and properly performed in correcting the static magnetic field, and the water/fat images are automatically discriminated. | 1. A magnetic resonance imaging device which is provided with a magnetic field generation means which respectively generates a high frequency magnetic field and a gradient magnetic field in a space where a static magnetic field is formed, a signal processing means which detects a nuclear magnetic resonance signal generated from an inspection subject placed in the space and reconstructs an image therefrom and a display means which displays the reconstructed image, wherein the signal processing means performs, by making use of at least two nuclear magnetic resonance signals having different times TE from irradiation of the high frequency magnetic field signal to generation of the nuclear magnetic resonance signals, a computation determining a phase offset distribution between the generated nuclear magnetic resonance signals, and wherein the signal processing means performs an unwrap processing in which an unwrap processing state map is generated and a principal value of rotation caused in the computation determining the phase offset distribution is corrected, and at the same time, performs a judgement as to whether the unwrap processing is proper based on a consistency of values of a same point obtained by the unwrap processing in at least two different directions of the unwrap processing state map. | 1. A magnetic resonance imaging device which is provided with a magnetic field generation means which respectively generates a high frequency magnetic field and a gradient magnetic field in a space where a static magnetic field is formed, a signal processing means which detects a nuclear magnetic resonance signal generated from an inspection subject placed in the space and reconstructs an image therefrom and a display means which displays the reconstructed image, wherein the signal processing means performs, by making use of at least two nuclear magnetic resonance signals having different times TE from irradiation of the high frequency magnetic field signal to generation of the nuclear magnetic resonance signals, a computation determining a phase offset distribution between the generated nuclear magnetic resonance signals, and wherein the signal processing means performs an unwrap processing in which an unwrap processing state map is generated and a principal value of rotation caused in the computation determining the phase offset distribution is corrected, and at the same time, performs a judgement as to whether the unwrap processing is proper based on a consistency of values of a same point obtained by the unwrap processing in at least two different directions of the unwrap processing state map. 2. A magnetic resonance imaging device according to claim 1, wherein after correcting the nuclear magnetic resonance signals based on the phase offset distribution, in which the unwrap processing has already been judged as having been performed properly, the signal processing means reconstructs another image with regard to two different types of nuclear spins having different chemical shifts through the computation utilizing the generated nuclear magnetic resonance signals. 3. A magnetic resonance imaging method in which an inspection subject is laid in a space where a static magnetic field is formed, a high frequency magnetic field is irradiated onto the inspection subject while applying a gradient field to the static magnetic field, a nuclear magnetic resonance signal generated from the inspection subject in response to the irradiation is detected, an image of a predetermined portion of the inspection subject is reconstructed through a processing means based on the detected signal and the reconstructed image is displayed, said method comprising: a step of calculating first, second and third echo image data by making use of respectively plural first, second and third echo signals each having different echo time from the high frequency magnetic field irradiation to the nuclear magnetic resonance signal generationa step of forming a phase rotation amount distribution map representing a static magnetic field inhomogeneity from the first and third echo image data;a step of applying a first mask prepared based on the first echo image data to the phase rotation amount distribution map;a step of further applying a first loop mask to the phase rotation amoun... | Hitachi Medical Corporation | G01R033/20 G01V003/00 A61B005/055
US06985769 | B2 | 2006-01-10 | Automated realtime interpretation of brain waves | 2002-02-13 | A method and system for automated real time interpretation of brain waves in an acute brain injury of a patient using correlations between brain wave frequency power ratio and wave morphology, determined by a measure of the rhythmicity and variability of the brain wave as a function of the slope of the brain wave upstroke, the arc of the brain wave, and the synchronicity of the brain wave. A system is provided with a central processing unit (16), an electroencephalogram acquisition unit (18), a reference database (24), a quantitative electroencephalogram analysis program (14), and a display device (26) for communicating the classification of the acquired electroencephalographic signals. Artifact rejection is also provided. | 1. A method for automated real time interpretation of brain waves in an acute brain injury of a patient, comprising: determining an alpha-beta/theta-delta power ratio for the brain waves; andcorrelating the brain wave power ratio and wave morphology. | 1. A method for automated real time interpretation of brain waves in an acute brain injury of a patient, comprising: determining an alpha-beta/theta-delta power ratio for the brain waves; andcorrelating the brain wave power ratio and wave morphology. 2. A method for automated real time interpretation of brain waves in an acute brain injury of a patient, comprising: determining a brain wave morphology by measuring the rhythmicity and variability of the brain waves as a function of the slope of the brain waves' upstrokes, the area of the brain waves, and the synchronicity of the brain waves; andcorrelating a brain wave frequency power ratio and the brain wave morphology. 3. A method for automated real time interpretation of brain waves in an acute brain injury of a patient, comprising the steps of: a) acquiring electroencephalographic signals from the patient;b) selecting one or more than one of the acquired electroencephalographic signal for analysis;d) transforming the one or more than one selected electroencephalographic signal into one or more than one frequency component;e) calculating a power ratio for each of the one or more than one selected electroencephalographic signal, wherein the power ratio is an alpha-beta/theta-delta power ratio;f) accessing a reference database storing reference power ratios calculated from normal brain wave activity data;g) comparing the one or more than one calculated power ratio to the reference power ratios to determine whether the one or more than one calculated power ratio is abnormally high or low;h) interpreting the one or more than one selected electroencephalographic signal based upon the comparison of the one or more than one calculated power ratio to the reference power ratios; andi) communicating the interpretation of the selected electroencephalographic signals as the interpretation of the acquired electroencephalographic signals. 4. The method of claim 3, further comprising the step of measuring the wave morphology of the one or more than one selected electroencephalographic signal, and wherein the interpreting step further comprises consideration of the measured wave morphology. 5. The method of claim 3, wherein the reference database further comprises classification data correlating acute brain injury electroencephalographic categories with specific types of brain dysfunction, wherein the method further comprises the step of classifying the one or more than one compared power ratio into one or more than one of the acute brain injury electroencephalographic categories, and wherein the communicating step further comprises communicating the classification. 6. The method of claim 3, wherein the step of comparing further comprises determining the spatial distribution of any abnormality. 7. The method of claim 3, wherein the communicating step further comprises displaying a quadrant schematic representation for communicating the interpretation of the electroencephalographic signals. 8. A method fo... | Jordan Neuroscience, Inc. | A61B005/04
USRE038680 | E1 | 2005-01-04 | Arrangement for determining a heart rate by automatically establishing the optimal refractory time for heart rate determination | 2002-04-30 | An arrangement for determining the heart rate (fRR) or the refractory time of the cardiac tissue, particularly for detecting tachycardia or fibrillation, has an electrode (E) for sensing heart action signals (SIG), an input stage, connected to the electrode, for processing the heart action signals, a refractory member for ascertaining a refractory time value of the arrangement, in each case after a predetermined segment of a heart action signal, and a processing device, connected to the output of the input stage or of the refractory member, for determining the rate of the heart action signals, processed with blanking of the component that occurs during the refractory time, or for determining the refractory time value of the cardiac tissue. The arrangement further includes a device which is connected at least indirectly to the output of the processing device or of the input stage for automatically adjusting the refractory time of the refractory member as a function of a previously determined value of the heart rate or of a signal parameter of a previously determined heart action signal. | 1. A system for determining the heart rate and for detecting a medical condition of a heart, comprising: an electrode for sensing heart action signals; an input stage connected to the electrode for processing the heart action signals; a refractory member for ascertaining a refractory time of the system after a predetermined segment of a heart action signal, said refractory member connected to an output of the input stage; a processing device connected to an output of the input stage for determining the rate of the heart action signals which includes a means responsive to the refractory member for processing the heart action signals by blanking a component of a heart action signal that occurs during the refractory time, and a counter of detected heart signals for determining the heart rate; and adjusting means connected between an output of the processing device and said refractory member for automatically adjusting the refractory time of the refractory member which includes an arithmetic calculation unit for calculating the refractory time (Tref) as a function of a previously determined value of the heart rate in accordance with the following equation (1): Tref=QT=QTc−1.75(Tn−60)Tref=QT=QTc−1.75(TRR−60) (1) wherein Tref is the refractory period, Trr is the previously calculated refractory periodTRR is the time period between adjacent R-peaks of a heart action signal, and QTc is a set period. | 1. A system for determining the heart rate and for detecting a medical condition of a heart, comprising: an electrode for sensing heart action signals; an input stage connected to the electrode for processing the heart action signals; a refractory member for ascertaining a refractory time of the system after a predetermined segment of a heart action signal, said refractory member connected to an output of the input stage; a processing device connected to an output of the input stage for determining the rate of the heart action signals which includes a means responsive to the refractory member for processing the heart action signals by blanking a component of a heart action signal that occurs during the refractory time, and a counter of detected heart signals for determining the heart rate; and adjusting means connected between an output of the processing device and said refractory member for automatically adjusting the refractory time of the refractory member which includes an arithmetic calculation unit for calculating the refractory time (Tref) as a function of a previously determined value of the heart rate in accordance with the following equation (1): Tref=QT=QTc−1.75(Tn−60)Tref=QT=QTc−1.75(TRR−60) (1) wherein Tref is the refractory period, Trr is the previously calculated refractory periodTRR is the time period between adjacent R-peaks of a heart action signal, and QTc is a set period. 2. The system of claim 1, wherein a value slightly greater than the value determined by equation (1) of the arithmetic calculation unit is used to adjust the refractory time. 3. The system of claim 1, wherein the value for QTc is determined for each patient based on an electrophysiological examination. 4. An implantable cardiac pacemaker, for treating tachyarrhythmias, comprising a system according to claim 1. 5. An implantable defibrillator comprising a system according to claim 1. 6. The system of claim 1, wherein the adjusting means includes a refractory time memory addressable at least indirectly through the processing device and connected on an output side to a control input of the refractory member, the memory having a plurality of memory spaces for predetermined refractory time values. | BIOTRONIK Mess-und Therapiegeraete GmbH & Co. Ingenieurbuero Berlin | A61B005/024
US06927843 | B2 | 2005-08-09 | Non-invasive measurement of skin bilirubin level | 2001-03-01 | The invention concerns a method and a device for non-invasive measurement of a tissue and in particular of the skin bilirubin level. The inventive device is characterized in that it comprises: a reading head (1) capable of sending several flashes of various specific wavelengths towards the tissue (2) to be examined and of receiving and measuring in return the reflected light; a calculator, such as a microprocessor, capable of calculating for each wavelength the amount of reflected light and bring it to a value calculated proportionally to a reference value identical for a predetermined wavelength; and a comparator for comparing the calculated value to a table of reference values. | 1. Process for the measurement of the level of a tissue constituent, characterized in that is/are measured a first light reflection with a first given wavelength that is representative of the measurement to be performed and at least a second measurement of the light reflection for another second predetermined wavelength that will serve as a reference unit value Rf, by calculating the coefficient k between the reflection measured on the tissue for this second wavelength and the reflection measured on a standard for this same second wavelength, and in that the level of the constituent to be evaluated is deduced from the measurement of the reflection of the first wavelength according to a table of predetermined values known for this same first wavelength, after correction by the factor k previously measured of said reflection measurement, so as to obtain the value to be compared to the table reference values, said unit value and said second wavelength determining a point through which will go all the LIR curves after standardization. | 1. Process for the measurement of the level of a tissue constituent, characterized in that is/are measured a first light reflection with a first given wavelength that is representative of the measurement to be performed and at least a second measurement of the light reflection for another second predetermined wavelength that will serve as a reference unit value Rf, by calculating the coefficient k between the reflection measured on the tissue for this second wavelength and the reflection measured on a standard for this same second wavelength, and in that the level of the constituent to be evaluated is deduced from the measurement of the reflection of the first wavelength according to a table of predetermined values known for this same first wavelength, after correction by the factor k previously measured of said reflection measurement, so as to obtain the value to be compared to the table reference values, said unit value and said second wavelength determining a point through which will go all the LIR curves after standardization. 2. Process according to claim 1, characterized in that, for an individual's skin, the second wavelength selected is λ=520 nm. 3. Process according to claim 2, characterized in that, for the measurement of the skin bilirubin level, the wavelength selected is λ=460 nm. 4. Device for the measurement of the level of a tissue constituent, characterized in that it comprises: a reading head capable of sending successively several flashes of various predefined wavelengths toward the tissue to be examined as a conical beam having a cone angle between 30° and 50° and of receiving and measuring in return the reflected light, a calculator, capable of calculating for each wavelength the amount of reflected light and of bringing it to a value calculated proportionally to a reference value identical for a predetermined wavelength, a comparator for comparing the calculated value to a table of reference values. 5. Device according to claim 4, characterized in that it comprises electro-luminescent diodes as elements to send successive flashes of various specific wavelengths. 6. Device according to claim 4, characterized in that the light wavelength used to measure the reference value is about 520 nm. 7. Device according to claim 6, characterized in that the light wavelength used to measure the bilirubin level is about 460 nm. 8. Device according to claim 6, characterized in that the light wavelength used to measure the skin pigmentation is comprised between 620 and 780 nm. 9. Device according to claim 6, characterized in that the light wavelength used to measure the skin hemoglobin are two wavelengths or about 545 and 575 nm. 10. Device according to claim 6, characterized in that the light wavelength used to measure the skin hemoglobin is a wavelength of about 550 nm. 11. Device according to claim 6, characterized in that the light wavelength used to measure the skin hemoglobin is a wavelength of about 430 nm. 12. Device acco... | Medick S.A. | G01N033/48 A61B005/00
US06969360 | B1 | 2005-11-29 | Spinal proprioception methods and related systems | 2002-02-13 | A system, device and/or method for measuring spinal proprioception, as can be used to screen proprioceptive loss and/or improvement after treatment. | 1. A clinical method of assessing spiral reposition sense, said method comprising: providing a test subject engaged in thoracolumbar movement, said movement comprising a plurality of flexion positions within a range of motion through the sagittal plane;selecting one said flexion position, said position having a vertical coordinate and a horizontal coordinate;recording a subject reposition in said sagittal plane, said reposition sensed by said subject responsive to said selected position, said reposition having a vertical coordinate and a horizontal coordinate; andcomparing said reposition with said selected position. | 1. A clinical method of assessing spiral reposition sense, said method comprising: providing a test subject engaged in thoracolumbar movement, said movement comprising a plurality of flexion positions within a range of motion through the sagittal plane;selecting one said flexion position, said position having a vertical coordinate and a horizontal coordinate;recording a subject reposition in said sagittal plane, said reposition sensed by said subject responsive to said selected position, said reposition having a vertical coordinate and a horizontal coordinate; andcomparing said reposition with said selected position. 2. The method of claim 1 wherein said selected position and said recorded reposition have a horizontal coordinate measured at C7-T4. 3. The method of claim 2 wherein said horizontal coordinate is measured at the T4 spinal process. 4. The method of claim 1 wherein said subject is asymptomatic. 5. The method of claim 1 wherein said plurality of flexion positions are selected from the group consisting of full flexion range of motion, one-third flexion range of motion and two-thirds flexion range of motion. 6. The method of claim 1 wherein said recorded reposition comprises a recordation average. 7. The method of claim 1, wherein said subject reposition is substantially without cutaneous input. 8. The method of claim 1 wherein said subject reposition is substantially without visual input. 9. A method of using linear coordinates to determine angle of thoracolumbar flexion in the sagittal plane, said method comprising: positioning a subject seated on a substantially level seating surface;providing a vertical reference and measuring a vertical linear coordinate at C7-T4 on said subject;measuring a horizontal linear coordinate from said vertical reference at C7-T4 on said subject, said horizontal coordinate at a right angle to said vertical reference; andtrigonometrically determining an angle of flexion of the thoracolumbar spine in the sagittal plane with said measured linear coordinates. 10. The method of claim 9 wherein said subject is seated at a distance apart from said vertical reference. 11. The method of claim 9 wherein said subject is positioned throughout a range of motion within the sagittal plane. 12. The method of claim 11 wherein said subject is positioned at least one of one-third full flexion and two-thirds full flexion range of motion. 13. The method of claim 9 wherein said determination assesses thoracolumbar reposition sense. 14. A method of using a system for determining thoracolumbar position in the sagittal plane to clinically assess spinal reposition sense, said method comprising: providing a thoracolumbar measurement system having a vertical reference component, a horizontal reference component and a reference housing component, said vertical and horizontal reference components positioned substantially perpendicular one to another within said housing component, and a substantially level seat component sup... | Northwestern University | A61B005/103
US06861045 | B1 | 2005-03-01 | Contrast-enhanced diagnostic imaging method for monitoring interventional therapies | 2000-09-08 | The present invention relates to a contrast-enhanced diagnostic imaging method for monitoring the efficacy of interventional therapies. The contrast agents useful in this method comprise an image-enhancing moiety (IEM) and a state-dependent tissue binding moiety (SDTBM). These contrast agents exhibit state-dependent binding to one or more components of a targeted tissue or tissue component and provide a detectable change in the signal characteristics of the agent once bound to the targeted tissue. As a result, these agents exhibit a binding affinity for, and thus image contrast of, the targeted tissue which changes as the tissue-state changes during therapy. | 1. A method for monitoring treatment of a tissue comprising HSA in a patient, said method comprising: a) administering a contrast agent to said patient, wherein said contrast agent is selected from the group consisting of MS-315, MS-317, MS-322, MS-323, MS-325, MS-326, MS-327, and MS-328; b) subjecting said patient to magnetic resonance imaging to determine an initial signal intensity value in a region of interest of said undesired tissue; c) applying an interventional therapy to at least a portion of said undesired tissue in order to treat said undesired tissue, said interventional therapy selected from the group consisting of a thermal energy generation, a cryoablation, an injection of a denaturing liquid, an injection of a chemotherapeutic agent, and a photodynamic therapy; and d) contemporaneously monitoring with magnetic resonance imaging change in said initial signal intensity value in said region of interest of said undesired tissue during said interventional therapy. | 1. A method for monitoring treatment of a tissue comprising HSA in a patient, said method comprising: a) administering a contrast agent to said patient, wherein said contrast agent is selected from the group consisting of MS-315, MS-317, MS-322, MS-323, MS-325, MS-326, MS-327, and MS-328; b) subjecting said patient to magnetic resonance imaging to determine an initial signal intensity value in a region of interest of said undesired tissue; c) applying an interventional therapy to at least a portion of said undesired tissue in order to treat said undesired tissue, said interventional therapy selected from the group consisting of a thermal energy generation, a cryoablation, an injection of a denaturing liquid, an injection of a chemotherapeutic agent, and a photodynamic therapy; and d) contemporaneously monitoring with magnetic resonance imaging change in said initial signal intensity value in said region of interest of said undesired tissue during said interventional therapy. 2. The method of claim 1, wherein said tissue is selected from the group consisting of cancerous tissue, tumorous tissue, and neoplastic tissue. 3. The method of claim 2, wherein said tissue is cancerous tissue. 4. The method of claim 1, wherein said interventional therapy application is said generation of thermal energy, and wherein said thermal energy is generated by a source selected from the group consisting of one or more focused ultrasound waves, radiofrequency waves, microwaves, and lasers. 5. The method of claim 1, further comprising: e) stopping said interventional therapy application when said change in said initial signal intensity value in said region of interest of said undesired tissue is more than about a 10% reduction in said initial signal intensity value. | - | A61B005/055
US06881191 | B2 | 2005-04-19 | Cardiac monitoring apparatus and method | 2002-12-09 | A monitor includes cardiac and movement sensors 36, 34 responsive to a user's heart beat and a user's movement. The monitor includes a processor coupled to the sensors for generating heart-rate or other cardiac data and user movement or activity data. These data can be stored in a memory and used to analyse the relationship between heart rate and physical exertion. | 1. A monitor for monitoring a user's heart, comprising; a housing; a cardiac sensor, an accelerometer, a memory and a communications module contained in the housing; a processor contained in the housing and coupled to the cardiac sensor, the accelerometer, the memory and the communications module; a support contained in the housing and comprising a first monitor contact coupled to the cardiac sensor and the communications module; a second monitor contact contained in the housing and coupled to the cardiac sensor and the communications module; and an electrical lead having two ends, a first end being couplable to the second monitor contact and a second end comprising a clip for connection to a standard electrocardiogram (ECG) electrode, wherein the monitor is operable in first and second modes such that in a first mode: the support means is removably securable to a first standard adhesive electrocardiogram (ECG) electrode, for securing the housing in position both to provide a first electrical ECG input to the first monitor contact and physically to support the monitor, for sensing the user's heart beat and movement, the lead is removably couplable at its second end to a second standard adhesive ECG electrode and at its first end to the second monitor contact to provide a second electrical ECG input to the second monitor contact, the cardiac sensor receives the first and second ECG inputs from the first and second monitor contacts, and the processor is coupled to the cardiac sensor and the accelerometer for generating cardiac and movement data, and has an output coupled to the memory for storing the cardiac and movement data, and in the second mode: the first and second monitor contacts are respectively couplable to first and second contacts of an interface for downloading the stored cardiac and movement data from the memory under the control of the communications module. | 1. A monitor for monitoring a user's heart, comprising; a housing; a cardiac sensor, an accelerometer, a memory and a communications module contained in the housing; a processor contained in the housing and coupled to the cardiac sensor, the accelerometer, the memory and the communications module; a support contained in the housing and comprising a first monitor contact coupled to the cardiac sensor and the communications module; a second monitor contact contained in the housing and coupled to the cardiac sensor and the communications module; and an electrical lead having two ends, a first end being couplable to the second monitor contact and a second end comprising a clip for connection to a standard electrocardiogram (ECG) electrode, wherein the monitor is operable in first and second modes such that in a first mode: the support means is removably securable to a first standard adhesive electrocardiogram (ECG) electrode, for securing the housing in position both to provide a first electrical ECG input to the first monitor contact and physically to support the monitor, for sensing the user's heart beat and movement, the lead is removably couplable at its second end to a second standard adhesive ECG electrode and at its first end to the second monitor contact to provide a second electrical ECG input to the second monitor contact, the cardiac sensor receives the first and second ECG inputs from the first and second monitor contacts, and the processor is coupled to the cardiac sensor and the accelerometer for generating cardiac and movement data, and has an output coupled to the memory for storing the cardiac and movement data, and in the second mode: the first and second monitor contacts are respectively couplable to first and second contacts of an interface for downloading the stored cardiac and movement data from the memory under the control of the communications module. 2. A monitor according to claim 1, wherein during use in the first mode, the housing is secured to the chest of the user so that the accelerometer is oriented to detect only vertical movements of the chest of the user. 3. A monitor according to claim 2, wherein the housing is marked to indicate correct orientation of the housing on the ECG electrode for use in the first mode. 4. A monitor according to claim 1, wherein the length of the lead is greater than the distance between the ECG electrodes, during use of the monitor in the first mode. 5. A monitor according to claim 1, wherein the processor generates inter-beat interval data from signals it receives from the cardiac sensor. 6. A monitor according to claim 1, wherein the processor processes signals it receives from the cardiac sensor according to a predetermined parameter in order to generate the cardiac data and modifies that parameter in response to signals it receives from the accelerometer. 7. A monitor according to claim 6, wherein the parameter is selected from the group comprising a gain parameter and a thresh... | Cambridge Neurotechnology Limited | A61B005/02 A61B005/04
US06934571 | B2 | 2005-08-23 | Integrated physiologic sensor system | 2002-08-16 | An embodiment of the present invention comprises a self-contained breathing apparatus (SCBA) mask that incorporates a series of physiologic sensors. These physiologic sensors are mounted on the interior of the SCBA mask in such a way as to contact the facial skin of an individual wearing the mask. The physiologic sensors monitor, among other things, heart rate and carbon monoxide and oxygen saturation levels. The physiologic sensors are connected to a transmitting apparatus attached to the SCBA. The SCBA transmitter sends output data from the physiologic sensors to a remote processor. The SCBA transmitter also sends location information of the individual wearing the SCBA mask to the remote processor. The remote processor compiles the physiologic and location data for the wearer of the SCBA and transmits a signal back to the SCBA receiver which, in turn, sends a signal to a display integrated into a SCBA mask. The display indicates to the individual his physiologic status. The display incorporates indicator lights are integrated into the mask of the SCBA to indicate, for example, when bottled air must be switched on, to indicate the emergency situation of the wearer and/or to indicate the direction that the wearer of the SCBA is going. The processor associated with the transmitter also monitors groups of individuals wearing the SCBA masks. In this way, the particular status of several individuals in hazardous environment can be monitored at once and their situation known over a period of time in the hazardous conditions. Specifically, the physiologic condition of the group can be monitored to track the changes in the group condition over the time that the group is exposed to the hazardous environment. | 1. A sensor system, comprising headgear comprising a headband; a first sensor attached to said headband wherein said first sensor detects an amount of carbon monoxide in a user's blood when wearing said headband and produces a first data signal indicative thereof; a second sensor attached to said headband wherein said second sensor detects an amount of oxygen in the user's blood when wearing said headband and produces a second data signal indicative thereof; a processor connected to said first and second sensors and attached to said headband wherein said processor receives the first and second data signals from said first and second sensors, respectively, and creates at least one condition signal indicating a hazard condition, wherein the condition signal is derived from at least one of the first and second data signals, a position sensor system connected to said headband, said position sensor system comprising: a light source that emits a beam of light upon receiving a source signal from said processor; and a light detector that detects the light emitted from said light source after the light reflects off an object, said light detector producing a detector signal for said processor, said processor calculating the distance between said object and said headband. | 1. A sensor system, comprising headgear comprising a headband; a first sensor attached to said headband wherein said first sensor detects an amount of carbon monoxide in a user's blood when wearing said headband and produces a first data signal indicative thereof; a second sensor attached to said headband wherein said second sensor detects an amount of oxygen in the user's blood when wearing said headband and produces a second data signal indicative thereof; a processor connected to said first and second sensors and attached to said headband wherein said processor receives the first and second data signals from said first and second sensors, respectively, and creates at least one condition signal indicating a hazard condition, wherein the condition signal is derived from at least one of the first and second data signals, a position sensor system connected to said headband, said position sensor system comprising: a light source that emits a beam of light upon receiving a source signal from said processor; and a light detector that detects the light emitted from said light source after the light reflects off an object, said light detector producing a detector signal for said processor, said processor calculating the distance between said object and said headband. 2. A sensor system, as in claim 1, further comprising: a transmitter/receiver connected to said processor and attached to said headband wherein said transmitter/receiver receives data from said processor and transmits said data to a base unit, wherein said base unit is located apart from said headband, said base unit having a base processor and storage unit for processing and storing data from a plurality of transmitter/receiver units, and said base unit transmits data to each of said transmitter/receiver units. | BioAsyst, L.L.C. | A61B005/00 G01C003/08
US06944490 | B1 | 2005-09-13 | Apparatus and method for positioning and delivering a therapeutic tool to the inside of a heart | 2002-09-25 | A medical device and method are provided for positioning and supporting a therapeutic tool within the heart or other body cavity. The device has a tubular member through which a cage can be extended. The cage supports the positioning and delivery of a tool to a targeted treatment or diagnosis area within the heart or other body cavity. Markers on the device facilitate accurate positioning and delivery of the tool. | 1. A medical device, comprising: an elongate tubular member for inserting into a body cavity or organ of a patient;at least two support arms capable of extending out from and retracting into the distal end of the elongate tubular member, the support arms converging at the distal end of the support arms and being outwardly bow-shaped when extended out from the elongate tubular member;a tool for performing a therapeutic or diagnostic function extending at least partially along one of the support arms; anda conduit extending at least partially along the support arm along which the tool extends, wherein the conduit houses the tool, the tool being configured to extend out from and retract into the conduit. | 1. A medical device, comprising: an elongate tubular member for inserting into a body cavity or organ of a patient;at least two support arms capable of extending out from and retracting into the distal end of the elongate tubular member, the support arms converging at the distal end of the support arms and being outwardly bow-shaped when extended out from the elongate tubular member;a tool for performing a therapeutic or diagnostic function extending at least partially along one of the support arms; anda conduit extending at least partially along the support arm along which the tool extends, wherein the conduit houses the tool, the tool being configured to extend out from and retract into the conduit. 2. The medical device of claim 1, additionally including a first stop attached to the inner surface of the conduit and a second stop attached to the tool, wherein the tool can be extended out from the conduit until the second stop contacts the first stop. 3. The medical device of claim 1, wherein the distal end of the elongate tubular member includes a segment having a larger diameter than the remaining portion of the elongate tubular member. 4. The medical device of claim 1, wherein the support arms bend to a generally linear shape from the bow-shape configuration when the support arms are retracted into the tubular member. 5. The medical device of claim 1, additionally including a shaft slidably disposed at least partially along the tubular member. 6. The medical device of claim 5, wherein the shaft includes a lumen for receiving a guidewire. 7. The medical device of claim 5, wherein the distal ends of the support arms converge by being attached to the shaft such that the slidable movement of the shaft along the elongate tubular member extends and retracts the support arms out from and into the elongate tubular member. 8. The medical device of claim 1, additionally including a tip member, wherein the distal ends of the support arms converge by being attached to the tip member. 9. The medical device of claim 1, wherein the distal ends of the support arms are coupled to one another. 10. The medical device of claim 1, wherein the support arms are made from an elastic material allowing the support arms to flex into a generally linear configuration from the bow-shaped configuration when the support arms are retracted into the elongate tubular member and allowing the support arms to bend back into the bow-shaped configuration when the support arms are extended out from the tubular member. 11. The medical device of claim 1, wherein the support arms are capable of being rotated with respect to the elongate tubular member. 12. The medical device of claim 1, wherein the support arms support markers for allowing a physician to locate the position of the support arms using an imaging apparatus. 13. The medical device of claim 1, wherein one of the support arms supports a marker having a different shape than a marker supported by another support ar... | Advanced Cardiovascular Systems, Inc. | A61B005/04
US06957095 | B2 | 2005-10-18 | Imaging system for medical diagnosis | 2002-10-04 | An apparatus for medical diagnostic imaging may include a database that relates help information with multiple types of attribute values and stores this information; an operations state acquisition part that acquires the attribute values that expresses the current operational status of the medical diagnostic imaging apparatus; a search part that searches the knowledge database, based on the attribute values; and a display that displays help information or a list of help information, based on the search results of the search part. As the present invention searches for and displays help information according to the operational status of the apparatus for medical diagnostic imaging, it is possible to provide the operator with help information suitable for a diagnosis currently being made. | 1. An apparatus for medical diagnostic imaging, comprising: a first database configured to relate help information to a plurality of attribute values; an operations state acquisition part configured to acquire the attribute values, each of which expresses an aspect of a current operational status of the medical diagnostic imaging apparatus, wherein the current operational status includes information that specifies a program being run by an operator; a search part configured to search the first database, based on the attribute values and generates search result; and a display configured to display at least a portion of the help information based on the search results. | 1. An apparatus for medical diagnostic imaging, comprising: a first database configured to relate help information to a plurality of attribute values; an operations state acquisition part configured to acquire the attribute values, each of which expresses an aspect of a current operational status of the medical diagnostic imaging apparatus, wherein the current operational status includes information that specifies a program being run by an operator; a search part configured to search the first database, based on the attribute values and generates search result; and a display configured to display at least a portion of the help information based on the search results. 2. The apparatus of claim 1, wherein the portion of the help information includes at least one of clinical information for a study region, information on exemplary clinical images, and information relating to at least one caution for the study. 3. The apparatus of claim 1, wherein the current operational status includes information concerning a display form of the image. 4. The apparatus of claim 1, wherein the current operational status comprises at least one of an ultrasonic scan transmission frequency, an imaging mode, a position of focus point, a gain, and a type of probe in an ultrasound diagnostic apparatus. 5. The apparatus of claim 1, wherein the program includes at least one of a patient registration program, an annotation program, a program configured to display at least one body mark, an image file management program, and a measurement program configured to measure values from an image. 6. The apparatus of claim 1, wherein the search part is configured to search the database based on a plurality of attribute values that express study information. 7. The apparatus of claim 6, wherein the study information includes at least one of a study name, a study region, a disease name, a patient name, a doctor name, and a pharmaceutical product used for the study. 8. The apparatus of claim 1, further comprising: a second database configured to store at least one study sequence that defines a workflow of multiple tasks; and a workflow execution part configured to switch the current operational status of the apparatus successively according to the workflow, wherein the search part is configured to search the first database based on information about the workflow. 9. The apparatus of claim 8, further comprising: an analysis part configured to find at least one analysis value determined by an order of the multiple tasks in the workflow, wherein the search part is configured to search the first database, based on the at least one analysis value. 10. An apparatus for medical diagnostic imaging, comprising: a knowledge database configured to store help information associated with scan setting information and body region information; an operational state acquisition part configured to acquire the scan setting information and the body region information; a search part configured to... | Kabushiki Kaisha Toshiba | A61B005/05 A61B008/08
US06885881 | B2 | 2005-04-26 | Device for measuring human blood sugar levels | 2002-01-08 | Device for measuring human blood sugar levels with a catheter, the free end of which is positioned in a blood vessel, wherein the catheter consists of at least one optical waveguide comprising a light source for coupling light into the at least one optical waveguide, a measurement point at the free end of the catheter at which point the light is emitted from the at least one optical waveguide, wherein the light is dispersed by the blood and/or transmitted by the blood and wherein the dispersed and/or transmitted light is coupled again into the minimum of one return optical waveguide, a detector to receive the light which is returned, and a computer unit for analysing the light received by the detector. Provision is made for a cleansing device to be located at the point of measurement for removing the tissue particles deposited from the blood in order to provide a device which is as accurate as possible and which delivers constant measurement values over time for blood sugar levels which can be used as the basis for further data analysis. | 1. Apparatus for measuring human blood sugar levels, comprising: a catheter, the free end of which is to be positioned in a blood vessel, wherein the catheter comprises at least one optical waveguide, a light source for coupling light into the at least one optical waveguide, a point of measurement at the free end of the catheter at which point the light is emitted from the at least one optical waveguide, wherein the light is dispersed by the blood and/or transmitted through the blood and wherein the dispersed and/or transmitted light is coupled again into the at least one optical waveguide, a detector to receive the light which is returned, a computer unit for analysing the light received by the detector, a cleansing device located at the point of measurement for removing the tissue particles deposited from the blood, the cleansing device comprises a controllable actuator which frees or wipes clear a light emission orifice of the at least one optical waveguide when activated, and the cleansing device is controllable by the computer unit. | 1. Apparatus for measuring human blood sugar levels, comprising: a catheter, the free end of which is to be positioned in a blood vessel, wherein the catheter comprises at least one optical waveguide, a light source for coupling light into the at least one optical waveguide, a point of measurement at the free end of the catheter at which point the light is emitted from the at least one optical waveguide, wherein the light is dispersed by the blood and/or transmitted through the blood and wherein the dispersed and/or transmitted light is coupled again into the at least one optical waveguide, a detector to receive the light which is returned, a computer unit for analysing the light received by the detector, a cleansing device located at the point of measurement for removing the tissue particles deposited from the blood, the cleansing device comprises a controllable actuator which frees or wipes clear a light emission orifice of the at least one optical waveguide when activated, and the cleansing device is controllable by the computer unit. 2. Apparatus according to claim 1, wherein a movement of the actuator is generated by a shape memory alloy, a thermopneumatic drive, an electrostatic drive (piezoelement) or a rotor. 3. Apparatus according to claim 1, wherein the energy required for control purposes is supplied to the point of measurement in electrical, thermal, optical, mechanical, hydraulic or pneumatic form. 4. Apparatus according to claim 1, wherein the measurement at the point of measurement is undertaken by means of transmission (irradiation over a specified sampling length) or by means of diffuse reflection from the surface of the catheter. 5. Apparatus according to claim 1, wherein the catheter is clad with a biocompatible material. 6. Apparatus according to claim 5, wherein the catheter and the optical waveguide are designed as a single piece of solid material. 7. Apparatus according to claim 1, wherein the light source, the detector and the computer unit are integrated into an implant together with an energy store for providing voltage. 8. Apparatus according to claim 7, wherein a telemetry unit is integrated into the implant with which it is possible to transmit data and/or energy between the computer unit and a control unit located extracorporeally, wherein for that purpose a probe is connected to the control unit. 9. Apparatus according to claim 8, wherein the telemetry unit and the probe have an oscillatory circuit each with an inductor, wherein the oscillatory circuit on the side of the telemetry unit and the oscillatory circuit on the side of the probe can be tuned for the transmission of data and wherein the data are transmitted without direct contact by coupling the inductors inductively. 10. Apparatus according to claim 8, wherein the energy content of the frequency transmitted by the control unit to the telemetry unit charges an accumulator or a capacitor for the purpose of providing energy to the implant. 11. Appar... | - | A61B005/00
US06846284 | B2 | 2005-01-25 | Method and apparatus for treatment of mono-frequency tinnitus | 2002-12-12 | Reciprocal noise cancellation of a patient's mono-frequency tinnitus tone is achieved utilizing an externally generated tone which is subjectively defined by a mono-frequency tinnitus patient to match his/her tinnitus tone in frequency and amplitude. An externally generated sound wave, selectively designated by subjective observations of a patient to match the patient's tinnitus tone is first applied to the tinnitus patient via earphones or a speaker system and then the same externally generated sound wave is sequentially phase shifted through a plurality of angularly shifted sequence steps to shift or slide the external sound wave through at least a 180 degree phase shift of the generated signal as it is applied to the patient to achieve a series of reductions of the patient's tinnitus tone and in one of such shifted steps a reciprocal, canceling relationship with the patient's tinnitus tone. The tinnitus treatment sequence of an externally generated sound wave and then the phase shifted externally generated tone achieve cancellation of the tinnitus tone of the patient as the sequential steps of the generated tone in effect slide across the tinnitus sound wave resulting in cancellation of the tinnitus tone. By replaying the sequential phase shifted segments of the patient treatment process a patient may utilize the previously recorded sequences in a patient self-treatment process. | 1. A process of treating mono-frequency tinnitus patients comprising the steps of: enabling each patient to subjectively determine his/her tinnitus tone relative to an externally generated sound wave having the same frequency and amplitude as the patient's tinnitus tone, applying to the patient via an audio transducer an externally generated sound wave having substantially the same tone and amplitude as the patient's tinnitus tone, and shifting said externally generated sound wave through a plurality of predetermined angular increments while said externally generated sound wave is applied to said patient whereby a progressive angular phase shift relationship is effected through said successive plurality of angular increments with reference to a predetermined initial point of said externally generated wave form thereby causing at least a partial cancellation or reduction in the amplitude of the patient's tinnitus tone. | 1. A process of treating mono-frequency tinnitus patients comprising the steps of: enabling each patient to subjectively determine his/her tinnitus tone relative to an externally generated sound wave having the same frequency and amplitude as the patient's tinnitus tone, applying to the patient via an audio transducer an externally generated sound wave having substantially the same tone and amplitude as the patient's tinnitus tone, and shifting said externally generated sound wave through a plurality of predetermined angular increments while said externally generated sound wave is applied to said patient whereby a progressive angular phase shift relationship is effected through said successive plurality of angular increments with reference to a predetermined initial point of said externally generated wave form thereby causing at least a partial cancellation or reduction in the amplitude of the patient's tinnitus tone. 2. The process of claim 1 wherein the step of shifting comprises additional steps of continuing said successive incremental shifting steps through a total of at least 180 degree angular phase shift of said external sound wave thus causing both a series of partial amplitude reductions and a reciprocal out-of-phase canceling relationship between the patient's tinnitus tone and said externally generated sound wave. 3. The process of claim 2 wherein each incremental angular phase shift step comprises a succession of substantially identical step segments which are substantially identical in angular displacement and in time duration. 4. The process of claim 2 including an additional step of having each patient maintain a daily tinnitus diary of self-assessment data of the frequency and amplitude of said patient's tinnitus tone following completion of a patient's tinnitus treatment. 5. The process of claim 2 comprising an additional step of performing a patient Brain Scan procedure to generate objective patient data to analyze the status or effectiveness of said patient tinnitus treatment. 6. The process of claim 2 comprising an additional step of electronically recording said externally generated sound wave as it is successively phase-shifted through said plurality of predetermined angular increments through at least said full 180 degree shift of said generated tinnitus treatment sound wave. 7. The process of claim 2 comprising an additional step of enabling a patient to perform a self treatment by playing a recording of said phase-shifted externally generated sound wave. 8. The process of claim 2 where the step of shifting comprises creating sequential phase shift segments having angular phase shifts in the order of six degrees to 30 degrees and each segment has a time duration in the order of one minute to 20 minutes. 9. The process of claim 3 wherein the step of shifting comprises the steps of defining each angular phase shift segment to substantially equal 6 degrees of angular shift and to define each subsequent segment equal... | Tinnitus Control, Inc. | A61M021/00 A61B005/00
US06941165 | B2 | 2005-09-06 | Cardiac magnetic field diagnosing apparatus by late ventricular potential and method of locating intramyocardial excitement uneven propagation portion | 2001-07-17 | A magnetic field distribution measurement device (1) provides a non-contact magnetic measurement on a subject's chest at a plurality of coordinates and forms therefrom time-series magnetic field distribution data. A first arithmetic device (2) in response generates image data representing an intramyocardial excitation conduction rate. A second arithmetic device (3) receives a plurality of tomographic image data separately obtained by a tomographic diagnosis apparatus and processes the data to generate three-dimensional, anatomical image data. A display device (4) receives these data and displays on an anatomical image an image representing an intramyocardial excitation current. This can facilitate identifying an anatomical, positional relationship of a ventricular late potential caused in heart muscle. Furthermore, the anatomical image may be replaced with an image representing a normal stimulation conduction path and serving as a template. | 1. A magnetocardiographic diagnosis apparatus for ventricular late potential, comprising: a magnetic field distribution measurement device performing a non-contact magnetic measurement on a subject's chest at a plurality of coordinates to obtain a plurality of time-series magnetic data corresponding to said plurality of coordinates, and also using said plurality of time-series magnetic data to generate time-series magnetic field distribution data on said chest; a first arithmetic device using said generated time-series magnetic field distribution data to generate data representative of a three-dimensional, intramyocardial, electrical behavior of said subject; a second arithmetic device processing separately provided, tomographic, thoracic data of said subject to generate data representative of a three-dimensional anatomical image; and a display device displaying an image of said three-dimensional, intramyocardial, electrical behavior represented by said data generated by said first arithmetic device, as superimposed on said three-dimensional anatomical image represented by said data generated by said second arithmetic device, thereby capable of three-dimensionally identifying localization of a ventricular late potential attributed to intramyocardial, un-uniform excitation conduction, wherein said data generated by said first arithmetic device and representative of said three-dimensional, intramyocardial, electrical behavior is data representative of an intramyocardial excitation conduction rate and wherein said first arithmetic device approximates by means of one or more small current element pieces a part in myocardium corresponding to an excitation conduction path and calculates a temporal variation of a position of said small current element piece to generate data representative of an intramyocardial excitation conduction rate. | 1. A magnetocardiographic diagnosis apparatus for ventricular late potential, comprising: a magnetic field distribution measurement device performing a non-contact magnetic measurement on a subject's chest at a plurality of coordinates to obtain a plurality of time-series magnetic data corresponding to said plurality of coordinates, and also using said plurality of time-series magnetic data to generate time-series magnetic field distribution data on said chest; a first arithmetic device using said generated time-series magnetic field distribution data to generate data representative of a three-dimensional, intramyocardial, electrical behavior of said subject; a second arithmetic device processing separately provided, tomographic, thoracic data of said subject to generate data representative of a three-dimensional anatomical image; and a display device displaying an image of said three-dimensional, intramyocardial, electrical behavior represented by said data generated by said first arithmetic device, as superimposed on said three-dimensional anatomical image represented by said data generated by said second arithmetic device, thereby capable of three-dimensionally identifying localization of a ventricular late potential attributed to intramyocardial, un-uniform excitation conduction, wherein said data generated by said first arithmetic device and representative of said three-dimensional, intramyocardial, electrical behavior is data representative of an intramyocardial excitation conduction rate and wherein said first arithmetic device approximates by means of one or more small current element pieces a part in myocardium corresponding to an excitation conduction path and calculates a temporal variation of a position of said small current element piece to generate data representative of an intramyocardial excitation conduction rate. 2. The apparatus of claim 1, wherein said first arithmetic device operates based on said calculated temporal variation of the position of said small current element piece to generate data representative of a difference in intramyocardial excitation conduction rate for each excitation conduction path. 3. A magnetocardiographic diagnosis apparatus for ventricular late potential, comprising: a magnetic field distribution measurement device performing a non-contact magnetic measurement on a subject's chest at a plurality of coordinates to obtain a plurality of time-series magnetic data corresponding to said plurality of coordinates, and also using said plurality of time-series magnetic data to generate time-series magnetic field distribution data on said chest; an arithmetic device using said generated time-series magnetic field distribution data to generate data representative of a three-dimensional, intramyocardial, electrical behavior of said subject; and a display device using the data generated by said arithmetic device to superimpose together an image representing a stimulation conduction path of said subject exten... | Japan Science and Technology Corporation Takenaka Corporation | A61B005/04
US06855112 | B2 | 2005-02-15 | Method of and system for health treatment | 2002-02-12 | The present invention relates to a method of and a system for treating a health condition of a user. The method includes determining a sensory signal for use to treat the user's health condition, measuring the change to the sensory signal while the user is engaging in an activity, and using the sensory signal change to regulate the user's activity to thereby treat the user's health condition. The system has a sensor device for measuring a sensory signal of the user while the user is engaging in an activity and a display device receiving the sensory signal measured by the sensor device and displaying the sensory signal. The change to the sensory signal is used to regulate the user's activity to thereby treat the user's health condition. | 1. A method of regulating the mental or psychosomatic condition of a user, comprising: providing a graphic element for the user to compose a graphic production comprising the graphic element; determining a sensory signal for use to treat the user's health condition; measuring a change of the sensory signal while the user is composing the graphic production; and using the sensory signal change to regulate the user's activities in composing the graphic production to thereby regulate the mental or psychosomatic condition to improve the user's health. | 1. A method of regulating the mental or psychosomatic condition of a user, comprising: providing a graphic element for the user to compose a graphic production comprising the graphic element; determining a sensory signal for use to treat the user's health condition; measuring a change of the sensory signal while the user is composing the graphic production; and using the sensory signal change to regulate the user's activities in composing the graphic production to thereby regulate the mental or psychosomatic condition to improve the user's health. 2. The method of claim 1, wherein the sensory signal is determined in accordance to the health condition to be treated. 3. The method of claim 2, wherein the sensory signal is at least one of blood pressure, digital pulse volume, skin conductance, skin temperature, heart rate and pulse rate when the health condition is the user's general health. 4. The method of claim 2, wherein the sensory signal is at least one of digital pulse volume, skin temperature, skin conductance, blood pressure, pulse rate and heart rate when the health condition is the user's emotional health. 5. The method of claim 2, wherein the sensory signal is electroencephalogram when the health condition is the user's cognitive health. 6. The method of claim 1, wherein the sensory signal is selected from the group consisting of blood pressure, digital pulse volume, electroencephalogram, electrocardiogram, electromyogram, galvanic heart rate, pulse rate, respiration, skin conductance, skin response, skin temperature, and any combination thereof. 7. The method of claim 1, wherein the graphic element is selected from the group consisting of characters, alphabets, scripts, numerals, geometric units, non-linguistic forms, and any combination thereof. 8. The method of claim 1 further comprising determining a second sensory signal, and displaying the second sensory signal while the user is composing the graphic production. 9. The method of claim 8 further comprising displaying the sensory signals in at least one of the forms of visual, auditory, tactile, and thermal signals. 10. The method of claim 1, wherein the change in the sensory signal is measured when the user is engaging in a handwriting activity. 11. The method of claim 10, wherein the handwriting activity comprising tracing a graphic element. 12. The method of claim 10, wherein the handwriting activity comprising copying a graphic element. 13. The method of claim 10, wherein the handwriting activity comprising tracing a graphic element, copying a graphic element, and free handwriting. 14. The method of claim 1 wherein the change in the sensory signal is measured when the user is engaging in a free handwriting. 15. The method of claim 1, wherein the change in the sensory signal is measured when the user is handwriting a graphic element. 16. The method of claim 15, wherein the graphic element is selected from the group consisting of characters, alphabets, scripts, nume... | The University of Hong Kong | A61B005/00
US06901285 | B2 | 2005-05-31 | System and method for synthesizing leads of an electrocardiogram | 2002-05-17 | A method for synthesizing electrocardiogram leads includes obtaining a sequence of voltage-time measurements for a set of electrocardiogram leads and subjecting the measurements to abstract factor analysis to obtain a set of eigenvalues and associated eigenvectors. A minimal subset of electrocardiogram leads is identified from which the voltage-time measurements can be calculated with acceptable error. Simplex optimization is performed on a subset of the voltage-time measurements measured with the minimal subset of electrocardiogram leads to obtain a universal transformation matrix, and the universal transformation matrix is multiplied by the subset of the voltage-time measurements to calculate the full set of voltage-time measurements. The full set of leads can be used to calculate a body surface map, and the eigenvalues can be tracked in time to predict the onset of pathology such as myocardial infarction. | 1. A method for synthesizing electrocardiogram leads-comprising the steps of: (a) calculating a univeral transformation matrix by obtaining a sequence of voltage-time measurements for a full set of n-lead electrocardiogram leads for a given n; measuring the voltage-time measurements of the full set of n-lead electrocardiogram leads; and performing simplex optimization on the full set of the voltage-time measurements for an n-lead electrocardiogram to obtain the universal transformation matrix; and (b) multiplying the universal transformation matrix by a minimal subset of the voltage-time measurements to calculate the full set of voltage-time measurements in the n-lead electrocardiogram. | 1. A method for synthesizing electrocardiogram leads-comprising the steps of: (a) calculating a univeral transformation matrix by obtaining a sequence of voltage-time measurements for a full set of n-lead electrocardiogram leads for a given n; measuring the voltage-time measurements of the full set of n-lead electrocardiogram leads; and performing simplex optimization on the full set of the voltage-time measurements for an n-lead electrocardiogram to obtain the universal transformation matrix; and (b) multiplying the universal transformation matrix by a minimal subset of the voltage-time measurements to calculate the full set of voltage-time measurements in the n-lead electrocardiogram. 2. The method of claim 1, further comprising the steps of subjecting the sequence of voltage-time measurements to abstract factor analysis to obtain a set of eigenvalues and associated eigenvectors; and identifying the minimal subset of electrocardiogram leads from which the full set of voltage-time measurements can be calculated with acceptable error. 3. The method of claim 2, further comprising the step of tracking functions of eigenvalues for successive voltage-time measurements in order to predict the onset of pathology, including myocardial infarction. 4. The method of claim 1, further comprising the step of calculating any segment of a cardiac cycle from the universal transformation matrix as applied to the minimal subset of electrocardiogram leads. 5. The method of claim 1,wherein the full set of electrocardiogram leads can comprise from 12 to at least 22 leads. 6. The method of claim 1, wherein the minimal subset of electrocardiogram leads comprises 3 leads. 7. The method of claim 6, wherein the 3 leads are the I, aVF, and V2 leads. 8. The method of claim 6, wherein the 3 leads are the I, II, and V2 leads. 9. The method of claim 6, wherein the 3 leads are the I, aVF, and V9 leads. 10. The method of claim 6, wherein the 3 leads are the I, II, and V9 leads. 11. The method of claim 1, further comprising the step of constructing a body surface map from the calculated full set of voltage-time measurements. 12. The method of claim 1, wherein the full set of electrocardiogram leads can comprise up to 80 or more leads, and further comprising the step of constructing a body surface map from this full set of voltage-time measurements. 13. The method of claim 1, using the technique of cumulative percentage variance to identify the minimal subset of electrocardiogram leads. 14. The method of claim 1, using any 3 measured leads of a conventional n-lead electrocardiogram to derive the complete electrocardiogram. 15. A method for synthesizing electrocardiogram leads comprising the steps of: obtaining a sequence of voltage-time measurements for a full set of from 12 to at least 22 electrocardiogram leads; measuring the voltage-time measurements of the full set of n-lead electrocardiogram leads; subjecting the sequence of voltage-time measurements to abstr... | - | A61B005/0402